## List1
Frequency response measurement template |  |  | measurement properties |  |  | hardware data |  |  |  | Actual number of readings: |  |  |  | 1000
Software will read settings from cells E2 to E7 and |  |  | AC voltage amplitude | 1.414 | V | AC source idn | TEKTRONIX,AFG3101C,C010192,SCPI:99.0 FV:1.0.2 |  |  | Crude estimate of measurement time (h): |  |  |  | 5.8056
measurement points from collumns A-C (yellow cells). |  |  | Sampling frequency | 15000000 | Hz | DC meter idn | HP3458A
Results will be written into collumns D-H (and K+). Sheet must |  |  | AC-DC settle time | 20 | s | Digitizer idn | NI PXI-5922, sn. F22D5B, channel 0
contain keywords START and STOP in collumn A, values of |  |  | AC-DC warm up time | 900 | s
frequency in collumn B, and record length in C. |  |  | DC meter readings | 5
Place keyword STOP to the row with half of the sampling frequency. |  |  | Algorithm | TWM-WRMS
|  |  | Simultaneous meas.? | 1
inputs |  |  | results
Command | Signal frequency | Record length | TIme of DC readings | Digitizer amplitude rms |  | DC voltage |  | Calculation |  | DC meter readings
(START, STOP) | Hz | multiple of periods | (time of writing this line into excel) | value (V) | unc (V) | value (V) | unc (V) | for figure |  | ¦---------------->
START | 15000 | 15000 | 2025-05-09 20:27 | 0.90108 | 0 | 0.75846 | 0.011487 | 15 | 1.0163 | 0.71257 | 0.77223 | 0.76973 | 0.76894 | 0.76881
| 1000 | 1000 | 2025-05-09 20:29 | 0.89871 | 0 | 0.76881 | 8.4188e-07 | #NAME? | #NAME? | 0.7688 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 30000 | 30000 | 2025-05-09 20:30 | 0.8991 | 0 | 0.76881 | 1.0354e-06 | 30 | 1.0004 | 0.76881 | 0.76881 | 0.76881 | 0.7688 | 0.7688
| 1000 | 1000 | 2025-05-09 20:32 | 0.89872 | 0 | 0.7688 | 6.8309e-07 | #NAME? | #NAME? | 0.76881 | 0.7688 | 0.76881 | 0.7688 | 0.7688
| 45000 | 45000 | 2025-05-09 20:34 | 0.89883 | 0 | 0.7688 | 9.3307e-07 | 45 | 1.0002 | 0.7688 | 0.7688 | 0.76881 | 0.76881 | 0.7688
| 1000 | 1000 | 2025-05-09 20:36 | 0.89867 | 0 | 0.7688 | 1.127e-06 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.7688
| 60000 | 60000 | 2025-05-09 20:38 | 0.89885 | 0 | 0.7688 | 1.0521e-06 | 60 | 1.0001 | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.76879
| 1000 | 1000 | 2025-05-09 20:39 | 0.89871 | 0 | 0.76879 | 8.6921e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 75000 | 75000 | 2025-05-09 20:41 | 0.89879 | 0 | 0.76878 | 1.5305e-06 | 75 | 1.0002 | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-09 20:43 | 0.89863 | 0 | 0.76878 | 9.5024e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 90000 | 90000 | 2025-05-09 20:45 | 0.89902 | 0 | 0.76878 | 1.2815e-06 | 90 | 1.0004 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-09 20:46 | 0.89866 | 0 | 0.76877 | 4.4508e-07 | #NAME? | #NAME? | 0.76878 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 105000 | 105000 | 2025-05-09 20:48 | 0.89924 | 0 | 0.76877 | 9.3056e-07 | 105 | 1.0006 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76878
| 1000 | 1000 | 2025-05-09 20:50 | 0.89872 | 0 | 0.76877 | 6.6067e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 120000 | 120000 | 2025-05-09 20:52 | 0.89943 | 0 | 0.76877 | 4.5237e-07 | 120 | 1.0008 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 20:54 | 0.89874 | 0 | 0.76877 | 8.9969e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 135000 | 135000 | 2025-05-09 20:55 | 0.89964 | 0 | 0.76876 | 4.9533e-07 | 135 | 1.001 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 1000 | 1000 | 2025-05-09 20:57 | 0.89871 | 0 | 0.76877 | 7.1352e-07 | #NAME? | #NAME? | 0.76876 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 150000 | 150000 | 2025-05-09 20:59 | 0.89988 | 0 | 0.76876 | 1.0462e-06 | 150 | 1.0014 | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-09 21:01 | 0.89864 | 0 | 0.76876 | 8.8557e-07 | #NAME? | #NAME? | 0.76876 | 0.76877 | 0.76876 | 0.76877 | 0.76876
| 165000 | 165000 | 2025-05-09 21:03 | 0.90012 | 0 | 0.76877 | 6.891e-07 | 165 | 1.0016 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 21:04 | 0.89867 | 0 | 0.76877 | 9.0564e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 180000 | 180000 | 2025-05-09 21:06 | 0.9005 | 0 | 0.76876 | 1.0676e-06 | 180 | 1.002 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:08 | 0.89869 | 0 | 0.76876 | 7.2365e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 195000 | 195000 | 2025-05-09 21:10 | 0.90079 | 0 | 0.76876 | 1.5303e-06 | 195 | 1.0024 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:12 | 0.89866 | 0 | 0.76876 | 5.5352e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 210000 | 210000 | 2025-05-09 21:13 | 0.90113 | 0 | 0.76876 | 1.3491e-06 | 210 | 1.0027 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-09 21:15 | 0.89867 | 0 | 0.76875 | 1.6304e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 225000 | 225000 | 2025-05-09 21:17 | 0.90148 | 0 | 0.76876 | 1.1105e-06 | 225 | 1.0031 | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-09 21:19 | 0.89866 | 0 | 0.76876 | 1.3068e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76875
| 240000 | 240000 | 2025-05-09 21:20 | 0.90188 | 0 | 0.76875 | 1.5737e-06 | 240 | 1.0036 | 0.76875 | 0.76876 | 0.76875 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-09 21:22 | 0.89867 | 0 | 0.76876 | 9.4307e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 255000 | 255000 | 2025-05-09 21:24 | 0.90224 | 0 | 0.76876 | 2.1853e-06 | 255 | 1.004 | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 21:26 | 0.89866 | 0 | 0.76876 | 8.4281e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76876
| 270000 | 270000 | 2025-05-09 21:28 | 0.90266 | 0 | 0.76876 | 1.074e-06 | 270 | 1.0044 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:29 | 0.89868 | 0 | 0.76876 | 1.1507e-06 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 285000 | 285000 | 2025-05-09 21:31 | 0.90298 | 0 | 0.76876 | 6.86e-07 | 285 | 1.0048 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:33 | 0.89869 | 0 | 0.76876 | 1.2678e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76875
| 300000 | 300000 | 2025-05-09 21:35 | 0.90334 | 0 | 0.76875 | 5.9973e-07 | 300 | 1.0052 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 21:36 | 0.8987 | 0 | 0.76876 | 1.2379e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76876
| 315000 | 315000 | 2025-05-09 21:38 | 0.90376 | 0 | 0.76876 | 9.7879e-07 | 315 | 1.0057 | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-09 21:40 | 0.89862 | 0 | 0.76876 | 4.8934e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 330000 | 330000 | 2025-05-09 21:42 | 0.90406 | 0 | 0.76876 | 5.0294e-07 | 330 | 1.0061 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:44 | 0.89862 | 0 | 0.76876 | 4.8276e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 345000 | 345000 | 2025-05-09 21:45 | 0.90448 | 0 | 0.76877 | 1.9511e-06 | 345 | 1.0064 | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-09 21:47 | 0.89871 | 0 | 0.76878 | 8.3502e-07 | #NAME? | #NAME? | 0.76877 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 360000 | 360000 | 2025-05-09 21:49 | 0.90493 | 0 | 0.76877 | 4.0557e-07 | 360 | 1.0069 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 21:51 | 0.89872 | 0 | 0.76877 | 8.2172e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76877
| 375000 | 375000 | 2025-05-09 21:53 | 0.90535 | 0 | 0.76877 | 1.4476e-06 | 375 | 1.0074 | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:54 | 0.89868 | 0 | 0.76876 | 1.0499e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 390000 | 390000 | 2025-05-09 21:56 | 0.90571 | 0 | 0.76876 | 1.1778e-06 | 390 | 1.0078 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 21:58 | 0.8987 | 0 | 0.76876 | 1.0224e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 405000 | 405000 | 2025-05-09 22:00 | 0.90603 | 0 | 0.76875 | 1.2131e-06 | 405 | 1.0081 | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 22:01 | 0.89873 | 0 | 0.76876 | 8.7206e-07 | #NAME? | #NAME? | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 420000 | 420000 | 2025-05-09 22:03 | 0.90639 | 0 | 0.76877 | 1.5282e-06 | 420 | 1.0086 | 0.76876 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 22:05 | 0.8987 | 0 | 0.76877 | 1.2276e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 435000 | 435000 | 2025-05-09 22:07 | 0.90674 | 0 | 0.76876 | 8.4146e-07 | 435 | 1.0089 | 0.76876 | 0.76876 | 0.76876 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-09 22:09 | 0.89871 | 0 | 0.76876 | 9.5323e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 450000 | 450000 | 2025-05-09 22:11 | 0.90699 | 0 | 0.76877 | 2.1399e-06 | 450 | 1.0092 | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 22:12 | 0.89871 | 0 | 0.76877 | 1.4568e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 465000 | 465000 | 2025-05-09 22:14 | 0.90722 | 0 | 0.76876 | 1.7859e-06 | 465 | 1.0095 | 0.76876 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 22:16 | 0.89869 | 0 | 0.76876 | 1.0228e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 480000 | 480000 | 2025-05-09 22:18 | 0.90754 | 0 | 0.76875 | 8.8114e-07 | 480 | 1.0099 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 22:19 | 0.89869 | 0 | 0.76875 | 1.1162e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76875
| 495000 | 495000 | 2025-05-09 22:21 | 0.90775 | 0 | 0.76875 | 1.3241e-06 | 495 | 1.0101 | 0.76875 | 0.76876 | 0.76876 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 22:23 | 0.89868 | 0 | 0.76875 | 5.5004e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 510000 | 510000 | 2025-05-09 22:25 | 0.90788 | 0 | 0.76875 | 7.6089e-07 | 510 | 1.0102 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 22:26 | 0.89875 | 0 | 0.76875 | 5.9753e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 525000 | 525000 | 2025-05-09 22:28 | 0.9081 | 0 | 0.76874 | 5.8686e-07 | 525 | 1.0104 | 0.76875 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 1000 | 1000 | 2025-05-09 22:30 | 0.89878 | 0 | 0.76874 | 1.7209e-06 | #NAME? | #NAME? | 0.76874 | 0.76874 | 0.76875 | 0.76874 | 0.76874
| 540000 | 540000 | 2025-05-09 22:32 | 0.90825 | 0 | 0.76874 | 6.7761e-07 | 540 | 1.0106 | 0.76874 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 1000 | 1000 | 2025-05-09 22:34 | 0.89872 | 0 | 0.76874 | 1.0975e-06 | #NAME? | #NAME? | 0.76874 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 555000 | 555000 | 2025-05-09 22:35 | 0.90831 | 0 | 0.76874 | 1.0657e-06 | 555 | 1.0106 | 0.76873 | 0.76874 | 0.76874 | 0.76874 | 0.76873
| 1000 | 1000 | 2025-05-09 22:37 | 0.89875 | 0 | 0.76874 | 1.0125e-06 | #NAME? | #NAME? | 0.76874 | 0.76874 | 0.76874 | 0.76873 | 0.76873
| 570000 | 570000 | 2025-05-09 22:39 | 0.90827 | 0 | 0.76873 | 8.4281e-07 | 570 | 1.0106 | 0.76874 | 0.76874 | 0.76873 | 0.76873 | 0.76873
| 1000 | 1000 | 2025-05-09 22:41 | 0.89872 | 0 | 0.76873 | 1.4267e-06 | #NAME? | #NAME? | 0.76873 | 0.76873 | 0.76873 | 0.76873 | 0.76873
| 585000 | 585000 | 2025-05-09 22:42 | 0.90829 | 0 | 0.76873 | 6.7498e-07 | 585 | 1.0107 | 0.76873 | 0.76873 | 0.76873 | 0.76873 | 0.76873
| 1000 | 1000 | 2025-05-09 22:44 | 0.89871 | 0 | 0.76873 | 1.1128e-06 | #NAME? | #NAME? | 0.76873 | 0.76873 | 0.76873 | 0.76874 | 0.76874
| 600000 | 600000 | 2025-05-09 22:46 | 0.90824 | 0 | 0.76873 | 5.7121e-07 | 600 | 1.0106 | 0.76873 | 0.76873 | 0.76873 | 0.76873 | 0.76874
| 1000 | 1000 | 2025-05-09 22:48 | 0.89867 | 0 | 0.76873 | 1.0369e-06 | #NAME? | #NAME? | 0.76873 | 0.76874 | 0.76873 | 0.76873 | 0.76873
| 615000 | 615000 | 2025-05-09 22:50 | 0.9081 | 0 | 0.76873 | 1.9259e-07 | 615 | 1.0105 | 0.76873 | 0.76873 | 0.76873 | 0.76873 | 0.76873
| 1000 | 1000 | 2025-05-09 22:51 | 0.8987 | 0 | 0.76873 | 7.3751e-07 | #NAME? | #NAME? | 0.76873 | 0.76874 | 0.76873 | 0.76874 | 0.76874
| 630000 | 630000 | 2025-05-09 22:53 | 0.90799 | 0 | 0.76874 | 1.4518e-06 | 630 | 1.0104 | 0.76873 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 1000 | 1000 | 2025-05-09 22:55 | 0.89868 | 0 | 0.76874 | 5.3558e-07 | #NAME? | #NAME? | 0.76874 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 645000 | 645000 | 2025-05-09 22:57 | 0.90784 | 0 | 0.76874 | 7.0666e-07 | 645 | 1.0102 | 0.76874 | 0.76875 | 0.76874 | 0.76875 | 0.76874
| 1000 | 1000 | 2025-05-09 22:58 | 0.89868 | 0 | 0.76875 | 5.2527e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 660000 | 660000 | 2025-05-09 23:00 | 0.90755 | 0 | 0.76876 | 2.271e-06 | 660 | 1.0099 | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 23:02 | 0.89868 | 0 | 0.76876 | 1.3208e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76875
| 675000 | 675000 | 2025-05-09 23:04 | 0.90729 | 0 | 0.76876 | 6.7766e-07 | 675 | 1.0096 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 23:05 | 0.89866 | 0 | 0.76876 | 6.5711e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 690000 | 690000 | 2025-05-09 23:07 | 0.90705 | 0 | 0.76876 | 1.0763e-06 | 690 | 1.0093 | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-09 23:09 | 0.89867 | 0 | 0.76876 | 6.3545e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 705000 | 705000 | 2025-05-09 23:11 | 0.90662 | 0 | 0.76876 | 7.8556e-07 | 705 | 1.0089 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 23:13 | 0.89862 | 0 | 0.76876 | 1.3144e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 720000 | 720000 | 2025-05-09 23:14 | 0.90626 | 0 | 0.76876 | 1.2766e-06 | 720 | 1.0084 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 23:16 | 0.89871 | 0 | 0.76877 | 7.2497e-07 | #NAME? | #NAME? | 0.76876 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 735000 | 735000 | 2025-05-09 23:18 | 0.90584 | 0 | 0.76876 | 1.4679e-06 | 735 | 1.0079 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-09 23:20 | 0.89871 | 0 | 0.76876 | 2.8147e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 750000 | 750000 | 2025-05-09 23:21 | 0.90545 | 0 | 0.76875 | 5.8686e-07 | 750 | 1.0075 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-09 23:23 | 0.8987 | 0 | 0.76875 | 8.0402e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 765000 | 765000 | 2025-05-09 23:25 | 0.90485 | 0 | 0.76875 | 8.2664e-07 | 765 | 1.0069 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-09 23:27 | 0.89863 | 0 | 0.76875 | 8.5942e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76876 | 0.76875
| 780000 | 780000 | 2025-05-09 23:28 | 0.90435 | 0 | 0.76876 | 1.0883e-06 | 780 | 1.0063 | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-09 23:30 | 0.89866 | 0 | 0.76876 | 5.9503e-07 | #NAME? | #NAME? | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 795000 | 795000 | 2025-05-09 23:32 | 0.90375 | 0 | 0.76876 | 7.8411e-07 | 795 | 1.0057 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-09 23:34 | 0.89862 | 0 | 0.76876 | 1.6227e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76876 | 0.76876 | 0.76876
| 810000 | 810000 | 2025-05-09 23:36 | 0.90314 | 0 | 0.76876 | 7.4765e-07 | 810 | 1.005 | 0.76876 | 0.76876 | 0.76876 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 23:37 | 0.89868 | 0 | 0.76876 | 1.2262e-06 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 825000 | 825000 | 2025-05-09 23:39 | 0.90257 | 0 | 0.76877 | 5.7991e-07 | 825 | 1.0043 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 23:41 | 0.89874 | 0 | 0.76877 | 6.7724e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 840000 | 840000 | 2025-05-09 23:43 | 0.90187 | 0 | 0.76877 | 1.5973e-06 | 840 | 1.0035 | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-09 23:44 | 0.89875 | 0 | 0.76877 | 9.0285e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 855000 | 855000 | 2025-05-09 23:46 | 0.90116 | 0 | 0.76877 | 1.074e-06 | 855 | 1.0027 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-09 23:48 | 0.89874 | 0 | 0.76877 | 1.9509e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76878
| 870000 | 870000 | 2025-05-09 23:50 | 0.90049 | 0 | 0.76878 | 1.2624e-06 | 870 | 1.002 | 0.76879 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-09 23:51 | 0.89871 | 0 | 0.76878 | 1.6212e-06 | #NAME? | #NAME? | 0.76878 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 885000 | 885000 | 2025-05-09 23:53 | 0.8997 | 0 | 0.76878 | 7.1536e-07 | 885 | 1.0012 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-09 23:55 | 0.89866 | 0 | 0.76878 | 6.1742e-07 | #NAME? | #NAME? | 0.76878 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 900000 | 900000 | 2025-05-09 23:57 | 0.89898 | 0 | 0.76878 | 9.876e-07 | 900 | 1.0003 | 0.76877 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-09 23:58 | 0.89872 | 0 | 0.76878 | 2.0833e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76879
| 915000 | 915000 | 2025-05-10 | 0.89825 | 0 | 0.76879 | 1.3115e-06 | 915 | 0.99945 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76878
| 1000 | 1000 | 2025-05-10 00:02 | 0.89875 | 0 | 0.76879 | 2.9645e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.7688 | 0.7688
| 930000 | 930000 | 2025-05-10 00:04 | 0.89743 | 0 | 0.76879 | 8.5922e-07 | 930 | 0.99854 | 0.7688 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 00:06 | 0.89873 | 0 | 0.76879 | 1.4772e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 945000 | 945000 | 2025-05-10 00:07 | 0.8966 | 0 | 0.76879 | 1.1165e-06 | 945 | 0.99761 | 0.76879 | 0.76879 | 0.76878 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 00:09 | 0.89874 | 0 | 0.76878 | 7.8592e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76879
| 960000 | 960000 | 2025-05-10 00:11 | 0.89582 | 0 | 0.76878 | 1.6412e-06 | 960 | 0.99672 | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 00:13 | 0.89877 | 0 | 0.76878 | 6.4846e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 975000 | 975000 | 2025-05-10 00:15 | 0.89502 | 0 | 0.76878 | 1.6127e-06 | 975 | 0.99591 | 0.76878 | 0.76879 | 0.76879 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 00:16 | 0.89869 | 0 | 0.76878 | 1.6622e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 990000 | 990000 | 2025-05-10 00:18 | 0.89421 | 0 | 0.76878 | 8.7831e-07 | 990 | 0.995 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 00:20 | 0.8987 | 0 | 0.76878 | 9.3163e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 1005000 | 1005000 | 2025-05-10 00:22 | 0.89347 | 0 | 0.76878 | 3.6422e-07 | 1005 | 0.99411 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 00:23 | 0.89877 | 0 | 0.76878 | 9.2251e-07 | #NAME? | #NAME? | 0.76878 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 1020000 | 1020000 | 2025-05-10 00:25 | 0.89261 | 0 | 0.76877 | 6.2098e-07 | 1020 | 0.9932 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 00:27 | 0.89872 | 0 | 0.76877 | 1.7953e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 1035000 | 1035000 | 2025-05-10 00:29 | 0.8918 | 0 | 0.76876 | 6.6185e-07 | 1035 | 0.99234 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 00:31 | 0.89869 | 0 | 0.76876 | 4.5977e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1050000 | 1050000 | 2025-05-10 00:32 | 0.89112 | 0 | 0.76877 | 8.4685e-07 | 1050 | 0.99158 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-10 00:34 | 0.89867 | 0 | 0.76876 | 8.5374e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1065000 | 1065000 | 2025-05-10 00:36 | 0.89038 | 0 | 0.76876 | 7.3306e-07 | 1065 | 0.99071 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 00:38 | 0.89873 | 0 | 0.76876 | 9.2004e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1080000 | 1080000 | 2025-05-10 00:40 | 0.88971 | 0 | 0.76876 | 1.3666e-06 | 1080 | 0.98994 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 00:41 | 0.89875 | 0 | 0.76875 | 4.7892e-07 | #NAME? | #NAME? | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1095000 | 1095000 | 2025-05-10 00:43 | 0.88897 | 0 | 0.76876 | 9.7073e-07 | 1095 | 0.98914 | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 00:45 | 0.89873 | 0 | 0.76876 | 8.8436e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 1110000 | 1110000 | 2025-05-10 00:47 | 0.88827 | 0 | 0.76876 | 8.5187e-07 | 1110 | 0.98838 | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-10 00:48 | 0.89871 | 0 | 0.76875 | 1.7903e-06 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1125000 | 1125000 | 2025-05-10 00:50 | 0.88764 | 0 | 0.76875 | 8.5453e-07 | 1125 | 0.98763 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 00:52 | 0.89875 | 0 | 0.76875 | 2.4765e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76874 | 0.76874 | 0.76874
| 1140000 | 1140000 | 2025-05-10 00:54 | 0.88703 | 0 | 0.76874 | 1.1384e-06 | 1140 | 0.98698 | 0.76874 | 0.76874 | 0.76875 | 0.76874 | 0.76875
| 1000 | 1000 | 2025-05-10 00:56 | 0.89873 | 0 | 0.76874 | 1.9186e-06 | #NAME? | #NAME? | 0.76875 | 0.76874 | 0.76874 | 0.76874 | 0.76875
| 1155000 | 1155000 | 2025-05-10 00:57 | 0.88647 | 0 | 0.76875 | 1.2239e-06 | 1155 | 0.98633 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 00:59 | 0.89875 | 0 | 0.76875 | 8.4869e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76874 | 0.76875 | 0.76875
| 1170000 | 1170000 | 2025-05-10 01:01 | 0.88592 | 0 | 0.76875 | 1.3096e-06 | 1170 | 0.98578 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 01:03 | 0.89869 | 0 | 0.76875 | 6.3652e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76874 | 0.76875 | 0.76875
| 1185000 | 1185000 | 2025-05-10 01:05 | 0.88547 | 0 | 0.76875 | 5.7801e-07 | 1185 | 0.98528 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 01:06 | 0.8987 | 0 | 0.76875 | 7.2635e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1200000 | 1200000 | 2025-05-10 01:08 | 0.88504 | 0 | 0.76874 | 6.5419e-07 | 1200 | 0.98475 | 0.76874 | 0.76875 | 0.76875 | 0.76874 | 0.76874
| 1000 | 1000 | 2025-05-10 01:10 | 0.89875 | 0 | 0.76874 | 1.0597e-06 | #NAME? | #NAME? | 0.76875 | 0.76874 | 0.76874 | 0.76874 | 0.76874
| 1215000 | 1215000 | 2025-05-10 01:12 | 0.88456 | 0 | 0.76874 | 1.6336e-06 | 1215 | 0.9843 | 0.76875 | 0.76874 | 0.76874 | 0.76875 | 0.76874
| 1000 | 1000 | 2025-05-10 01:13 | 0.89868 | 0 | 0.76875 | 9.5899e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1230000 | 1230000 | 2025-05-10 01:15 | 0.88421 | 0 | 0.76875 | 1.1666e-06 | 1230 | 0.98391 | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 01:17 | 0.89868 | 0 | 0.76876 | 1.5104e-06 | #NAME? | #NAME? | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1245000 | 1245000 | 2025-05-10 01:19 | 0.88401 | 0 | 0.76876 | 9.1555e-07 | 1245 | 0.98369 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 01:20 | 0.89867 | 0 | 0.76876 | 8.8098e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 1260000 | 1260000 | 2025-05-10 01:22 | 0.88366 | 0 | 0.76877 | 1.8086e-06 | 1260 | 0.98324 | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 01:24 | 0.89872 | 0 | 0.76876 | 1.6156e-06 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1275000 | 1275000 | 2025-05-10 01:26 | 0.88359 | 0 | 0.76876 | 1.2156e-06 | 1275 | 0.98316 | 0.76876 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 01:28 | 0.89872 | 0 | 0.76876 | 7.1411e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1290000 | 1290000 | 2025-05-10 01:29 | 0.88342 | 0 | 0.76875 | 1.3898e-06 | 1290 | 0.98299 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-10 01:31 | 0.89871 | 0 | 0.76876 | 9.1306e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1305000 | 1305000 | 2025-05-10 01:33 | 0.88331 | 0 | 0.76876 | 1.1196e-06 | 1305 | 0.9828 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 01:35 | 0.89876 | 0 | 0.76875 | 7.6457e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76876 | 0.76875 | 0.76875
| 1320000 | 1320000 | 2025-05-10 01:36 | 0.88326 | 0 | 0.76876 | 1.7105e-06 | 1320 | 0.98285 | 0.76875 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 01:38 | 0.89868 | 0 | 0.76876 | 4.3865e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1335000 | 1335000 | 2025-05-10 01:40 | 0.88323 | 0 | 0.76877 | 1.0899e-06 | 1335 | 0.98282 | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76877
| 1000 | 1000 | 2025-05-10 01:42 | 0.89867 | 0 | 0.76877 | 8.1189e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1350000 | 1350000 | 2025-05-10 01:44 | 0.88332 | 0 | 0.76877 | 1.1829e-06 | 1350 | 0.98286 | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 01:45 | 0.89872 | 0 | 0.76877 | 1.2011e-06 | #NAME? | #NAME? | 0.76877 | 0.76878 | 0.76877 | 0.76877 | 0.76877
| 1365000 | 1365000 | 2025-05-10 01:47 | 0.88349 | 0 | 0.76878 | 2.0399e-06 | 1365 | 0.98308 | 0.76877 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 01:49 | 0.8987 | 0 | 0.76878 | 1.2136e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 1380000 | 1380000 | 2025-05-10 01:51 | 0.88368 | 0 | 0.76877 | 1.0105e-06 | 1380 | 0.98326 | 0.76877 | 0.76877 | 0.76878 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 01:52 | 0.89873 | 0 | 0.76877 | 1.0386e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76877 | 0.76877
| 1395000 | 1395000 | 2025-05-10 01:54 | 0.88382 | 0 | 0.76878 | 1.9869e-06 | 1395 | 0.98344 | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 01:56 | 0.89871 | 0 | 0.76878 | 2.0669e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76879 | 0.76878 | 0.76878
| 1410000 | 1410000 | 2025-05-10 01:58 | 0.88411 | 0 | 0.76878 | 1.305e-06 | 1410 | 0.98377 | 0.76878 | 0.76878 | 0.76878 | 0.76879 | 0.76878
| 1000 | 1000 | 2025-05-10 01:59 | 0.8987 | 0 | 0.76879 | 4.2466e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76878 | 0.76878
| 1425000 | 1425000 | 2025-05-10 02:01 | 0.88444 | 0 | 0.76878 | 8.6896e-07 | 1425 | 0.98411 | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:03 | 0.89873 | 0 | 0.76879 | 1.2522e-06 | #NAME? | #NAME? | 0.76878 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1440000 | 1440000 | 2025-05-10 02:05 | 0.88483 | 0 | 0.76879 | 1.4519e-06 | 1440 | 0.98448 | 0.76879 | 0.7688 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 02:06 | 0.89877 | 0 | 0.76879 | 1.1703e-06 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.76879 | 0.76879 | 0.76879
| 1455000 | 1455000 | 2025-05-10 02:08 | 0.88525 | 0 | 0.76879 | 1.0325e-06 | 1455 | 0.98503 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 02:10 | 0.8987 | 0 | 0.76879 | 1.5193e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76878
| 1470000 | 1470000 | 2025-05-10 02:12 | 0.88569 | 0 | 0.76878 | 1.4538e-06 | 1470 | 0.98547 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:13 | 0.89874 | 0 | 0.76877 | 9.0304e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76877 | 0.76878
| 1485000 | 1485000 | 2025-05-10 02:15 | 0.88617 | 0 | 0.76878 | 6.7986e-07 | 1485 | 0.98601 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:17 | 0.89874 | 0 | 0.76878 | 1.1204e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1500000 | 1500000 | 2025-05-10 02:19 | 0.88673 | 0 | 0.76878 | 1.2403e-06 | 1500 | 0.98667 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:21 | 0.89871 | 0 | 0.76878 | 1.1823e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 1515000 | 1515000 | 2025-05-10 02:22 | 0.88731 | 0 | 0.76877 | 1.5124e-06 | 1515 | 0.98725 | 0.76878 | 0.76878 | 0.76877 | 0.76877 | 0.76878
| 1000 | 1000 | 2025-05-10 02:24 | 0.89876 | 0 | 0.76877 | 1.1947e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1530000 | 1530000 | 2025-05-10 02:26 | 0.88787 | 0 | 0.76877 | 1.8587e-06 | 1530 | 0.98797 | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76877
| 1000 | 1000 | 2025-05-10 02:28 | 0.89868 | 0 | 0.76877 | 4.5135e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76877
| 1545000 | 1545000 | 2025-05-10 02:29 | 0.88849 | 0 | 0.76876 | 1.1538e-06 | 1545 | 0.98868 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 02:31 | 0.89865 | 0 | 0.76875 | 6.484e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76876
| 1560000 | 1560000 | 2025-05-10 02:33 | 0.8892 | 0 | 0.76875 | 9.7059e-07 | 1560 | 0.98939 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-10 02:35 | 0.89874 | 0 | 0.76875 | 7.578e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 1575000 | 1575000 | 2025-05-10 02:36 | 0.88994 | 0 | 0.76875 | 7.3447e-07 | 1575 | 0.99024 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 02:38 | 0.89872 | 0 | 0.76875 | 7.7783e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1590000 | 1590000 | 2025-05-10 02:40 | 0.89067 | 0 | 0.76876 | 1.7949e-06 | 1590 | 0.9911 | 0.76876 | 0.76876 | 0.76875 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 02:42 | 0.89868 | 0 | 0.76877 | 3.453e-07 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 1605000 | 1605000 | 2025-05-10 02:44 | 0.89133 | 0 | 0.76877 | 9.3513e-07 | 1605 | 0.99176 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 02:45 | 0.89874 | 0 | 0.76877 | 5.6872e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1620000 | 1620000 | 2025-05-10 02:47 | 0.89214 | 0 | 0.76877 | 1.2942e-06 | 1620 | 0.9927 | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-10 02:49 | 0.8987 | 0 | 0.76878 | 1.0029e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1635000 | 1635000 | 2025-05-10 02:51 | 0.89292 | 0 | 0.76878 | 7.1451e-07 | 1635 | 0.99361 | 0.76878 | 0.76878 | 0.76878 | 0.76879 | 0.76878
| 1000 | 1000 | 2025-05-10 02:52 | 0.89866 | 0 | 0.76878 | 6.1077e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1650000 | 1650000 | 2025-05-10 02:54 | 0.89369 | 0 | 0.76878 | 1.269e-06 | 1650 | 0.9943 | 0.76878 | 0.76879 | 0.76879 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:56 | 0.8988 | 0 | 0.76878 | 9.8832e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1665000 | 1665000 | 2025-05-10 02:58 | 0.89443 | 0 | 0.76878 | 8.1756e-07 | 1665 | 0.99529 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 02:59 | 0.89867 | 0 | 0.76879 | 8.7884e-07 | #NAME? | #NAME? | 0.76879 | 0.76878 | 0.76878 | 0.76878 | 0.76879
| 1680000 | 1680000 | 2025-05-10 03:01 | 0.89519 | 0 | 0.76879 | 7.1689e-07 | 1680 | 0.99609 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 03:03 | 0.8987 | 0 | 0.76879 | 9.3612e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1695000 | 1695000 | 2025-05-10 03:05 | 0.8959 | 0 | 0.76879 | 8.2642e-07 | 1695 | 0.9969 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 03:06 | 0.89869 | 0 | 0.7688 | 1.2319e-06 | #NAME? | #NAME? | 0.76879 | 0.7688 | 0.7688 | 0.7688 | 0.7688
| 1710000 | 1710000 | 2025-05-10 03:08 | 0.89667 | 0 | 0.7688 | 8.3353e-07 | 1710 | 0.99771 | 0.7688 | 0.7688 | 0.76879 | 0.7688 | 0.7688
| 1000 | 1000 | 2025-05-10 03:10 | 0.89874 | 0 | 0.7688 | 7.5573e-07 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.7688
| 1725000 | 1725000 | 2025-05-10 03:12 | 0.89743 | 0 | 0.7688 | 9.0893e-07 | 1725 | 0.99856 | 0.7688 | 0.7688 | 0.76881 | 0.7688 | 0.7688
| 1000 | 1000 | 2025-05-10 03:13 | 0.89873 | 0 | 0.76881 | 1.4508e-06 | #NAME? | #NAME? | 0.7688 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1740000 | 1740000 | 2025-05-10 03:15 | 0.89816 | 0 | 0.76881 | 1.7338e-06 | 1740 | 0.99935 | 0.76881 | 0.7688 | 0.76881 | 0.76881 | 0.76881
| 1000 | 1000 | 2025-05-10 03:17 | 0.89875 | 0 | 0.76881 | 7.5489e-07 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76882
| 1755000 | 1755000 | 2025-05-10 03:19 | 0.89886 | 0 | 0.76881 | 8.1084e-07 | 1755 | 1.0002 | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1000 | 1000 | 2025-05-10 03:20 | 0.89868 | 0 | 0.76881 | 1.4046e-06 | #NAME? | #NAME? | 0.7688 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1770000 | 1770000 | 2025-05-10 03:22 | 0.89952 | 0 | 0.76881 | 6.5228e-07 | 1770 | 1.0009 | 0.76881 | 0.76881 | 0.7688 | 0.76881 | 0.76881
| 1000 | 1000 | 2025-05-10 03:24 | 0.89871 | 0 | 0.76881 | 9.671e-07 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.7688 | 0.76881 | 0.76881
| 1785000 | 1785000 | 2025-05-10 03:26 | 0.90019 | 0 | 0.76881 | 7.3428e-07 | 1785 | 1.0016 | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1000 | 1000 | 2025-05-10 03:27 | 0.89873 | 0 | 0.76881 | 3.681e-07 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1800000 | 1800000 | 2025-05-10 03:29 | 0.90084 | 0 | 0.76881 | 1.0278e-06 | 1800 | 1.0024 | 0.76881 | 0.7688 | 0.76881 | 0.7688 | 0.76881
| 1000 | 1000 | 2025-05-10 03:31 | 0.89871 | 0 | 0.76881 | 1.5584e-06 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1815000 | 1815000 | 2025-05-10 03:33 | 0.90135 | 0 | 0.76882 | 5.6953e-07 | 1815 | 1.0029 | 0.76882 | 0.76882 | 0.76882 | 0.76882 | 0.76882
| 1000 | 1000 | 2025-05-10 03:34 | 0.89872 | 0 | 0.76881 | 8.1259e-07 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1830000 | 1830000 | 2025-05-10 03:36 | 0.90212 | 0 | 0.76881 | 1.2972e-06 | 1830 | 1.0038 | 0.76881 | 0.76881 | 0.7688 | 0.76881 | 0.76881
| 1000 | 1000 | 2025-05-10 03:38 | 0.89872 | 0 | 0.76881 | 1.3362e-06 | #NAME? | #NAME? | 0.76881 | 0.76881 | 0.76881 | 0.76881 | 0.76881
| 1845000 | 1845000 | 2025-05-10 03:40 | 0.90254 | 0 | 0.7688 | 1.4485e-06 | 1845 | 1.0043 | 0.7688 | 0.76881 | 0.7688 | 0.7688 | 0.7688
| 1000 | 1000 | 2025-05-10 03:42 | 0.89869 | 0 | 0.7688 | 9.2539e-07 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.7688
| 1860000 | 1860000 | 2025-05-10 03:44 | 0.90303 | 0 | 0.7688 | 1.1621e-06 | 1860 | 1.0048 | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.76879
| 1000 | 1000 | 2025-05-10 03:46 | 0.89873 | 0 | 0.76879 | 9.9513e-07 | #NAME? | #NAME? | 0.7688 | 0.76879 | 0.7688 | 0.76879 | 0.76879
| 1875000 | 1875000 | 2025-05-10 03:47 | 0.90342 | 0 | 0.76879 | 1.1749e-06 | 1875 | 1.0053 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 03:49 | 0.89868 | 0 | 0.76879 | 1.3751e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.7688 | 0.7688 | 0.76879
| 1890000 | 1890000 | 2025-05-10 03:51 | 0.90386 | 0 | 0.76879 | 1.4033e-06 | 1890 | 1.0058 | 0.76879 | 0.76879 | 0.76878 | 0.76879 | 0.76878
| 1000 | 1000 | 2025-05-10 03:53 | 0.89868 | 0 | 0.76878 | 1.0691e-06 | #NAME? | #NAME? | 0.76878 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 1905000 | 1905000 | 2025-05-10 03:54 | 0.90414 | 0 | 0.76879 | 5.3031e-07 | 1905 | 1.0061 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 03:56 | 0.89868 | 0 | 0.76879 | 1.578e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.7688
| 1920000 | 1920000 | 2025-05-10 03:58 | 0.90453 | 0 | 0.76879 | 1.3964e-06 | 1920 | 1.0064 | 0.7688 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 04:00 | 0.89874 | 0 | 0.76879 | 6.4676e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1935000 | 1935000 | 2025-05-10 04:02 | 0.90474 | 0 | 0.76879 | 9.9855e-07 | 1935 | 1.0067 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 04:03 | 0.8987 | 0 | 0.76879 | 1.9312e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 1950000 | 1950000 | 2025-05-10 04:05 | 0.90491 | 0 | 0.76878 | 1.1358e-06 | 1950 | 1.0069 | 0.76878 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:07 | 0.89875 | 0 | 0.76878 | 1.0959e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 1965000 | 1965000 | 2025-05-10 04:09 | 0.90516 | 0 | 0.76878 | 7.4087e-07 | 1965 | 1.0071 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:10 | 0.89876 | 0 | 0.76878 | 8.2707e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1980000 | 1980000 | 2025-05-10 04:12 | 0.90518 | 0 | 0.76878 | 8.053e-07 | 1980 | 1.0072 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:14 | 0.8987 | 0 | 0.76878 | 1.2942e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76877
| 1995000 | 1995000 | 2025-05-10 04:16 | 0.90526 | 0 | 0.76877 | 9.4533e-07 | 1995 | 1.0072 | 0.76878 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 04:17 | 0.89874 | 0 | 0.76877 | 1.6154e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76878
| 2010000 | 2010000 | 2025-05-10 04:19 | 0.90523 | 0 | 0.76878 | 8.5975e-07 | 2010 | 1.0072 | 0.76877 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 1000 | 1000 | 2025-05-10 04:21 | 0.89873 | 0 | 0.76878 | 9.5218e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 2025000 | 2025000 | 2025-05-10 04:23 | 0.90506 | 0 | 0.76878 | 8.3289e-07 | 2025 | 1.007 | 0.76878 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:24 | 0.89874 | 0 | 0.76878 | 6.0451e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2040000 | 2040000 | 2025-05-10 04:26 | 0.90494 | 0 | 0.76878 | 6.3035e-07 | 2040 | 1.0069 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:28 | 0.89877 | 0 | 0.76878 | 1.1136e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2055000 | 2055000 | 2025-05-10 04:30 | 0.90478 | 0 | 0.76878 | 1.2255e-06 | 2055 | 1.0068 | 0.76878 | 0.76878 | 0.76879 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:32 | 0.89868 | 0 | 0.76879 | 8.5312e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 2070000 | 2070000 | 2025-05-10 04:33 | 0.90446 | 0 | 0.76879 | 1.2973e-06 | 2070 | 1.0064 | 0.76879 | 0.76879 | 0.76879 | 0.7688 | 0.76879
| 1000 | 1000 | 2025-05-10 04:35 | 0.89874 | 0 | 0.7688 | 1.3319e-06 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.76879 | 0.7688 | 0.7688
| 2085000 | 2085000 | 2025-05-10 04:37 | 0.90417 | 0 | 0.7688 | 6.3902e-07 | 2085 | 1.0061 | 0.7688 | 0.7688 | 0.7688 | 0.7688 | 0.7688
| 1000 | 1000 | 2025-05-10 04:39 | 0.89871 | 0 | 0.7688 | 1.6477e-06 | #NAME? | #NAME? | 0.7688 | 0.7688 | 0.76879 | 0.76879 | 0.76879
| 2100000 | 2100000 | 2025-05-10 04:40 | 0.90396 | 0 | 0.76879 | 7.1117e-07 | 2100 | 1.0058 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 04:42 | 0.89876 | 0 | 0.76879 | 1.654e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 2115000 | 2115000 | 2025-05-10 04:44 | 0.90345 | 0 | 0.76879 | 4.7884e-07 | 2115 | 1.0052 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 04:46 | 0.8988 | 0 | 0.76879 | 1.1696e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76879
| 2130000 | 2130000 | 2025-05-10 04:47 | 0.90299 | 0 | 0.76879 | 1.3309e-06 | 2130 | 1.0047 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 04:49 | 0.89876 | 0 | 0.76879 | 7.0197e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76878
| 2145000 | 2145000 | 2025-05-10 04:51 | 0.90249 | 0 | 0.76879 | 6.697e-07 | 2145 | 1.0042 | 0.76879 | 0.76879 | 0.76879 | 0.76878 | 0.76879
| 1000 | 1000 | 2025-05-10 04:53 | 0.89871 | 0 | 0.76878 | 6.1644e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76879 | 0.76878 | 0.76878
| 2160000 | 2160000 | 2025-05-10 04:55 | 0.90197 | 0 | 0.76878 | 1.1061e-06 | 2160 | 1.0036 | 0.76878 | 0.76878 | 0.76879 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 04:56 | 0.89875 | 0 | 0.76878 | 1.7285e-06 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 2175000 | 2175000 | 2025-05-10 04:58 | 0.90131 | 0 | 0.76878 | 1.0974e-06 | 2175 | 1.0029 | 0.76878 | 0.76877 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 05:00 | 0.89872 | 0 | 0.76878 | 1.5616e-06 | #NAME? | #NAME? | 0.76877 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 2190000 | 2190000 | 2025-05-10 05:02 | 0.90072 | 0 | 0.76878 | 6.6996e-07 | 2190 | 1.0021 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 05:03 | 0.89881 | 0 | 0.76878 | 6.3136e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2205000 | 2205000 | 2025-05-10 05:05 | 0.90004 | 0 | 0.76878 | 1.3058e-06 | 2205 | 1.0015 | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:07 | 0.89871 | 0 | 0.76877 | 7.6065e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76877
| 2220000 | 2220000 | 2025-05-10 05:09 | 0.89931 | 0 | 0.76877 | 9.867e-07 | 2220 | 1.0007 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:10 | 0.89868 | 0 | 0.76877 | 9.1197e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 2235000 | 2235000 | 2025-05-10 05:12 | 0.89854 | 0 | 0.76877 | 8.5103e-07 | 2235 | 0.99976 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-10 05:14 | 0.89875 | 0 | 0.76876 | 4.5961e-07 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2250000 | 2250000 | 2025-05-10 05:16 | 0.89775 | 0 | 0.76877 | 1.0018e-06 | 2250 | 0.99888 | 0.76877 | 0.76877 | 0.76876 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:17 | 0.89876 | 0 | 0.76877 | 1.0169e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76877
| 2265000 | 2265000 | 2025-05-10 05:19 | 0.89693 | 0 | 0.76877 | 9.0033e-07 | 2265 | 0.99804 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:21 | 0.89869 | 0 | 0.76877 | 8.5682e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76877
| 2280000 | 2280000 | 2025-05-10 05:23 | 0.8961 | 0 | 0.76877 | 4.5331e-07 | 2280 | 0.99704 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:25 | 0.89876 | 0 | 0.76877 | 9.7329e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 2295000 | 2295000 | 2025-05-10 05:26 | 0.89526 | 0 | 0.76877 | 9.1322e-07 | 2295 | 0.99619 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 05:28 | 0.89867 | 0 | 0.76876 | 1.7896e-06 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76877 | 0.76877 | 0.76876
| 2310000 | 2310000 | 2025-05-10 05:30 | 0.89442 | 0 | 0.76876 | 2.0657e-06 | 2310 | 0.99523 | 0.76877 | 0.76877 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:32 | 0.8987 | 0 | 0.76876 | 4.174e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2325000 | 2325000 | 2025-05-10 05:33 | 0.89348 | 0 | 0.76876 | 5.8321e-07 | 2325 | 0.99419 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:35 | 0.8987 | 0 | 0.76876 | 4.7966e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76876
| 2340000 | 2340000 | 2025-05-10 05:37 | 0.89272 | 0 | 0.76876 | 8.7002e-07 | 2340 | 0.99333 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:39 | 0.89872 | 0 | 0.76876 | 2.0305e-06 | #NAME? | #NAME? | 0.76876 | 0.76877 | 0.76876 | 0.76876 | 0.76875
| 2355000 | 2355000 | 2025-05-10 05:40 | 0.89185 | 0 | 0.76875 | 1.2053e-06 | 2355 | 0.99233 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-10 05:42 | 0.89875 | 0 | 0.76876 | 2.0221e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 2370000 | 2370000 | 2025-05-10 05:44 | 0.89098 | 0 | 0.76876 | 2.0123e-06 | 2370 | 0.99134 | 0.76877 | 0.76877 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:46 | 0.89876 | 0 | 0.76876 | 1.4179e-06 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76876 | 0.76876 | 0.76876
| 2385000 | 2385000 | 2025-05-10 05:47 | 0.89009 | 0 | 0.76876 | 7.4983e-07 | 2385 | 0.99041 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:49 | 0.89871 | 0 | 0.76876 | 1.5347e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2400000 | 2400000 | 2025-05-10 05:51 | 0.88918 | 0 | 0.76876 | 1.1313e-06 | 2400 | 0.9894 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76877
| 1000 | 1000 | 2025-05-10 05:53 | 0.8987 | 0 | 0.76876 | 1.1285e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2415000 | 2415000 | 2025-05-10 05:55 | 0.88832 | 0 | 0.76876 | 7.0389e-07 | 2415 | 0.98846 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 05:56 | 0.89869 | 0 | 0.76876 | 3.0741e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2430000 | 2430000 | 2025-05-10 05:58 | 0.88751 | 0 | 0.76876 | 1.2621e-06 | 2430 | 0.98755 | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-10 06:00 | 0.89869 | 0 | 0.76876 | 1.7049e-06 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2445000 | 2445000 | 2025-05-10 06:02 | 0.88669 | 0 | 0.76876 | 6.7551e-07 | 2445 | 0.98661 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 06:03 | 0.89873 | 0 | 0.76876 | 1.165e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2460000 | 2460000 | 2025-05-10 06:05 | 0.88595 | 0 | 0.76876 | 1.4982e-06 | 2460 | 0.98582 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 06:07 | 0.8987 | 0 | 0.76876 | 1.0058e-06 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76876 | 0.76875
| 2475000 | 2475000 | 2025-05-10 06:09 | 0.88533 | 0 | 0.76875 | 1.6508e-06 | 2475 | 0.98499 | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 06:10 | 0.89881 | 0 | 0.76875 | 5.2044e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 2490000 | 2490000 | 2025-05-10 06:12 | 0.8845 | 0 | 0.76875 | 1.3582e-06 | 2490 | 0.98416 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 06:14 | 0.89874 | 0 | 0.76875 | 3.5293e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 2505000 | 2505000 | 2025-05-10 06:16 | 0.88372 | 0 | 0.76876 | 1.225e-06 | 2505 | 0.98338 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:18 | 0.89865 | 0 | 0.76876 | 5.9884e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2520000 | 2520000 | 2025-05-10 06:19 | 0.88314 | 0 | 0.76876 | 6.2758e-07 | 2520 | 0.98268 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:21 | 0.89871 | 0 | 0.76875 | 8.2771e-07 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 2535000 | 2535000 | 2025-05-10 06:23 | 0.88245 | 0 | 0.76875 | 9.2351e-07 | 2535 | 0.98192 | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76876
| 1000 | 1000 | 2025-05-10 06:25 | 0.8987 | 0 | 0.76875 | 4.6872e-07 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 2550000 | 2550000 | 2025-05-10 06:26 | 0.88194 | 0 | 0.76876 | 9.1578e-07 | 2550 | 0.98132 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:28 | 0.89873 | 0 | 0.76876 | 8.9419e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2565000 | 2565000 | 2025-05-10 06:30 | 0.88141 | 0 | 0.76875 | 5.2864e-07 | 2565 | 0.98075 | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 06:32 | 0.89871 | 0 | 0.76876 | 1.7688e-06 | #NAME? | #NAME? | 0.76876 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 2580000 | 2580000 | 2025-05-10 06:33 | 0.88099 | 0 | 0.76876 | 6.6816e-07 | 2580 | 0.98025 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:35 | 0.89874 | 0 | 0.76876 | 7.2771e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76876
| 2595000 | 2595000 | 2025-05-10 06:37 | 0.8805 | 0 | 0.76876 | 1.08e-06 | 2595 | 0.97979 | 0.76876 | 0.76876 | 0.76875 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:39 | 0.89867 | 0 | 0.76877 | 9.5936e-07 | #NAME? | #NAME? | 0.76876 | 0.76877 | 0.76877 | 0.76876 | 0.76877
| 2610000 | 2610000 | 2025-05-10 06:40 | 0.8802 | 0 | 0.76876 | 2.2222e-06 | 2610 | 0.97938 | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 1000 | 1000 | 2025-05-10 06:42 | 0.89873 | 0 | 0.76876 | 6.5419e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2625000 | 2625000 | 2025-05-10 06:44 | 0.87979 | 0 | 0.76876 | 8.5673e-07 | 2625 | 0.97906 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 06:46 | 0.89861 | 0 | 0.76876 | 2.0638e-06 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76876 | 0.76876
| 2640000 | 2640000 | 2025-05-10 06:48 | 0.8796 | 0 | 0.76876 | 8.4576e-07 | 2640 | 0.97877 | 0.76876 | 0.76877 | 0.76877 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-10 06:49 | 0.89867 | 0 | 0.76876 | 3.6955e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2655000 | 2655000 | 2025-05-10 06:51 | 0.87934 | 0 | 0.76877 | 1.8936e-06 | 2655 | 0.97848 | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 06:53 | 0.89867 | 0 | 0.76877 | 9.1033e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76876 | 0.76877 | 0.76876
| 2670000 | 2670000 | 2025-05-10 06:55 | 0.87925 | 0 | 0.76877 | 1.6329e-06 | 2670 | 0.97841 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 06:56 | 0.89865 | 0 | 0.76877 | 1.0037e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 2685000 | 2685000 | 2025-05-10 06:58 | 0.87928 | 0 | 0.76877 | 6.7472e-07 | 2685 | 0.97839 | 0.76877 | 0.76878 | 0.76878 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 07:00 | 0.89871 | 0 | 0.76878 | 6.4445e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2700000 | 2700000 | 2025-05-10 07:02 | 0.87917 | 0 | 0.76878 | 1.1714e-06 | 2700 | 0.97827 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:03 | 0.8987 | 0 | 0.76879 | 1.4143e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76879 | 0.76879 | 0.76879
| 2715000 | 2715000 | 2025-05-10 07:05 | 0.87923 | 0 | 0.76879 | 7.2791e-07 | 2715 | 0.9783 | 0.76878 | 0.76879 | 0.76878 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 07:07 | 0.89873 | 0 | 0.76879 | 8.0026e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 2730000 | 2730000 | 2025-05-10 07:09 | 0.87939 | 0 | 0.76878 | 8.2427e-07 | 2730 | 0.97849 | 0.76878 | 0.76878 | 0.76878 | 0.76879 | 0.76878
| 1000 | 1000 | 2025-05-10 07:11 | 0.89871 | 0 | 0.76878 | 1.0456e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2745000 | 2745000 | 2025-05-10 07:12 | 0.8796 | 0 | 0.76878 | 1.0428e-06 | 2745 | 0.97864 | 0.76878 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:14 | 0.8988 | 0 | 0.76878 | 3.9113e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2760000 | 2760000 | 2025-05-10 07:16 | 0.87977 | 0 | 0.76878 | 4.2256e-07 | 2760 | 0.9789 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:18 | 0.89874 | 0 | 0.76878 | 1.6011e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76879 | 0.76879
| 2775000 | 2775000 | 2025-05-10 07:19 | 0.87997 | 0 | 0.76878 | 1.4943e-06 | 2775 | 0.97913 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:21 | 0.89873 | 0 | 0.76878 | 1.4116e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2790000 | 2790000 | 2025-05-10 07:23 | 0.88027 | 0 | 0.76878 | 8.9148e-07 | 2790 | 0.97949 | 0.76877 | 0.76878 | 0.76878 | 0.76877 | 0.76878
| 1000 | 1000 | 2025-05-10 07:25 | 0.89871 | 0 | 0.76878 | 6.3175e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2805000 | 2805000 | 2025-05-10 07:27 | 0.88065 | 0 | 0.76878 | 1.1946e-06 | 2805 | 0.97989 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-10 07:28 | 0.89871 | 0 | 0.76877 | 1.2364e-06 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76878 | 0.76877 | 0.76877
| 2820000 | 2820000 | 2025-05-10 07:30 | 0.88114 | 0 | 0.76877 | 1.2288e-06 | 2820 | 0.98053 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 07:32 | 0.89864 | 0 | 0.76878 | 7.4623e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76877
| 2835000 | 2835000 | 2025-05-10 07:34 | 0.88156 | 0 | 0.76877 | 1.322e-06 | 2835 | 0.98091 | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:35 | 0.89872 | 0 | 0.76878 | 1.7215e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 2850000 | 2850000 | 2025-05-10 07:37 | 0.88206 | 0 | 0.76879 | 9.6257e-07 | 2850 | 0.98148 | 0.76879 | 0.76879 | 0.76878 | 0.76879 | 0.76879
| 1000 | 1000 | 2025-05-10 07:39 | 0.89871 | 0 | 0.76879 | 7.5663e-07 | #NAME? | #NAME? | 0.76879 | 0.76879 | 0.76878 | 0.76879 | 0.76879
| 2865000 | 2865000 | 2025-05-10 07:41 | 0.88261 | 0 | 0.76878 | 6.4977e-07 | 2865 | 0.98216 | 0.76879 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:42 | 0.89864 | 0 | 0.76878 | 8.9617e-07 | #NAME? | #NAME? | 0.76878 | 0.76879 | 0.76878 | 0.76878 | 0.76878
| 2880000 | 2880000 | 2025-05-10 07:44 | 0.88326 | 0 | 0.76878 | 1.0591e-06 | 2880 | 0.98272 | 0.76878 | 0.76879 | 0.76878 | 0.76878 | 0.76879
| 1000 | 1000 | 2025-05-10 07:46 | 0.89879 | 0 | 0.76878 | 7.5111e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 2895000 | 2895000 | 2025-05-10 07:48 | 0.88382 | 0 | 0.76878 | 1.0735e-06 | 2895 | 0.98343 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 07:50 | 0.89871 | 0 | 0.76877 | 1.7242e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76877 | 0.76877
| 2910000 | 2910000 | 2025-05-10 07:51 | 0.88447 | 0 | 0.76877 | 1.3535e-06 | 2910 | 0.98411 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76878
| 1000 | 1000 | 2025-05-10 07:53 | 0.89875 | 0 | 0.76877 | 1.1694e-06 | #NAME? | #NAME? | 0.76878 | 0.76877 | 0.76877 | 0.76877 | 0.76878
| 2925000 | 2925000 | 2025-05-10 07:55 | 0.88522 | 0 | 0.76878 | 1.5596e-06 | 2925 | 0.98496 | 0.76877 | 0.76877 | 0.76878 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-10 07:57 | 0.89874 | 0 | 0.76877 | 7.3616e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 2940000 | 2940000 | 2025-05-10 07:58 | 0.88589 | 0 | 0.76877 | 9.7311e-07 | 2940 | 0.98564 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 08:00 | 0.89878 | 0 | 0.76876 | 9.742e-07 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 2955000 | 2955000 | 2025-05-10 08:02 | 0.88667 | 0 | 0.76876 | 1.6175e-06 | 2955 | 0.9866 | 0.76876 | 0.76876 | 0.76877 | 0.76876 | 0.76877
| 1000 | 1000 | 2025-05-10 08:04 | 0.89871 | 0 | 0.76877 | 9.4364e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76877 | 0.76877 | 0.76876
| 2970000 | 2970000 | 2025-05-10 08:06 | 0.88742 | 0 | 0.76877 | 4.2458e-07 | 2970 | 0.98746 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 08:07 | 0.89869 | 0 | 0.76876 | 2.0566e-06 | #NAME? | #NAME? | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76875
| 2985000 | 2985000 | 2025-05-10 08:09 | 0.88817 | 0 | 0.76876 | 8.6732e-07 | 2985 | 0.98822 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 08:11 | 0.89876 | 0 | 0.76876 | 7.3161e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3000000 | 3000000 | 2025-05-10 08:13 | 0.88898 | 0 | 0.76876 | 9.8512e-07 | 3000 | 0.98912 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 08:14 | 0.89876 | 0 | 0.76876 | 8.3604e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76876
| 3015000 | 3015000 | 2025-05-10 08:16 | 0.88979 | 0 | 0.76875 | 1.0768e-06 | 3015 | 0.99007 | 0.76876 | 0.76876 | 0.76875 | 0.76875 | 0.76875
| 1000 | 1000 | 2025-05-10 08:18 | 0.89872 | 0 | 0.76876 | 8.6342e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3030000 | 3030000 | 2025-05-10 08:20 | 0.89047 | 0 | 0.76876 | 1.241e-06 | 3030 | 0.9908 | 0.76876 | 0.76876 | 0.76876 | 0.76877 | 0.76876
| 1000 | 1000 | 2025-05-10 08:22 | 0.89874 | 0 | 0.76876 | 7.3064e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3045000 | 3045000 | 2025-05-10 08:23 | 0.89123 | 0 | 0.76876 | 8.5561e-07 | 3045 | 0.99163 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 08:25 | 0.89876 | 0 | 0.76876 | 2.7178e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76877 | 0.76877
| 3060000 | 3060000 | 2025-05-10 08:27 | 0.89199 | 0 | 0.76877 | 1.5878e-06 | 3060 | 0.99249 | 0.76877 | 0.76876 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 08:29 | 0.89874 | 0 | 0.76877 | 9.117e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 3075000 | 3075000 | 2025-05-10 08:30 | 0.89275 | 0 | 0.76877 | 1.1606e-06 | 3075 | 0.99338 | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 1000 | 1000 | 2025-05-10 08:32 | 0.8987 | 0 | 0.76877 | 7.6624e-07 | #NAME? | #NAME? | 0.76877 | 0.76877 | 0.76877 | 0.76877 | 0.76877
| 3090000 | 3090000 | 2025-05-10 08:34 | 0.89349 | 0 | 0.76877 | 7.8656e-07 | 3090 | 0.99421 | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-10 08:36 | 0.89871 | 0 | 0.76878 | 2.8426e-06 | #NAME? | #NAME? | 0.76877 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 3105000 | 3105000 | 2025-05-10 08:37 | 0.89423 | 0 | 0.76878 | 1.5837e-07 | 3105 | 0.99499 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 08:39 | 0.89873 | 0 | 0.76878 | 1.6004e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76877 | 0.76877
| 3120000 | 3120000 | 2025-05-10 08:41 | 0.89498 | 0 | 0.76877 | 1.3392e-06 | 3120 | 0.99582 | 0.76877 | 0.76877 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 08:43 | 0.89875 | 0 | 0.76878 | 7.5192e-07 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76878
| 3135000 | 3135000 | 2025-05-10 08:44 | 0.8956 | 0 | 0.76878 | 1.7617e-06 | 3135 | 0.99655 | 0.76878 | 0.76878 | 0.76878 | 0.76878 | 0.76877
| 1000 | 1000 | 2025-05-10 08:46 | 0.8987 | 0 | 0.76878 | 1.1991e-06 | #NAME? | #NAME? | 0.76877 | 0.76878 | 0.76877 | 0.76878 | 0.76878
| 3150000 | 3150000 | 2025-05-10 08:48 | 0.89623 | 0 | 0.76878 | 8.4957e-07 | 3150 | 0.99725 | 0.76878 | 0.76877 | 0.76877 | 0.76878 | 0.76878
| 1000 | 1000 | 2025-05-10 08:50 | 0.8987 | 0 | 0.76877 | 2.0512e-06 | #NAME? | #NAME? | 0.76878 | 0.76878 | 0.76877 | 0.76877 | 0.76877
| 3165000 | 3165000 | 2025-05-10 08:52 | 0.89686 | 0 | 0.76876 | 1.0762e-06 | 3165 | 0.99793 | 0.76877 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 08:53 | 0.89872 | 0 | 0.76876 | 6.5943e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3180000 | 3180000 | 2025-05-10 08:55 | 0.8974 | 0 | 0.76876 | 6.065e-07 | 3180 | 0.99855 | 0.76876 | 0.76876 | 0.76877 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 08:57 | 0.89869 | 0 | 0.76876 | 1.1403e-06 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3195000 | 3195000 | 2025-05-10 08:59 | 0.89792 | 0 | 0.76876 | 8.1761e-07 | 3195 | 0.99916 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 09:00 | 0.89867 | 0 | 0.76876 | 7.3156e-07 | #NAME? | #NAME? | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 3210000 | 3210000 | 2025-05-10 09:02 | 0.89849 | 0 | 0.76876 | 4.3581e-07 | 3210 | 0.99976 | 0.76876 | 0.76876 | 0.76876 | 0.76876 | 0.76876
| 1000 | 1000 | 2025-05-10 09:04 | 0.8987 | 0 | 0.76875 | 9.0171e-07 | #NAME? | #NAME? | 0.76875 | 0.76875 | 0.76875 | 0.76875 | 0.76875
| 3225000 | 3225000 | 2025-05-10 09:06 | 0.89895 | 0 | 0.76836 | 0.00040168 | 3225 | 0.99823 | 0.76876 | 0.76876 | 0.76876 | 0.76875 | 0.76675
| 1000 | 1000 | 2025-05-10 09:08 | 0.89873 | 0 | 0.76681 | 5.6096e-05 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76703
| 3240000 | 3240000 | 2025-05-10 09:09 | 0.89934 | 0 | 0.76698 | 5.4115e-05 | 3240 | 1.0004 | 0.76703 | 0.76703 | 0.76703 | 0.76703 | 0.76676
| 1000 | 1000 | 2025-05-10 09:11 | 0.89874 | 0 | 0.76675 | 9.0643e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76672
| 3255000 | 3255000 | 2025-05-10 09:13 | 0.89963 | 0 | 0.76671 | 1.1408e-06 | 3255 | 1.001 | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1000 | 1000 | 2025-05-10 09:15 | 0.89869 | 0 | 0.7667 | 2.1279e-05 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76662
| 3270000 | 3270000 | 2025-05-10 09:16 | 0.90004 | 0 | 0.76665 | 3.0719e-05 | 3270 | 1.0016 | 0.76662 | 0.76662 | 0.76662 | 0.76662 | 0.76678
| 1000 | 1000 | 2025-05-10 09:18 | 0.89867 | 0 | 0.7667 | 6.472e-05 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76676 | 0.76677 | 0.76644
| 3285000 | 3285000 | 2025-05-10 09:20 | 0.90021 | 0 | 0.76649 | 5.0393e-05 | 3285 | 1.0019 | 0.76644 | 0.76644 | 0.76644 | 0.76644 | 0.76669
| 1000 | 1000 | 2025-05-10 09:22 | 0.89871 | 0 | 0.76665 | 4.8911e-05 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76646
| 3300000 | 3300000 | 2025-05-10 09:24 | 0.90043 | 0 | 0.7665 | 5.2208e-05 | 3300 | 1.0021 | 0.76646 | 0.76646 | 0.76645 | 0.76645 | 0.76671
| 1000 | 1000 | 2025-05-10 09:25 | 0.89873 | 0 | 0.76666 | 5.4162e-05 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76645
| 3315000 | 3315000 | 2025-05-10 09:27 | 0.90053 | 0 | 0.76651 | 6.2646e-05 | 3315 | 1.0022 | 0.76645 | 0.76645 | 0.76644 | 0.76644 | 0.76676
| 1000 | 1000 | 2025-05-10 09:29 | 0.89872 | 0 | 0.76669 | 7.5628e-05 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76676 | 0.76639
| 3330000 | 3330000 | 2025-05-10 09:31 | 0.90059 | 0 | 0.76647 | 7.9046e-05 | 3330 | 1.0024 | 0.76639 | 0.76639 | 0.76639 | 0.76639 | 0.76679
| 1000 | 1000 | 2025-05-10 09:32 | 0.89866 | 0 | 0.7667 | 9.3836e-05 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.76632
| 3345000 | 3345000 | 2025-05-10 09:34 | 0.90054 | 0 | 0.76641 | 8.5098e-05 | 3345 | 1.0024 | 0.76632 | 0.76632 | 0.76632 | 0.76632 | 0.76675
| 1000 | 1000 | 2025-05-10 09:36 | 0.89869 | 0 | 0.76666 | 0.00010386 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76624
| 3360000 | 3360000 | 2025-05-10 09:38 | 0.90058 | 0 | 0.76633 | 9.0827e-05 | 3360 | 1.0024 | 0.76624 | 0.76625 | 0.76624 | 0.76624 | 0.7667
| 1000 | 1000 | 2025-05-10 09:39 | 0.89872 | 0 | 0.7666 | 0.00010707 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76617
| 3375000 | 3375000 | 2025-05-10 09:41 | 0.90039 | 0 | 0.76627 | 0.00011186 | 3375 | 1.0024 | 0.76617 | 0.76616 | 0.76616 | 0.76615 | 0.76672
| 1000 | 1000 | 2025-05-10 09:43 | 0.89863 | 0 | 0.76662 | 0.0001122 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76617
| 3390000 | 3390000 | 2025-05-10 09:45 | 0.90021 | 0 | 0.76628 | 0.00011343 | 3390 | 1.0021 | 0.76617 | 0.76617 | 0.76617 | 0.76617 | 0.76673
| 1000 | 1000 | 2025-05-10 09:47 | 0.89868 | 0 | 0.76661 | 0.00013137 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76608
| 3405000 | 3405000 | 2025-05-10 09:48 | 0.90002 | 0 | 0.7662 | 0.00012992 | 3405 | 1.002 | 0.76608 | 0.76607 | 0.76607 | 0.76607 | 0.76672
| 1000 | 1000 | 2025-05-10 09:50 | 0.89864 | 0 | 0.76658 | 0.0001416 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76602
| 3420000 | 3420000 | 2025-05-10 09:52 | 0.89976 | 0 | 0.76615 | 0.00014361 | 3420 | 1.0017 | 0.76602 | 0.76601 | 0.76601 | 0.766 | 0.76673
| 1000 | 1000 | 2025-05-10 09:54 | 0.89869 | 0 | 0.76657 | 0.00015487 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76595
| 3435000 | 3435000 | 2025-05-10 09:55 | 0.8994 | 0 | 0.7661 | 0.00015501 | 3435 | 1.0013 | 0.76595 | 0.76594 | 0.76594 | 0.76594 | 0.76672
| 1000 | 1000 | 2025-05-10 09:57 | 0.89873 | 0 | 0.76655 | 0.00016268 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76671 | 0.76672 | 0.7659
| 3450000 | 3450000 | 2025-05-10 09:59 | 0.89903 | 0 | 0.76606 | 0.00016599 | 3450 | 1.001 | 0.7659 | 0.7659 | 0.7659 | 0.76589 | 0.76673
| 1000 | 1000 | 2025-05-10 10:01 | 0.89872 | 0 | 0.76655 | 0.00018084 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76582
| 3465000 | 3465000 | 2025-05-10 10:03 | 0.89861 | 0 | 0.76599 | 0.00017801 | 3465 | 1.0006 | 0.76582 | 0.76582 | 0.76581 | 0.76582 | 0.76671
| 1000 | 1000 | 2025-05-10 10:04 | 0.89871 | 0 | 0.76653 | 0.00018642 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76578
| 3480000 | 3480000 | 2025-05-10 10:06 | 0.89816 | 0 | 0.76597 | 0.00019172 | 3480 | 1.0001 | 0.76578 | 0.76578 | 0.76577 | 0.76577 | 0.76674
| 1000 | 1000 | 2025-05-10 10:08 | 0.89876 | 0 | 0.76653 | 0.0002189 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76676 | 0.76676 | 0.76566
| 3495000 | 3495000 | 2025-05-10 10:10 | 0.8975 | 0 | 0.76586 | 0.0002185 | 3495 | 0.99946 | 0.76565 | 0.76565 | 0.76564 | 0.76563 | 0.76674
| 1000 | 1000 | 2025-05-10 10:11 | 0.89873 | 0 | 0.7665 | 0.00023992 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76554
| 3510000 | 3510000 | 2025-05-10 10:13 | 0.89689 | 0 | 0.76578 | 0.0002418 | 3510 | 0.99889 | 0.76554 | 0.76554 | 0.76553 | 0.76553 | 0.76674
| 1000 | 1000 | 2025-05-10 10:15 | 0.89872 | 0 | 0.76649 | 0.00025203 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76674 | 0.76549
| 3525000 | 3525000 | 2025-05-10 10:17 | 0.89621 | 0 | 0.76572 | 0.00023723 | 3525 | 0.9982 | 0.76549 | 0.76549 | 0.76548 | 0.76548 | 0.76667
| 1000 | 1000 | 2025-05-10 10:18 | 0.89864 | 0 | 0.76641 | 0.00026635 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76535
| 3540000 | 3540000 | 2025-05-10 10:20 | 0.89562 | 0 | 0.76561 | 0.00026721 | 3540 | 0.99764 | 0.76534 | 0.76534 | 0.76534 | 0.76534 | 0.76667
| 1000 | 1000 | 2025-05-10 10:22 | 0.89869 | 0 | 0.76641 | 0.00027771 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.7653
| 3555000 | 3555000 | 2025-05-10 10:24 | 0.89485 | 0 | 0.76558 | 0.00029069 | 3555 | 0.99683 | 0.7653 | 0.7653 | 0.76529 | 0.76529 | 0.76674
| 1000 | 1000 | 2025-05-10 10:25 | 0.89873 | 0 | 0.76646 | 0.00029406 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76528
| 3570000 | 3570000 | 2025-05-10 10:27 | 0.89408 | 0 | 0.76558 | 0.00029769 | 3570 | 0.99594 | 0.76528 | 0.76528 | 0.76528 | 0.76528 | 0.76677
| 1000 | 1000 | 2025-05-10 10:29 | 0.89876 | 0 | 0.76647 | 0.00030317 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76525
| 3585000 | 3585000 | 2025-05-10 10:31 | 0.89326 | 0 | 0.76555 | 0.00029552 | 3585 | 0.9951 | 0.76525 | 0.76525 | 0.76525 | 0.76526 | 0.76673
| 1000 | 1000 | 2025-05-10 10:33 | 0.89868 | 0 | 0.76643 | 0.00031154 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76518
| 3600000 | 3600000 | 2025-05-10 10:34 | 0.89254 | 0 | 0.76549 | 0.00031403 | 3600 | 0.99434 | 0.76517 | 0.76517 | 0.76517 | 0.76517 | 0.76674
| 1000 | 1000 | 2025-05-10 10:36 | 0.89872 | 0 | 0.76642 | 0.00033172 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76509
| 3615000 | 3615000 | 2025-05-10 10:38 | 0.89163 | 0 | 0.76542 | 0.00033668 | 3615 | 0.99348 | 0.76509 | 0.76509 | 0.76509 | 0.76508 | 0.76677
| 1000 | 1000 | 2025-05-10 10:40 | 0.89867 | 0 | 0.76644 | 0.00034521 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76679 | 0.76505
| 3630000 | 3630000 | 2025-05-10 10:41 | 0.89069 | 0 | 0.76539 | 0.00033817 | 3630 | 0.9924 | 0.76506 | 0.76506 | 0.76506 | 0.76505 | 0.76675
| 1000 | 1000 | 2025-05-10 10:43 | 0.89871 | 0 | 0.76641 | 0.00034549 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76503
| 3645000 | 3645000 | 2025-05-10 10:45 | 0.88978 | 0 | 0.76537 | 0.00034743 | 3645 | 0.99144 | 0.76503 | 0.76502 | 0.76502 | 0.76502 | 0.76676
| 1000 | 1000 | 2025-05-10 10:47 | 0.89867 | 0 | 0.7664 | 0.00036629 | #NAME? | #NAME? | 0.76677 | 0.76676 | 0.76676 | 0.76676 | 0.76493
| 3660000 | 3660000 | 2025-05-10 10:49 | 0.889 | 0 | 0.76529 | 0.00036257 | 3660 | 0.99056 | 0.76493 | 0.76493 | 0.76493 | 0.76493 | 0.76674
| 1000 | 1000 | 2025-05-10 10:50 | 0.89875 | 0 | 0.76638 | 0.00038387 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76677 | 0.76484
| 3675000 | 3675000 | 2025-05-10 10:52 | 0.88795 | 0 | 0.76522 | 0.00038106 | 3675 | 0.98956 | 0.76484 | 0.76484 | 0.76483 | 0.76483 | 0.76674
| 1000 | 1000 | 2025-05-10 10:54 | 0.89868 | 0 | 0.76638 | 0.00037864 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76486
| 3690000 | 3690000 | 2025-05-10 10:56 | 0.88706 | 0 | 0.76523 | 0.00037474 | 3690 | 0.98859 | 0.76486 | 0.76486 | 0.76486 | 0.76485 | 0.76673
| 1000 | 1000 | 2025-05-10 10:57 | 0.89862 | 0 | 0.76636 | 0.00038633 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76675 | 0.76675 | 0.76481
| 3705000 | 3705000 | 2025-05-10 10:59 | 0.88611 | 0 | 0.76519 | 0.00038534 | 3705 | 0.98746 | 0.76481 | 0.7648 | 0.76479 | 0.7648 | 0.76673
| 1000 | 1000 | 2025-05-10 11:01 | 0.8987 | 0 | 0.76633 | 0.00040438 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76471
| 3720000 | 3720000 | 2025-05-10 11:03 | 0.88521 | 0 | 0.76512 | 0.00041611 | 3720 | 0.98657 | 0.76471 | 0.76471 | 0.7647 | 0.7647 | 0.76678
| 1000 | 1000 | 2025-05-10 11:04 | 0.89873 | 0 | 0.76638 | 0.00041294 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.7668 | 0.76473
| 3735000 | 3735000 | 2025-05-10 11:06 | 0.88429 | 0 | 0.76513 | 0.00041675 | 3735 | 0.9856 | 0.76472 | 0.76472 | 0.76471 | 0.76471 | 0.7668
| 1000 | 1000 | 2025-05-10 11:08 | 0.89868 | 0 | 0.76639 | 0.0004216 | #NAME? | #NAME? | 0.76681 | 0.76682 | 0.76682 | 0.76682 | 0.76471
| 3750000 | 3750000 | 2025-05-10 11:10 | 0.88345 | 0 | 0.76511 | 0.0004092 | 3750 | 0.9846 | 0.76471 | 0.7647 | 0.7647 | 0.7647 | 0.76675
| 1000 | 1000 | 2025-05-10 11:12 | 0.89871 | 0 | 0.76635 | 0.00042095 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76466
| 3765000 | 3765000 | 2025-05-10 11:13 | 0.88244 | 0 | 0.76508 | 0.00042714 | 3765 | 0.9835 | 0.76465 | 0.76465 | 0.76465 | 0.76465 | 0.76679
| 1000 | 1000 | 2025-05-10 11:15 | 0.89874 | 0 | 0.76634 | 0.00044897 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.7668 | 0.76455
| 3780000 | 3780000 | 2025-05-10 11:17 | 0.88167 | 0 | 0.76498 | 0.0004436 | 3780 | 0.98272 | 0.76454 | 0.76454 | 0.76453 | 0.76453 | 0.76675
| 1000 | 1000 | 2025-05-10 11:19 | 0.89874 | 0 | 0.76632 | 0.00044868 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76452
| 3795000 | 3795000 | 2025-05-10 11:21 | 0.88077 | 0 | 0.76496 | 0.00044749 | 3795 | 0.98171 | 0.76452 | 0.76452 | 0.76451 | 0.76451 | 0.76675
| 1000 | 1000 | 2025-05-10 11:22 | 0.89875 | 0 | 0.7663 | 0.00045776 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76447
| 3810000 | 3810000 | 2025-05-10 11:24 | 0.87991 | 0 | 0.76491 | 0.00044976 | 3810 | 0.98082 | 0.76447 | 0.76447 | 0.76447 | 0.76446 | 0.76671
| 1000 | 1000 | 2025-05-10 11:26 | 0.89868 | 0 | 0.76625 | 0.00046831 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76438
| 3825000 | 3825000 | 2025-05-10 11:28 | 0.87914 | 0 | 0.76485 | 0.00047421 | 3825 | 0.9801 | 0.76438 | 0.76438 | 0.76437 | 0.76437 | 0.76675
| 1000 | 1000 | 2025-05-10 11:29 | 0.89866 | 0 | 0.76627 | 0.00047743 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76436
| 3840000 | 3840000 | 2025-05-10 11:31 | 0.87845 | 0 | 0.76482 | 0.00047413 | 3840 | 0.97923 | 0.76436 | 0.76435 | 0.76434 | 0.76434 | 0.76672
| 1000 | 1000 | 2025-05-10 11:33 | 0.89876 | 0 | 0.76625 | 0.00047685 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76435
| 3855000 | 3855000 | 2025-05-10 11:35 | 0.87775 | 0 | 0.76481 | 0.00047755 | 3855 | 0.9784 | 0.76434 | 0.76434 | 0.76433 | 0.76433 | 0.76672
| 1000 | 1000 | 2025-05-10 11:37 | 0.89879 | 0 | 0.76623 | 0.00049715 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76674 | 0.76425
| 3870000 | 3870000 | 2025-05-10 11:38 | 0.87701 | 0 | 0.76474 | 0.00049726 | 3870 | 0.97773 | 0.76424 | 0.76424 | 0.76424 | 0.76423 | 0.76672
| 1000 | 1000 | 2025-05-10 11:40 | 0.89874 | 0 | 0.76623 | 0.00050705 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.7642
| 3885000 | 3885000 | 2025-05-10 11:42 | 0.87631 | 0 | 0.7647 | 0.00050298 | 3885 | 0.97701 | 0.7642 | 0.76419 | 0.76419 | 0.76419 | 0.76671
| 1000 | 1000 | 2025-05-10 11:44 | 0.89872 | 0 | 0.76622 | 0.00049906 | #NAME? | #NAME? | 0.76672 | 0.76671 | 0.76672 | 0.76673 | 0.76423
| 3900000 | 3900000 | 2025-05-10 11:46 | 0.87572 | 0 | 0.76471 | 0.00050095 | 3900 | 0.9763 | 0.76422 | 0.76421 | 0.76421 | 0.7642 | 0.76671
| 1000 | 1000 | 2025-05-10 11:47 | 0.89872 | 0 | 0.7662 | 0.00052389 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.7641
| 3915000 | 3915000 | 2025-05-10 11:49 | 0.87512 | 0 | 0.76461 | 0.00051622 | 3915 | 0.97574 | 0.7641 | 0.7641 | 0.76409 | 0.76409 | 0.76668
| 1000 | 1000 | 2025-05-10 11:51 | 0.8987 | 0 | 0.76616 | 0.0005177 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76668 | 0.76409
| 3930000 | 3930000 | 2025-05-10 11:53 | 0.87465 | 0 | 0.76462 | 0.0005358 | 3930 | 0.97529 | 0.76409 | 0.76409 | 0.76408 | 0.76408 | 0.76676
| 1000 | 1000 | 2025-05-10 11:55 | 0.8987 | 0 | 0.76623 | 0.0005374 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76408
| 3945000 | 3945000 | 2025-05-10 11:56 | 0.8741 | 0 | 0.7646 | 0.00053744 | 3945 | 0.9747 | 0.76407 | 0.76406 | 0.76406 | 0.76405 | 0.76675
| 1000 | 1000 | 2025-05-10 11:58 | 0.8987 | 0 | 0.76623 | 0.00053436 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76409
| 3960000 | 3960000 | 2025-05-10 12:00 | 0.87383 | 0 | 0.76462 | 0.00053319 | 3960 | 0.97429 | 0.76408 | 0.76409 | 0.76408 | 0.76408 | 0.76675
| 1000 | 1000 | 2025-05-10 12:02 | 0.89873 | 0 | 0.76619 | 0.00055754 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76396
| 3975000 | 3975000 | 2025-05-10 12:04 | 0.87346 | 0 | 0.76451 | 0.00054404 | 3975 | 0.97397 | 0.76396 | 0.76396 | 0.76396 | 0.76396 | 0.76668
| 1000 | 1000 | 2025-05-10 12:05 | 0.89873 | 0 | 0.76615 | 0.00054511 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.7667 | 0.76397
| 3990000 | 3990000 | 2025-05-10 12:07 | 0.87321 | 0 | 0.76451 | 0.0005498 | 3990 | 0.97363 | 0.76397 | 0.76396 | 0.76396 | 0.76396 | 0.76671
| 1000 | 1000 | 2025-05-10 12:09 | 0.89879 | 0 | 0.76616 | 0.00056049 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76391
| 4005000 | 4005000 | 2025-05-10 12:11 | 0.87294 | 0 | 0.76447 | 0.00055327 | 4005 | 0.97347 | 0.76391 | 0.76391 | 0.76391 | 0.76391 | 0.76668
| 1000 | 1000 | 2025-05-10 12:12 | 0.89868 | 0 | 0.76613 | 0.00056041 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76388
| 4020000 | 4020000 | 2025-05-10 12:14 | 0.87282 | 0 | 0.76445 | 0.00056984 | 4020 | 0.97337 | 0.76389 | 0.76388 | 0.76388 | 0.76388 | 0.76673
| 1000 | 1000 | 2025-05-10 12:16 | 0.89872 | 0 | 0.76617 | 0.00056565 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76391
| 4035000 | 4035000 | 2025-05-10 12:18 | 0.8727 | 0 | 0.76447 | 0.00057423 | 4035 | 0.97321 | 0.76391 | 0.7639 | 0.7639 | 0.76389 | 0.76677
| 1000 | 1000 | 2025-05-10 12:20 | 0.89873 | 0 | 0.76619 | 0.00058571 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76384
| 4050000 | 4050000 | 2025-05-10 12:22 | 0.87266 | 0 | 0.76443 | 0.00059001 | 4050 | 0.9732 | 0.76384 | 0.76384 | 0.76383 | 0.76383 | 0.76679
| 1000 | 1000 | 2025-05-10 12:23 | 0.89876 | 0 | 0.76619 | 0.00059954 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.76379
| 4065000 | 4065000 | 2025-05-10 12:25 | 0.87263 | 0 | 0.76439 | 0.00059701 | 4065 | 0.97328 | 0.76379 | 0.7638 | 0.76379 | 0.76379 | 0.76678
| 1000 | 1000 | 2025-05-10 12:27 | 0.89871 | 0 | 0.76619 | 0.00059571 | #NAME? | #NAME? | 0.76679 | 0.76678 | 0.76678 | 0.76679 | 0.76381
| 4080000 | 4080000 | 2025-05-10 12:29 | 0.87276 | 0 | 0.76438 | 0.00058854 | 4080 | 0.97333 | 0.7638 | 0.7638 | 0.76379 | 0.76379 | 0.76674
| 1000 | 1000 | 2025-05-10 12:30 | 0.89875 | 0 | 0.76615 | 0.00060488 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76373
| 4095000 | 4095000 | 2025-05-10 12:32 | 0.87287 | 0 | 0.76432 | 0.00059557 | 4095 | 0.97352 | 0.76373 | 0.76372 | 0.76372 | 0.76372 | 0.7667
| 1000 | 1000 | 2025-05-10 12:34 | 0.89872 | 0 | 0.76612 | 0.00059638 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76373
| 4110000 | 4110000 | 2025-05-10 12:36 | 0.87305 | 0 | 0.76433 | 0.00060725 | 4110 | 0.97379 | 0.76373 | 0.76373 | 0.76373 | 0.76372 | 0.76676
| 1000 | 1000 | 2025-05-10 12:38 | 0.89869 | 0 | 0.76616 | 0.00060695 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76373
| 4125000 | 4125000 | 2025-05-10 12:40 | 0.87329 | 0 | 0.76433 | 0.0006125 | 4125 | 0.97404 | 0.76373 | 0.76372 | 0.76372 | 0.76371 | 0.76678
| 1000 | 1000 | 2025-05-10 12:41 | 0.89871 | 0 | 0.76617 | 0.00062813 | #NAME? | #NAME? | 0.76678 | 0.7668 | 0.7668 | 0.7668 | 0.76365
| 4140000 | 4140000 | 2025-05-10 12:43 | 0.87357 | 0 | 0.76428 | 0.00062449 | 4140 | 0.97443 | 0.76365 | 0.76365 | 0.76365 | 0.76365 | 0.76677
| 1000 | 1000 | 2025-05-10 12:45 | 0.89867 | 0 | 0.76614 | 0.00063636 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76359
| 4155000 | 4155000 | 2025-05-10 12:47 | 0.87383 | 0 | 0.76422 | 0.00063767 | 4155 | 0.9748 | 0.76359 | 0.76359 | 0.76358 | 0.76357 | 0.76677
| 1000 | 1000 | 2025-05-10 12:48 | 0.89867 | 0 | 0.76614 | 0.00064301 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76679 | 0.76679 | 0.76357
| 4170000 | 4170000 | 2025-05-10 12:50 | 0.87431 | 0 | 0.76421 | 0.00064701 | 4170 | 0.97535 | 0.76357 | 0.76357 | 0.76357 | 0.76356 | 0.7668
| 1000 | 1000 | 2025-05-10 12:52 | 0.89868 | 0 | 0.76615 | 0.00065293 | #NAME? | #NAME? | 0.76681 | 0.7668 | 0.76681 | 0.76681 | 0.76354
| 4185000 | 4185000 | 2025-05-10 12:54 | 0.87474 | 0 | 0.76417 | 0.00063996 | 4185 | 0.97577 | 0.76354 | 0.76353 | 0.76353 | 0.76352 | 0.76673
| 1000 | 1000 | 2025-05-10 12:56 | 0.89871 | 0 | 0.76608 | 0.00065192 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76673 | 0.76674 | 0.76348
| 4200000 | 4200000 | 2025-05-10 12:57 | 0.87526 | 0 | 0.76413 | 0.00064898 | 4200 | 0.97645 | 0.76348 | 0.76348 | 0.76348 | 0.76348 | 0.76672
| 1000 | 1000 | 2025-05-10 12:59 | 0.89868 | 0 | 0.7661 | 0.00064737 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76675 | 0.76675 | 0.76351
| 4215000 | 4215000 | 2025-05-10 13:01 | 0.87576 | 0 | 0.76415 | 0.00066062 | 4215 | 0.97695 | 0.7635 | 0.76349 | 0.7635 | 0.76349 | 0.7668
| 1000 | 1000 | 2025-05-10 13:03 | 0.89873 | 0 | 0.76612 | 0.00068328 | #NAME? | #NAME? | 0.76681 | 0.7668 | 0.7668 | 0.76681 | 0.76339
| 4230000 | 4230000 | 2025-05-10 13:05 | 0.87638 | 0 | 0.76406 | 0.0006738 | 4230 | 0.9777 | 0.76339 | 0.76339 | 0.76339 | 0.76338 | 0.76675
| 1000 | 1000 | 2025-05-10 13:06 | 0.89874 | 0 | 0.76608 | 0.00068681 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76333
| 4245000 | 4245000 | 2025-05-10 13:08 | 0.8769 | 0 | 0.764 | 0.00068061 | 4245 | 0.9783 | 0.76333 | 0.76333 | 0.76332 | 0.76331 | 0.76672
| 1000 | 1000 | 2025-05-10 13:10 | 0.89876 | 0 | 0.76606 | 0.00067526 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76336
| 4260000 | 4260000 | 2025-05-10 13:12 | 0.8776 | 0 | 0.76402 | 0.00067493 | 4260 | 0.97907 | 0.76335 | 0.76335 | 0.76335 | 0.76334 | 0.76672
| 1000 | 1000 | 2025-05-10 13:14 | 0.89874 | 0 | 0.76605 | 0.0006811 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76674 | 0.76332
| 4275000 | 4275000 | 2025-05-10 13:15 | 0.87818 | 0 | 0.76401 | 0.00068889 | 4275 | 0.97983 | 0.76332 | 0.76332 | 0.76332 | 0.76331 | 0.76676
| 1000 | 1000 | 2025-05-10 13:17 | 0.89869 | 0 | 0.76608 | 0.00069184 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76331
| 4290000 | 4290000 | 2025-05-10 13:19 | 0.87894 | 0 | 0.764 | 0.00069433 | 4290 | 0.9807 | 0.76331 | 0.76331 | 0.76331 | 0.76331 | 0.76678
| 1000 | 1000 | 2025-05-10 13:21 | 0.89868 | 0 | 0.76608 | 0.00070932 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.76324
| 4305000 | 4305000 | 2025-05-10 13:23 | 0.87965 | 0 | 0.76394 | 0.00070624 | 4305 | 0.98153 | 0.76324 | 0.76324 | 0.76323 | 0.76323 | 0.76677
| 1000 | 1000 | 2025-05-10 13:24 | 0.89868 | 0 | 0.76605 | 0.00071802 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76318
| 4320000 | 4320000 | 2025-05-10 13:26 | 0.88043 | 0 | 0.76389 | 0.00071554 | 4320 | 0.98241 | 0.76318 | 0.76318 | 0.76317 | 0.76318 | 0.76676
| 1000 | 1000 | 2025-05-10 13:28 | 0.89871 | 0 | 0.76603 | 0.00072301 | #NAME? | #NAME? | 0.76676 | 0.76675 | 0.76676 | 0.76676 | 0.76314
| 4335000 | 4335000 | 2025-05-10 13:30 | 0.8811 | 0 | 0.76386 | 0.00072931 | 4335 | 0.98322 | 0.76314 | 0.76314 | 0.76313 | 0.76313 | 0.76678
| 1000 | 1000 | 2025-05-10 13:32 | 0.8987 | 0 | 0.76605 | 0.00073764 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.7631
| 4350000 | 4350000 | 2025-05-10 13:33 | 0.88182 | 0 | 0.76382 | 0.00073466 | 4350 | 0.98403 | 0.7631 | 0.76309 | 0.76309 | 0.76308 | 0.76676
| 1000 | 1000 | 2025-05-10 13:35 | 0.89872 | 0 | 0.76603 | 0.00074201 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76306
| 4365000 | 4365000 | 2025-05-10 13:37 | 0.88256 | 0 | 0.7638 | 0.00074275 | 4365 | 0.98489 | 0.76306 | 0.76306 | 0.76306 | 0.76305 | 0.76677
| 1000 | 1000 | 2025-05-10 13:39 | 0.89872 | 0 | 0.76603 | 0.00075498 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76679 | 0.76301
| 4380000 | 4380000 | 2025-05-10 13:41 | 0.88335 | 0 | 0.76375 | 0.00074599 | 4380 | 0.98571 | 0.76301 | 0.763 | 0.763 | 0.763 | 0.76673
| 1000 | 1000 | 2025-05-10 13:42 | 0.8988 | 0 | 0.76599 | 0.00074772 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.763
| 4395000 | 4395000 | 2025-05-10 13:44 | 0.88398 | 0 | 0.76374 | 0.00075066 | 4395 | 0.98642 | 0.763 | 0.76299 | 0.76299 | 0.76298 | 0.76674
| 1000 | 1000 | 2025-05-10 13:46 | 0.89879 | 0 | 0.76599 | 0.0007542 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76298
| 4410000 | 4410000 | 2025-05-10 13:48 | 0.88479 | 0 | 0.76373 | 0.00076385 | 4410 | 0.98737 | 0.76297 | 0.76297 | 0.76297 | 0.76296 | 0.76679
| 1000 | 1000 | 2025-05-10 13:49 | 0.89877 | 0 | 0.766 | 0.00078634 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.7668 | 0.76286
| 4425000 | 4425000 | 2025-05-10 13:51 | 0.88532 | 0 | 0.76363 | 0.00077051 | 4425 | 0.98807 | 0.76286 | 0.76286 | 0.76285 | 0.76285 | 0.76671
| 1000 | 1000 | 2025-05-10 13:53 | 0.89873 | 0 | 0.76594 | 0.00078069 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76282
| 4440000 | 4440000 | 2025-05-10 13:55 | 0.88603 | 0 | 0.76359 | 0.00077968 | 4440 | 0.98901 | 0.76282 | 0.76281 | 0.76281 | 0.76281 | 0.76671
| 1000 | 1000 | 2025-05-10 13:57 | 0.89865 | 0 | 0.76595 | 0.00077268 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76286
| 4455000 | 4455000 | 2025-05-10 13:58 | 0.88661 | 0 | 0.76363 | 0.00077288 | 4455 | 0.98953 | 0.76286 | 0.76286 | 0.76286 | 0.76285 | 0.76672
| 1000 | 1000 | 2025-05-10 14:00 | 0.89869 | 0 | 0.76593 | 0.00079604 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76274
| 4470000 | 4470000 | 2025-05-10 14:02 | 0.88736 | 0 | 0.76353 | 0.0008056 | 4470 | 0.9905 | 0.76274 | 0.76273 | 0.76273 | 0.76273 | 0.76676
| 1000 | 1000 | 2025-05-10 14:04 | 0.89873 | 0 | 0.76597 | 0.00078918 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76282
| 4485000 | 4485000 | 2025-05-10 14:06 | 0.88785 | 0 | 0.7636 | 0.00078757 | 4485 | 0.99105 | 0.76281 | 0.76281 | 0.76281 | 0.76281 | 0.76675
| 1000 | 1000 | 2025-05-10 14:07 | 0.89864 | 0 | 0.76596 | 0.00080231 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76275
| 4500000 | 4500000 | 2025-05-10 14:09 | 0.88836 | 0 | 0.76354 | 0.00080156 | 4500 | 0.99161 | 0.76275 | 0.76274 | 0.76274 | 0.76274 | 0.76675
| 1000 | 1000 | 2025-05-10 14:11 | 0.89869 | 0 | 0.76594 | 0.00081223 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76269
| 4515000 | 4515000 | 2025-05-10 14:13 | 0.88885 | 0 | 0.76351 | 0.0008218 | 4515 | 0.9922 | 0.76269 | 0.76269 | 0.76269 | 0.76269 | 0.7668
| 1000 | 1000 | 2025-05-10 14:15 | 0.89873 | 0 | 0.76598 | 0.00083045 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.76681 | 0.76681 | 0.76265
| 4530000 | 4530000 | 2025-05-10 14:16 | 0.88936 | 0 | 0.76346 | 0.00081449 | 4530 | 0.99273 | 0.76265 | 0.76264 | 0.76264 | 0.76264 | 0.76671
| 1000 | 1000 | 2025-05-10 14:18 | 0.89874 | 0 | 0.76589 | 0.00083236 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76257
| 4545000 | 4545000 | 2025-05-10 14:20 | 0.88978 | 0 | 0.76339 | 0.00083196 | 4545 | 0.99333 | 0.76257 | 0.76256 | 0.76255 | 0.76255 | 0.76671
| 1000 | 1000 | 2025-05-10 14:22 | 0.89869 | 0 | 0.76589 | 0.00083299 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76256
| 4560000 | 4560000 | 2025-05-10 14:24 | 0.89013 | 0 | 0.76339 | 0.00084244 | 4560 | 0.99375 | 0.76256 | 0.76255 | 0.76254 | 0.76254 | 0.76676
| 1000 | 1000 | 2025-05-10 14:25 | 0.8987 | 0 | 0.76592 | 0.00084329 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76255
| 4575000 | 4575000 | 2025-05-10 14:27 | 0.89052 | 0 | 0.76338 | 0.00083881 | 4575 | 0.99411 | 0.76254 | 0.76254 | 0.76254 | 0.76254 | 0.76674
| 1000 | 1000 | 2025-05-10 14:29 | 0.89877 | 0 | 0.76591 | 0.00084667 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76676 | 0.76676 | 0.76252
| 4590000 | 4590000 | 2025-05-10 14:31 | 0.89073 | 0 | 0.76337 | 0.00084758 | 4590 | 0.99443 | 0.76252 | 0.76252 | 0.76252 | 0.76252 | 0.76676
| 1000 | 1000 | 2025-05-10 14:33 | 0.89869 | 0 | 0.7659 | 0.00086982 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76242
| 4605000 | 4605000 | 2025-05-10 14:34 | 0.89089 | 0 | 0.76328 | 0.00086675 | 4605 | 0.99474 | 0.76241 | 0.76241 | 0.76242 | 0.76242 | 0.76675
| 1000 | 1000 | 2025-05-10 14:36 | 0.89866 | 0 | 0.76588 | 0.00087412 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76239
| 4620000 | 4620000 | 2025-05-10 14:38 | 0.89094 | 0 | 0.76325 | 0.00087274 | 4620 | 0.99483 | 0.76238 | 0.76238 | 0.76238 | 0.76238 | 0.76674
| 1000 | 1000 | 2025-05-10 14:40 | 0.89864 | 0 | 0.76587 | 0.00087828 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76236
| 4635000 | 4635000 | 2025-05-10 14:42 | 0.89102 | 0 | 0.76323 | 0.00088493 | 4635 | 0.99493 | 0.76236 | 0.76235 | 0.76234 | 0.76234 | 0.76677
| 1000 | 1000 | 2025-05-10 14:44 | 0.89868 | 0 | 0.76589 | 0.00089022 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76233
| 4650000 | 4650000 | 2025-05-10 14:45 | 0.89108 | 0 | 0.76322 | 0.00089468 | 4650 | 0.995 | 0.76232 | 0.76232 | 0.76232 | 0.76232 | 0.7668
| 1000 | 1000 | 2025-05-10 14:47 | 0.8987 | 0 | 0.7659 | 0.00090268 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.76681 | 0.76681 | 0.76229
| 4665000 | 4665000 | 2025-05-10 14:49 | 0.89101 | 0 | 0.76317 | 0.00089165 | 4665 | 0.99488 | 0.76228 | 0.76228 | 0.76227 | 0.76227 | 0.76673
| 1000 | 1000 | 2025-05-10 14:51 | 0.89873 | 0 | 0.76584 | 0.00090118 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76673 | 0.76223
| 4680000 | 4680000 | 2025-05-10 14:52 | 0.89096 | 0 | 0.76313 | 0.00091049 | 4680 | 0.9949 | 0.76223 | 0.76222 | 0.76222 | 0.76221 | 0.76677
| 1000 | 1000 | 2025-05-10 14:54 | 0.89872 | 0 | 0.76585 | 0.00092271 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76677 | 0.76216
| 4695000 | 4695000 | 2025-05-10 14:56 | 0.89065 | 0 | 0.76308 | 0.0009214 | 4695 | 0.99468 | 0.76216 | 0.76216 | 0.76216 | 0.76216 | 0.76677
| 1000 | 1000 | 2025-05-10 14:58 | 0.89867 | 0 | 0.76585 | 0.00092703 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76214
| 4710000 | 4710000 | 2025-05-10 15:00 | 0.89051 | 0 | 0.76306 | 0.00092205 | 4710 | 0.99443 | 0.76214 | 0.76213 | 0.76214 | 0.76213 | 0.76674
| 1000 | 1000 | 2025-05-10 15:01 | 0.89875 | 0 | 0.76583 | 0.00092978 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76211
| 4725000 | 4725000 | 2025-05-10 15:03 | 0.89013 | 0 | 0.76303 | 0.00093967 | 4725 | 0.99409 | 0.7621 | 0.7621 | 0.76209 | 0.76209 | 0.76679
| 1000 | 1000 | 2025-05-10 15:05 | 0.89873 | 0 | 0.76584 | 0.00095644 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.76202
| 4740000 | 4740000 | 2025-05-10 15:07 | 0.88979 | 0 | 0.76295 | 0.00094459 | 4740 | 0.99373 | 0.76201 | 0.76201 | 0.76201 | 0.762 | 0.76673
| 1000 | 1000 | 2025-05-10 15:09 | 0.89873 | 0 | 0.76578 | 0.00095761 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76195
| 4755000 | 4755000 | 2025-05-10 15:10 | 0.88934 | 0 | 0.76289 | 0.00095144 | 4755 | 0.99326 | 0.76194 | 0.76194 | 0.76194 | 0.76193 | 0.7667
| 1000 | 1000 | 2025-05-10 15:12 | 0.89874 | 0 | 0.76575 | 0.0009522 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76194
| 4770000 | 4770000 | 2025-05-10 15:14 | 0.8888 | 0 | 0.7629 | 0.00096926 | 4770 | 0.99278 | 0.76194 | 0.76193 | 0.76193 | 0.76193 | 0.76678
| 1000 | 1000 | 2025-05-10 15:16 | 0.89867 | 0 | 0.76581 | 0.00097087 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76192
| 4785000 | 4785000 | 2025-05-10 15:18 | 0.88819 | 0 | 0.76289 | 0.00096386 | 4785 | 0.99212 | 0.76193 | 0.76193 | 0.76193 | 0.76193 | 0.76675
| 1000 | 1000 | 2025-05-10 15:19 | 0.89864 | 0 | 0.76579 | 0.00097257 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.7619
| 4800000 | 4800000 | 2025-05-10 15:21 | 0.88764 | 0 | 0.76286 | 0.00096728 | 4800 | 0.99144 | 0.76189 | 0.7619 | 0.76189 | 0.76189 | 0.76673
| 1000 | 1000 | 2025-05-10 15:23 | 0.8987 | 0 | 0.76575 | 0.00098956 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76674 | 0.76179
| 4815000 | 4815000 | 2025-05-10 15:25 | 0.88699 | 0 | 0.76278 | 0.0009993 | 4815 | 0.99091 | 0.76178 | 0.76178 | 0.76178 | 0.76177 | 0.76678
| 1000 | 1000 | 2025-05-10 15:27 | 0.89866 | 0 | 0.76579 | 0.00099757 | #NAME? | #NAME? | 0.76679 | 0.76678 | 0.76679 | 0.76679 | 0.7618
| 4830000 | 4830000 | 2025-05-10 15:28 | 0.88624 | 0 | 0.76278 | 0.00098932 | 4830 | 0.99 | 0.7618 | 0.76179 | 0.76179 | 0.76179 | 0.76674
| 1000 | 1000 | 2025-05-10 15:30 | 0.89868 | 0 | 0.76575 | 0.00099777 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76176
| 4845000 | 4845000 | 2025-05-10 15:32 | 0.88556 | 0 | 0.76275 | 0.00098933 | 4845 | 0.98915 | 0.76176 | 0.76176 | 0.76176 | 0.76175 | 0.76671
| 1000 | 1000 | 2025-05-10 15:34 | 0.89876 | 0 | 0.76572 | 0.0010018 | #NAME? | #NAME? | 0.76672 | 0.76671 | 0.76672 | 0.76673 | 0.76171
| 4860000 | 4860000 | 2025-05-10 15:36 | 0.88462 | 0 | 0.76273 | 0.0010221 | 4860 | 0.98831 | 0.76171 | 0.76171 | 0.76171 | 0.76171 | 0.76682
| 1000 | 1000 | 2025-05-10 15:37 | 0.89868 | 0 | 0.76579 | 0.0010359 | #NAME? | #NAME? | 0.76682 | 0.76683 | 0.76683 | 0.76683 | 0.76165
| 4875000 | 4875000 | 2025-05-10 15:39 | 0.88376 | 0 | 0.76266 | 0.0010136 | 4875 | 0.98736 | 0.76165 | 0.76165 | 0.76165 | 0.76164 | 0.76672
| 1000 | 1000 | 2025-05-10 15:41 | 0.89864 | 0 | 0.7657 | 0.0010275 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76159
| 4890000 | 4890000 | 2025-05-10 15:43 | 0.88285 | 0 | 0.76261 | 0.0010311 | 4890 | 0.98631 | 0.76158 | 0.76158 | 0.76158 | 0.76157 | 0.76674
| 1000 | 1000 | 2025-05-10 15:45 | 0.89872 | 0 | 0.7657 | 0.0010525 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76149
| 4905000 | 4905000 | 2025-05-10 15:47 | 0.88196 | 0 | 0.76253 | 0.0010501 | 4905 | 0.98544 | 0.76148 | 0.76149 | 0.76148 | 0.76148 | 0.76673
| 1000 | 1000 | 2025-05-10 15:48 | 0.89868 | 0 | 0.76568 | 0.0010553 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76146
| 4920000 | 4920000 | 2025-05-10 15:50 | 0.88103 | 0 | 0.76251 | 0.0010619 | 4920 | 0.9845 | 0.76146 | 0.76145 | 0.76145 | 0.76144 | 0.76676
| 1000 | 1000 | 2025-05-10 15:52 | 0.89865 | 0 | 0.7657 | 0.0010692 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76142
| 4935000 | 4935000 | 2025-05-10 15:54 | 0.8801 | 0 | 0.76248 | 0.001067 | 4935 | 0.98338 | 0.76142 | 0.76141 | 0.76142 | 0.76141 | 0.76675
| 1000 | 1000 | 2025-05-10 15:56 | 0.89874 | 0 | 0.76569 | 0.0010784 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76676 | 0.76677 | 0.76137
| 4950000 | 4950000 | 2025-05-10 15:57 | 0.8792 | 0 | 0.76245 | 0.0010837 | 4950 | 0.98249 | 0.76137 | 0.76137 | 0.76136 | 0.76136 | 0.76678
| 1000 | 1000 | 2025-05-10 15:59 | 0.89868 | 0 | 0.76569 | 0.0010968 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.76679 | 0.7613
| 4965000 | 4965000 | 2025-05-10 16:01 | 0.87814 | 0 | 0.76238 | 0.0010856 | 4965 | 0.98125 | 0.76131 | 0.7613 | 0.7613 | 0.76129 | 0.76673
| 1000 | 1000 | 2025-05-10 16:03 | 0.89873 | 0 | 0.76563 | 0.0010999 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76124
| 4980000 | 4980000 | 2025-05-10 16:05 | 0.87713 | 0 | 0.76233 | 0.0011062 | 4980 | 0.98024 | 0.76123 | 0.76123 | 0.76123 | 0.76123 | 0.76676
| 1000 | 1000 | 2025-05-10 16:06 | 0.89869 | 0 | 0.76565 | 0.0011136 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76119
| 4995000 | 4995000 | 2025-05-10 16:08 | 0.87598 | 0 | 0.7623 | 0.0011166 | 4995 | 0.97901 | 0.76119 | 0.76119 | 0.76118 | 0.76118 | 0.76677
| 1000 | 1000 | 2025-05-10 16:10 | 0.89869 | 0 | 0.76565 | 0.0011271 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76114
| 5010000 | 5010000 | 2025-05-10 16:12 | 0.87498 | 0 | 0.76226 | 0.0011241 | 5010 | 0.97792 | 0.76114 | 0.76114 | 0.76113 | 0.76113 | 0.76675
| 1000 | 1000 | 2025-05-10 16:14 | 0.89867 | 0 | 0.76561 | 0.0011472 | #NAME? | #NAME? | 0.76676 | 0.76675 | 0.76676 | 0.76676 | 0.76102
| 5025000 | 5025000 | 2025-05-10 16:16 | 0.87395 | 0 | 0.76216 | 0.0011497 | 5025 | 0.9769 | 0.76102 | 0.76102 | 0.76101 | 0.761 | 0.76676
| 1000 | 1000 | 2025-05-10 16:17 | 0.89871 | 0 | 0.76565 | 0.001126 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.76114
| 5040000 | 5040000 | 2025-05-10 16:19 | 0.87298 | 0 | 0.76226 | 0.0011237 | 5040 | 0.97565 | 0.76113 | 0.76114 | 0.76113 | 0.76113 | 0.76675
| 1000 | 1000 | 2025-05-10 16:21 | 0.8987 | 0 | 0.76561 | 0.0011549 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76099
| 5055000 | 5055000 | 2025-05-10 16:23 | 0.87196 | 0 | 0.76214 | 0.001149 | 5055 | 0.97467 | 0.76099 | 0.76099 | 0.76098 | 0.76098 | 0.76673
| 1000 | 1000 | 2025-05-10 16:24 | 0.89866 | 0 | 0.76558 | 0.0011572 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76095
| 5070000 | 5070000 | 2025-05-10 16:26 | 0.87102 | 0 | 0.7621 | 0.0011638 | 5070 | 0.97357 | 0.76094 | 0.76094 | 0.76093 | 0.76093 | 0.76675
| 1000 | 1000 | 2025-05-10 16:28 | 0.89875 | 0 | 0.76558 | 0.0011846 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76676 | 0.76677 | 0.76084
| 5085000 | 5085000 | 2025-05-10 16:30 | 0.87011 | 0 | 0.76201 | 0.0011771 | 5085 | 0.97266 | 0.76084 | 0.76083 | 0.76083 | 0.76083 | 0.76672
| 1000 | 1000 | 2025-05-10 16:32 | 0.89873 | 0 | 0.76555 | 0.0011815 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76082
| 5100000 | 5100000 | 2025-05-10 16:34 | 0.86923 | 0 | 0.762 | 0.0011799 | 5100 | 0.9717 | 0.76082 | 0.76081 | 0.76081 | 0.76081 | 0.76672
| 1000 | 1000 | 2025-05-10 16:35 | 0.89869 | 0 | 0.76552 | 0.0012035 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76071
| 5115000 | 5115000 | 2025-05-10 16:37 | 0.86831 | 0 | 0.7619 | 0.0012056 | 5115 | 0.97076 | 0.76071 | 0.7607 | 0.7607 | 0.76069 | 0.76673
| 1000 | 1000 | 2025-05-10 16:39 | 0.89873 | 0 | 0.76553 | 0.0011985 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76074
| 5130000 | 5130000 | 2025-05-10 16:41 | 0.86755 | 0 | 0.76194 | 0.0012099 | 5130 | 0.96998 | 0.76074 | 0.76073 | 0.76073 | 0.76072 | 0.76678
| 1000 | 1000 | 2025-05-10 16:43 | 0.89865 | 0 | 0.76556 | 0.0012244 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76679 | 0.76066
| 5145000 | 5145000 | 2025-05-10 16:45 | 0.86674 | 0 | 0.76186 | 0.0012214 | 5145 | 0.96919 | 0.76065 | 0.76064 | 0.76064 | 0.76064 | 0.76675
| 1000 | 1000 | 2025-05-10 16:46 | 0.89858 | 0 | 0.76551 | 0.001237 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76056
| 5160000 | 5160000 | 2025-05-10 16:48 | 0.86608 | 0 | 0.76181 | 0.0012441 | 5160 | 0.96849 | 0.76056 | 0.76056 | 0.76056 | 0.76056 | 0.76678
| 1000 | 1000 | 2025-05-10 16:50 | 0.89866 | 0 | 0.76555 | 0.001238 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.7606
| 5175000 | 5175000 | 2025-05-10 16:52 | 0.86535 | 0 | 0.76183 | 0.0012372 | 5175 | 0.96763 | 0.7606 | 0.7606 | 0.76059 | 0.76059 | 0.76678
| 1000 | 1000 | 2025-05-10 16:54 | 0.89864 | 0 | 0.76553 | 0.0012609 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.76679 | 0.76049
| 5190000 | 5190000 | 2025-05-10 16:55 | 0.86482 | 0 | 0.76174 | 0.0012523 | 5190 | 0.96705 | 0.76049 | 0.76049 | 0.76049 | 0.76048 | 0.76675
| 1000 | 1000 | 2025-05-10 16:57 | 0.89871 | 0 | 0.7655 | 0.0012569 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76047
| 5205000 | 5205000 | 2025-05-10 16:59 | 0.86423 | 0 | 0.76173 | 0.0012644 | 5205 | 0.96645 | 0.76047 | 0.76047 | 0.76046 | 0.76046 | 0.76679
| 1000 | 1000 | 2025-05-10 17:01 | 0.89868 | 0 | 0.76551 | 0.0012728 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.76679 | 0.76042
| 5220000 | 5220000 | 2025-05-10 17:03 | 0.86379 | 0 | 0.76168 | 0.0012709 | 5220 | 0.9659 | 0.76041 | 0.76041 | 0.76041 | 0.76041 | 0.76677
| 1000 | 1000 | 2025-05-10 17:05 | 0.89873 | 0 | 0.76547 | 0.001296 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76029
| 5235000 | 5235000 | 2025-05-10 17:07 | 0.86335 | 0 | 0.76158 | 0.0012999 | 5235 | 0.96562 | 0.76028 | 0.76028 | 0.76028 | 0.76027 | 0.76678
| 1000 | 1000 | 2025-05-10 17:08 | 0.89867 | 0 | 0.76548 | 0.0013104 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.7668 | 0.76679 | 0.76024
| 5250000 | 5250000 | 2025-05-10 17:10 | 0.86302 | 0 | 0.76154 | 0.0013134 | 5250 | 0.9653 | 0.76023 | 0.76023 | 0.76023 | 0.76023 | 0.7668
| 1000 | 1000 | 2025-05-10 17:12 | 0.89868 | 0 | 0.76549 | 0.0013206 | #NAME? | #NAME? | 0.7668 | 0.76681 | 0.76681 | 0.76681 | 0.7602
| 5265000 | 5265000 | 2025-05-10 17:14 | 0.86275 | 0 | 0.76151 | 0.0013064 | 5265 | 0.96498 | 0.7602 | 0.7602 | 0.7602 | 0.7602 | 0.76673
| 1000 | 1000 | 2025-05-10 17:16 | 0.89865 | 0 | 0.76541 | 0.0013294 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76009
| 5280000 | 5280000 | 2025-05-10 17:18 | 0.86259 | 0 | 0.76142 | 0.0013391 | 5280 | 0.96493 | 0.76009 | 0.76009 | 0.76008 | 0.76008 | 0.76678
| 1000 | 1000 | 2025-05-10 17:20 | 0.89866 | 0 | 0.76544 | 0.0013496 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.7668 | 0.76004
| 5295000 | 5295000 | 2025-05-10 17:22 | 0.86251 | 0 | 0.76138 | 0.0013438 | 5295 | 0.96487 | 0.76004 | 0.76004 | 0.76003 | 0.76003 | 0.76675
| 1000 | 1000 | 2025-05-10 17:24 | 0.89867 | 0 | 0.76542 | 0.0013311 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.7601
| 5310000 | 5310000 | 2025-05-10 17:26 | 0.86245 | 0 | 0.76143 | 0.0013416 | 5310 | 0.96478 | 0.76009 | 0.7601 | 0.76009 | 0.76009 | 0.7668
| 1000 | 1000 | 2025-05-10 17:27 | 0.89864 | 0 | 0.76544 | 0.0013682 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.76681 | 0.76681 | 0.75996
| 5325000 | 5325000 | 2025-05-10 17:29 | 0.86254 | 0 | 0.76132 | 0.001375 | 5325 | 0.96496 | 0.75995 | 0.75995 | 0.75995 | 0.75994 | 0.76682
| 1000 | 1000 | 2025-05-10 17:31 | 0.8987 | 0 | 0.76544 | 0.0013897 | #NAME? | #NAME? | 0.76683 | 0.76683 | 0.76683 | 0.76684 | 0.75988
| 5340000 | 5340000 | 2025-05-10 17:33 | 0.86264 | 0 | 0.76126 | 0.0013806 | 5340 | 0.96513 | 0.75988 | 0.75988 | 0.75988 | 0.75987 | 0.76678
| 1000 | 1000 | 2025-05-10 17:35 | 0.89867 | 0 | 0.7654 | 0.0013917 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.75984
| 5355000 | 5355000 | 2025-05-10 17:37 | 0.86293 | 0 | 0.76121 | 0.0013842 | 5355 | 0.96543 | 0.75984 | 0.75983 | 0.75983 | 0.75983 | 0.76675
| 1000 | 1000 | 2025-05-10 17:39 | 0.89871 | 0 | 0.76536 | 0.001391 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.7598
| 5370000 | 5370000 | 2025-05-10 17:41 | 0.8632 | 0 | 0.76119 | 0.0013981 | 5370 | 0.96571 | 0.7598 | 0.75979 | 0.75979 | 0.75979 | 0.76678
| 1000 | 1000 | 2025-05-10 17:43 | 0.89875 | 0 | 0.76537 | 0.0014216 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.75968
| 5385000 | 5385000 | 2025-05-10 17:45 | 0.86355 | 0 | 0.76109 | 0.0014152 | 5385 | 0.96622 | 0.75968 | 0.75967 | 0.75967 | 0.75967 | 0.76675
| 1000 | 1000 | 2025-05-10 17:46 | 0.89872 | 0 | 0.76534 | 0.0014223 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.75965
| 5400000 | 5400000 | 2025-05-10 17:48 | 0.86398 | 0 | 0.76109 | 0.0014358 | 5400 | 0.96684 | 0.75965 | 0.75965 | 0.75965 | 0.75965 | 0.76683
| 1000 | 1000 | 2025-05-10 17:50 | 0.89867 | 0 | 0.76539 | 0.001449 | #NAME? | #NAME? | 0.76683 | 0.76684 | 0.76684 | 0.76685 | 0.7596
| 5415000 | 5415000 | 2025-05-10 17:52 | 0.86439 | 0 | 0.76102 | 0.0014326 | 5415 | 0.96725 | 0.75959 | 0.75959 | 0.75959 | 0.75958 | 0.76675
| 1000 | 1000 | 2025-05-10 17:54 | 0.89869 | 0 | 0.76531 | 0.0014511 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.7595
| 5430000 | 5430000 | 2025-05-10 17:56 | 0.86509 | 0 | 0.76094 | 0.0014464 | 5430 | 0.96811 | 0.7595 | 0.7595 | 0.7595 | 0.75949 | 0.76673
| 1000 | 1000 | 2025-05-10 17:58 | 0.89867 | 0 | 0.76527 | 0.0014615 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76673 | 0.76673 | 0.75942
| 5445000 | 5445000 | 2025-05-10 18:00 | 0.86563 | 0 | 0.76089 | 0.0014767 | 5445 | 0.96883 | 0.75942 | 0.75941 | 0.75941 | 0.75941 | 0.7668
| 1000 | 1000 | 2025-05-10 18:02 | 0.89867 | 0 | 0.76531 | 0.0014914 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.75934
| 5460000 | 5460000 | 2025-05-10 18:04 | 0.86637 | 0 | 0.76082 | 0.0014939 | 5460 | 0.96972 | 0.75934 | 0.75933 | 0.75933 | 0.75932 | 0.7668
| 1000 | 1000 | 2025-05-10 18:06 | 0.89866 | 0 | 0.76529 | 0.0015099 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.75925
| 5475000 | 5475000 | 2025-05-10 18:08 | 0.86713 | 0 | 0.76074 | 0.0015001 | 5475 | 0.9706 | 0.75925 | 0.75923 | 0.75923 | 0.75923 | 0.76674
| 1000 | 1000 | 2025-05-10 18:09 | 0.89867 | 0 | 0.76523 | 0.0015146 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.75917
| 5490000 | 5490000 | 2025-05-10 18:12 | 0.86799 | 0 | 0.76069 | 0.0015244 | 5490 | 0.97157 | 0.75917 | 0.75916 | 0.75916 | 0.75915 | 0.76678
| 1000 | 1000 | 2025-05-10 18:13 | 0.89875 | 0 | 0.76525 | 0.0015415 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.75909
| 5505000 | 5505000 | 2025-05-10 18:15 | 0.86877 | 0 | 0.76061 | 0.0015251 | 5505 | 0.97249 | 0.75909 | 0.75909 | 0.75909 | 0.75908 | 0.76671
| 1000 | 1000 | 2025-05-10 18:17 | 0.89872 | 0 | 0.76519 | 0.0015381 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.75903
| 5520000 | 5520000 | 2025-05-10 18:19 | 0.86965 | 0 | 0.76057 | 0.0015444 | 5520 | 0.97354 | 0.75903 | 0.75903 | 0.75903 | 0.75903 | 0.76675
| 1000 | 1000 | 2025-05-10 18:21 | 0.8987 | 0 | 0.76518 | 0.0015699 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76675 | 0.7589
| 5535000 | 5535000 | 2025-05-10 18:23 | 0.87049 | 0 | 0.76047 | 0.0015679 | 5535 | 0.97468 | 0.7589 | 0.7589 | 0.7589 | 0.7589 | 0.76674
| 1000 | 1000 | 2025-05-10 18:25 | 0.89864 | 0 | 0.76518 | 0.0015613 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76674 | 0.75894
| 5550000 | 5550000 | 2025-05-10 18:27 | 0.87139 | 0 | 0.76049 | 0.0015679 | 5550 | 0.97564 | 0.75893 | 0.75893 | 0.75892 | 0.75892 | 0.76676
| 1000 | 1000 | 2025-05-10 18:29 | 0.89866 | 0 | 0.76518 | 0.001581 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.75886
| 5565000 | 5565000 | 2025-05-10 18:31 | 0.87234 | 0 | 0.76044 | 0.001584 | 5565 | 0.97668 | 0.75886 | 0.75885 | 0.75885 | 0.75885 | 0.76677
| 1000 | 1000 | 2025-05-10 18:32 | 0.89874 | 0 | 0.76519 | 0.0015989 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.75879
| 5580000 | 5580000 | 2025-05-10 18:35 | 0.87334 | 0 | 0.76039 | 0.001601 | 5580 | 0.9779 | 0.7588 | 0.75879 | 0.75879 | 0.75878 | 0.76679
| 1000 | 1000 | 2025-05-10 18:36 | 0.8987 | 0 | 0.76518 | 0.0016184 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.7587
| 5595000 | 5595000 | 2025-05-10 18:38 | 0.87428 | 0 | 0.7603 | 0.0016108 | 5595 | 0.979 | 0.7587 | 0.75869 | 0.75869 | 0.75869 | 0.76675
| 1000 | 1000 | 2025-05-10 18:40 | 0.89869 | 0 | 0.76512 | 0.0016345 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.75858
| 5610000 | 5610000 | 2025-05-10 18:42 | 0.87512 | 0 | 0.76021 | 0.0016317 | 5610 | 0.98003 | 0.75858 | 0.75858 | 0.75858 | 0.75858 | 0.76674
| 1000 | 1000 | 2025-05-10 18:44 | 0.89869 | 0 | 0.7651 | 0.0016513 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76676 | 0.75849
| 5625000 | 5625000 | 2025-05-10 18:46 | 0.87598 | 0 | 0.76014 | 0.0016449 | 5625 | 0.98109 | 0.7585 | 0.7585 | 0.75849 | 0.75849 | 0.76672
| 1000 | 1000 | 2025-05-10 18:48 | 0.89867 | 0 | 0.76508 | 0.0016573 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.75845
| 5640000 | 5640000 | 2025-05-10 18:50 | 0.87681 | 0 | 0.7601 | 0.0016571 | 5640 | 0.98208 | 0.75845 | 0.75845 | 0.75844 | 0.75844 | 0.76673
| 1000 | 1000 | 2025-05-10 18:52 | 0.89864 | 0 | 0.76507 | 0.0016729 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.75838
| 5655000 | 5655000 | 2025-05-10 18:54 | 0.87762 | 0 | 0.76005 | 0.00168 | 5655 | 0.9831 | 0.75837 | 0.75837 | 0.75836 | 0.75836 | 0.76677
| 1000 | 1000 | 2025-05-10 18:56 | 0.89863 | 0 | 0.76509 | 0.0016883 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76679 | 0.76679 | 0.75834
| 5670000 | 5670000 | 2025-05-10 18:58 | 0.87836 | 0 | 0.76002 | 0.0016823 | 5670 | 0.98389 | 0.75834 | 0.75834 | 0.75833 | 0.75833 | 0.76675
| 1000 | 1000 | 2025-05-10 18:59 | 0.89867 | 0 | 0.76507 | 0.0016734 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.75838
| 5685000 | 5685000 | 2025-05-10 19:01 | 0.87897 | 0 | 0.76006 | 0.0016941 | 5685 | 0.98462 | 0.75837 | 0.75837 | 0.75837 | 0.75837 | 0.76684
| 1000 | 1000 | 2025-05-10 19:03 | 0.89863 | 0 | 0.76512 | 0.0017241 | #NAME? | #NAME? | 0.76685 | 0.76684 | 0.76684 | 0.76684 | 0.75822
| 5700000 | 5700000 | 2025-05-10 19:05 | 0.87952 | 0 | 0.75993 | 0.0017118 | 5700 | 0.98532 | 0.75822 | 0.75822 | 0.75822 | 0.75821 | 0.76678
| 1000 | 1000 | 2025-05-10 19:07 | 0.89863 | 0 | 0.76504 | 0.0017412 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.75807
| 5715000 | 5715000 | 2025-05-10 19:09 | 0.88003 | 0 | 0.75979 | 0.0017259 | 5715 | 0.98596 | 0.75807 | 0.75807 | 0.75807 | 0.75806 | 0.7667
| 1000 | 1000 | 2025-05-10 19:11 | 0.89866 | 0 | 0.76499 | 0.0017287 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.75807
| 5730000 | 5730000 | 2025-05-10 19:13 | 0.88042 | 0 | 0.7598 | 0.0017361 | 5730 | 0.98641 | 0.75807 | 0.75807 | 0.75806 | 0.75807 | 0.76675
| 1000 | 1000 | 2025-05-10 19:15 | 0.89864 | 0 | 0.76499 | 0.001762 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.75794
| 5745000 | 5745000 | 2025-05-10 19:17 | 0.88069 | 0 | 0.7597 | 0.0017597 | 5745 | 0.98683 | 0.75794 | 0.75794 | 0.75794 | 0.75793 | 0.76673
| 1000 | 1000 | 2025-05-10 19:19 | 0.89865 | 0 | 0.76498 | 0.0017639 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.75792
| 5760000 | 5760000 | 2025-05-10 19:21 | 0.88083 | 0 | 0.75968 | 0.0017691 | 5760 | 0.98691 | 0.75791 | 0.75791 | 0.75791 | 0.75791 | 0.76675
| 1000 | 1000 | 2025-05-10 19:23 | 0.89874 | 0 | 0.76498 | 0.0017866 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76678 | 0.75783
| 5775000 | 5775000 | 2025-05-10 19:25 | 0.88078 | 0 | 0.75961 | 0.0017841 | 5775 | 0.98697 | 0.75783 | 0.75783 | 0.75783 | 0.75782 | 0.76675
| 1000 | 1000 | 2025-05-10 19:26 | 0.8987 | 0 | 0.76497 | 0.0017863 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76676 | 0.75783
| 5790000 | 5790000 | 2025-05-10 19:29 | 0.88065 | 0 | 0.75961 | 0.0017886 | 5790 | 0.98686 | 0.75782 | 0.75782 | 0.75782 | 0.75782 | 0.76676
| 1000 | 1000 | 2025-05-10 19:30 | 0.89867 | 0 | 0.76496 | 0.0018158 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.7577
| 5805000 | 5805000 | 2025-05-10 19:32 | 0.88031 | 0 | 0.7595 | 0.0018048 | 5805 | 0.98656 | 0.7577 | 0.7577 | 0.75769 | 0.75769 | 0.76672
| 1000 | 1000 | 2025-05-10 19:34 | 0.89867 | 0 | 0.76492 | 0.0018057 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.7577
| 5820000 | 5820000 | 2025-05-10 19:36 | 0.87982 | 0 | 0.75951 | 0.0018137 | 5820 | 0.98595 | 0.7577 | 0.75769 | 0.75769 | 0.75769 | 0.76676
| 1000 | 1000 | 2025-05-10 19:38 | 0.89872 | 0 | 0.76492 | 0.0018426 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.75755
| 5835000 | 5835000 | 2025-05-10 19:40 | 0.87907 | 0 | 0.75937 | 0.0018233 | 5835 | 0.98529 | 0.75755 | 0.75755 | 0.75754 | 0.75754 | 0.76666
| 1000 | 1000 | 2025-05-10 19:42 | 0.89864 | 0 | 0.76485 | 0.0018236 | #NAME? | #NAME? | 0.76666 | 0.76667 | 0.76667 | 0.76669 | 0.75755
| 5850000 | 5850000 | 2025-05-10 19:44 | 0.87822 | 0 | 0.75938 | 0.0018341 | 5850 | 0.98432 | 0.75755 | 0.75755 | 0.75755 | 0.75754 | 0.76672
| 1000 | 1000 | 2025-05-10 19:46 | 0.89866 | 0 | 0.76487 | 0.0018549 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.75745
| 5865000 | 5865000 | 2025-05-10 19:48 | 0.87707 | 0 | 0.75929 | 0.0018565 | 5865 | 0.98313 | 0.75745 | 0.75744 | 0.75743 | 0.75743 | 0.76672
| 1000 | 1000 | 2025-05-10 19:50 | 0.89867 | 0 | 0.76487 | 0.001857 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76673 | 0.75744
| 5880000 | 5880000 | 2025-05-10 19:52 | 0.87579 | 0 | 0.75928 | 0.0018546 | 5880 | 0.98165 | 0.75743 | 0.75743 | 0.75742 | 0.75742 | 0.7667
| 1000 | 1000 | 2025-05-10 19:54 | 0.89872 | 0 | 0.76486 | 0.0018458 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76672 | 0.75748
| 5895000 | 5895000 | 2025-05-10 19:56 | 0.87417 | 0 | 0.75932 | 0.0018538 | 5895 | 0.97987 | 0.75748 | 0.75747 | 0.75747 | 0.75746 | 0.76674
| 1000 | 1000 | 2025-05-10 19:57 | 0.89865 | 0 | 0.76488 | 0.0018817 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76677 | 0.76677 | 0.75735
| 5910000 | 5910000 | 2025-05-10 20:00 | 0.87234 | 0 | 0.75922 | 0.0018785 | 5910 | 0.97791 | 0.75735 | 0.75734 | 0.75734 | 0.75734 | 0.76674
| 1000 | 1000 | 2025-05-10 20:01 | 0.89868 | 0 | 0.76486 | 0.0018854 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.75732
| 5925000 | 5925000 | 2025-05-10 20:03 | 0.87018 | 0 | 0.7592 | 0.0018913 | 5925 | 0.97555 | 0.75732 | 0.75731 | 0.75731 | 0.75731 | 0.76677
| 1000 | 1000 | 2025-05-10 20:05 | 0.89867 | 0 | 0.76488 | 0.0018877 | #NAME? | #NAME? | 0.76678 | 0.76677 | 0.76677 | 0.76677 | 0.75733
| 5940000 | 5940000 | 2025-05-10 20:07 | 0.86783 | 0 | 0.75921 | 0.0018861 | 5940 | 0.97285 | 0.75733 | 0.75732 | 0.75732 | 0.75732 | 0.76675
| 1000 | 1000 | 2025-05-10 20:09 | 0.89871 | 0 | 0.76487 | 0.0018866 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.75733
| 5955000 | 5955000 | 2025-05-10 20:11 | 0.86519 | 0 | 0.75922 | 0.001901 | 5955 | 0.96999 | 0.75732 | 0.75732 | 0.75732 | 0.75732 | 0.76683
| 1000 | 1000 | 2025-05-10 20:13 | 0.89866 | 0 | 0.76492 | 0.0019151 | #NAME? | #NAME? | 0.76683 | 0.76683 | 0.76684 | 0.76684 | 0.75726
| 5970000 | 5970000 | 2025-05-10 20:15 | 0.86223 | 0 | 0.75915 | 0.0019023 | 5970 | 0.96666 | 0.75726 | 0.75725 | 0.75725 | 0.75724 | 0.76676
| 1000 | 1000 | 2025-05-10 20:17 | 0.89867 | 0 | 0.76485 | 0.0019219 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.75717
| 5985000 | 5985000 | 2025-05-10 20:19 | 0.85895 | 0 | 0.75908 | 0.0019236 | 5985 | 0.96303 | 0.75716 | 0.75715 | 0.75715 | 0.75715 | 0.76677
| 1000 | 1000 | 2025-05-10 20:21 | 0.89869 | 0 | 0.76484 | 0.0019335 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.7571
| 6000000 | 6000000 | 2025-05-10 20:23 | 0.85547 | 0 | 0.75903 | 0.001929 | 6000 | 0.95916 | 0.7571 | 0.7571 | 0.7571 | 0.75709 | 0.76674
| 1000 | 1000 | 2025-05-10 20:24 | 0.89869 | 0 | 0.76481 | 0.0019374 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.75706
| 6015000 | 6015000 | 2025-05-10 20:27 | 0.85159 | 0 | 0.759 | 0.0019404 | 6015 | 0.95486 | 0.75706 | 0.75706 | 0.75706 | 0.75705 | 0.76676
| 1000 | 1000 | 2025-05-10 20:28 | 0.8987 | 0 | 0.76483 | 0.0019357 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.75708
| 6030000 | 6030000 | 2025-05-10 20:31 | 0.84747 | 0 | 0.759 | 0.0019229 | 6030 | 0.9501 | 0.75708 | 0.75708 | 0.75707 | 0.75707 | 0.76669
| 1000 | 1000 | 2025-05-10 20:32 | 0.89875 | 0 | 0.76476 | 0.001939 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.75701
| 6045000 | 6045000 | 2025-05-10 20:34 | 0.84288 | 0 | 0.75896 | 0.0019532 | 6045 | 0.94514 | 0.757 | 0.757 | 0.757 | 0.757 | 0.76677
| 1000 | 1000 | 2025-05-10 20:36 | 0.89868 | 0 | 0.76481 | 0.0019623 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.75696
| 6060000 | 6060000 | 2025-05-10 20:38 | 0.83797 | 0 | 0.75891 | 0.0019603 | 6060 | 0.93966 | 0.75695 | 0.75695 | 0.75694 | 0.75694 | 0.76675
| 1000 | 1000 | 2025-05-10 20:40 | 0.8987 | 0 | 0.76479 | 0.0019702 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.75691
| 6075000 | 6075000 | 2025-05-10 20:42 | 0.83274 | 0 | 0.75886 | 0.0019585 | 6075 | 0.93382 | 0.75691 | 0.75691 | 0.7569 | 0.7569 | 0.7667
| 1000 | 1000 | 2025-05-10 20:44 | 0.89867 | 0 | 0.76475 | 0.0019583 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.75691
| 6090000 | 6090000 | 2025-05-10 20:46 | 0.82719 | 0 | 0.75889 | 0.0019818 | 6090 | 0.9276 | 0.75691 | 0.75691 | 0.75691 | 0.7569 | 0.76682
| 1000 | 1000 | 2025-05-10 20:48 | 0.89873 | 0 | 0.76482 | 0.0020017 | #NAME? | #NAME? | 0.76682 | 0.76682 | 0.76683 | 0.76682 | 0.75681
| 6105000 | 6105000 | 2025-05-10 20:50 | 0.8212 | 0 | 0.7588 | 0.001989 | 6105 | 0.92101 | 0.75681 | 0.75681 | 0.75681 | 0.75681 | 0.76676
| 1000 | 1000 | 2025-05-10 20:52 | 0.89863 | 0 | 0.76476 | 0.0019995 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76677 | 0.75677
| 6120000 | 6120000 | 2025-05-10 20:54 | 0.815 | 0 | 0.75876 | 0.0020052 | 6120 | 0.91409 | 0.75676 | 0.75676 | 0.75676 | 0.75676 | 0.76679
| 1000 | 1000 | 2025-05-10 20:55 | 0.89869 | 0 | 0.76479 | 0.0020039 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.75678
| 6135000 | 6135000 | 2025-05-10 20:58 | 0.80839 | 0 | 0.75877 | 0.002002 | 6135 | 0.90666 | 0.75677 | 0.75677 | 0.75676 | 0.75676 | 0.76677
| 1000 | 1000 | 2025-05-10 20:59 | 0.89866 | 0 | 0.76476 | 0.002019 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.75668
| 6150000 | 6150000 | 2025-05-10 21:01 | 0.80138 | 0 | 0.75871 | 0.0020272 | 6150 | 0.89889 | 0.75668 | 0.75668 | 0.75668 | 0.75668 | 0.76682
| 1000 | 1000 | 2025-05-10 21:03 | 0.89867 | 0 | 0.76479 | 0.0020287 | #NAME? | #NAME? | 0.76682 | 0.76682 | 0.76682 | 0.76682 | 0.75668
| 6165000 | 6165000 | 2025-05-10 21:05 | 0.79396 | 0 | 0.75869 | 0.0020164 | 6165 | 0.89056 | 0.75668 | 0.75667 | 0.75667 | 0.75667 | 0.76675
| 1000 | 1000 | 2025-05-10 21:07 | 0.89864 | 0 | 0.76474 | 0.0020251 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.75663
| 6180000 | 6180000 | 2025-05-10 21:09 | 0.78632 | 0 | 0.75864 | 0.0020221 | 6180 | 0.88198 | 0.75663 | 0.75663 | 0.75662 | 0.75662 | 0.76673
| 1000 | 1000 | 2025-05-10 21:11 | 0.89866 | 0 | 0.7647 | 0.0020346 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.75657
| 6195000 | 6195000 | 2025-05-10 21:13 | 0.77826 | 0 | 0.75861 | 0.0020457 | 6195 | 0.87299 | 0.75656 | 0.75656 | 0.75656 | 0.75656 | 0.76679
| 1000 | 1000 | 2025-05-10 21:15 | 0.89871 | 0 | 0.76475 | 0.0020517 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.76681 | 0.76681 | 0.75654
| 6210000 | 6210000 | 2025-05-10 21:17 | 0.76988 | 0 | 0.75859 | 0.002053 | 6210 | 0.86359 | 0.75654 | 0.75654 | 0.75653 | 0.75653 | 0.7668
| 1000 | 1000 | 2025-05-10 21:19 | 0.89873 | 0 | 0.76475 | 0.0020564 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.75652
| 6225000 | 6225000 | 2025-05-10 21:21 | 0.76106 | 0 | 0.75856 | 0.0020514 | 6225 | 0.85369 | 0.75652 | 0.75651 | 0.75651 | 0.7565 | 0.76677
| 1000 | 1000 | 2025-05-10 21:23 | 0.8987 | 0 | 0.76469 | 0.0020801 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.75637
| 6240000 | 6240000 | 2025-05-10 21:25 | 0.75193 | 0 | 0.75844 | 0.0020808 | 6240 | 0.84359 | 0.75637 | 0.75637 | 0.75636 | 0.75636 | 0.76677
| 1000 | 1000 | 2025-05-10 21:27 | 0.89868 | 0 | 0.76468 | 0.0020947 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76677 | 0.7563
| 6255000 | 6255000 | 2025-05-10 21:29 | 0.74237 | 0 | 0.7584 | 0.0020975 | 6255 | 0.83298 | 0.7563 | 0.7563 | 0.7563 | 0.7563 | 0.76679
| 1000 | 1000 | 2025-05-10 21:31 | 0.89861 | 0 | 0.76468 | 0.0021129 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.75623
| 6270000 | 6270000 | 2025-05-10 21:33 | 0.73256 | 0 | 0.75833 | 0.0021092 | 6270 | 0.82196 | 0.75623 | 0.75622 | 0.75622 | 0.75622 | 0.76677
| 1000 | 1000 | 2025-05-10 21:35 | 0.89867 | 0 | 0.76467 | 0.0021227 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.7668 | 0.75617
| 6285000 | 6285000 | 2025-05-10 21:37 | 0.72239 | 0 | 0.75828 | 0.0021153 | 6285 | 0.81057 | 0.75617 | 0.75617 | 0.75616 | 0.75616 | 0.76674
| 1000 | 1000 | 2025-05-10 21:38 | 0.89865 | 0 | 0.76462 | 0.0021265 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76675 | 0.75611
| 6300000 | 6300000 | 2025-05-10 21:41 | 0.71187 | 0 | 0.75823 | 0.0021292 | 6300 | 0.79881 | 0.7561 | 0.7561 | 0.7561 | 0.75609 | 0.76674
| 1000 | 1000 | 2025-05-10 21:42 | 0.89867 | 0 | 0.76461 | 0.0021408 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.75605
| 6315000 | 6315000 | 2025-05-10 21:45 | 0.701 | 0 | 0.75818 | 0.0021461 | 6315 | 0.78669 | 0.75604 | 0.75604 | 0.75604 | 0.75603 | 0.76677
| 1000 | 1000 | 2025-05-10 21:46 | 0.89865 | 0 | 0.76463 | 0.0021491 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.75603
| 6330000 | 6330000 | 2025-05-10 21:49 | 0.68985 | 0 | 0.75817 | 0.0021515 | 6330 | 0.77413 | 0.75602 | 0.75602 | 0.75602 | 0.75601 | 0.76677
| 1000 | 1000 | 2025-05-10 21:50 | 0.89871 | 0 | 0.76461 | 0.0021771 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.76679 | 0.7559
| 6345000 | 6345000 | 2025-05-10 21:53 | 0.67843 | 0 | 0.75806 | 0.0021676 | 6345 | 0.76138 | 0.7559 | 0.7559 | 0.75589 | 0.75589 | 0.76673
| 1000 | 1000 | 2025-05-10 21:54 | 0.89869 | 0 | 0.76456 | 0.0021854 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.75582
| 6360000 | 6360000 | 2025-05-10 21:56 | 0.66661 | 0 | 0.758 | 0.0021833 | 6360 | 0.74815 | 0.75582 | 0.75582 | 0.75582 | 0.75581 | 0.76673
| 1000 | 1000 | 2025-05-10 21:58 | 0.89871 | 0 | 0.76455 | 0.0021982 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.75576
| 6375000 | 6375000 | 2025-05-10 22:00 | 0.65454 | 0 | 0.75795 | 0.0021991 | 6375 | 0.7347 | 0.75575 | 0.75575 | 0.75575 | 0.75574 | 0.76674
| 1000 | 1000 | 2025-05-10 22:02 | 0.89865 | 0 | 0.76454 | 0.0022227 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76677 | 0.75565
| 6390000 | 6390000 | 2025-05-10 22:04 | 0.64217 | 0 | 0.75786 | 0.0022194 | 6390 | 0.72088 | 0.75564 | 0.75564 | 0.75564 | 0.75564 | 0.76674
| 1000 | 1000 | 2025-05-10 22:06 | 0.89863 | 0 | 0.76451 | 0.0022355 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.75557
| 6405000 | 6405000 | 2025-05-10 22:08 | 0.62961 | 0 | 0.75779 | 0.002232 | 6405 | 0.70679 | 0.75557 | 0.75556 | 0.75556 | 0.75555 | 0.76672
| 1000 | 1000 | 2025-05-10 22:10 | 0.89868 | 0 | 0.76449 | 0.002238 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.75554
| 6420000 | 6420000 | 2025-05-10 22:12 | 0.6167 | 0 | 0.75778 | 0.0022416 | 6420 | 0.69232 | 0.75554 | 0.75554 | 0.75554 | 0.75553 | 0.76675
| 1000 | 1000 | 2025-05-10 22:14 | 0.89867 | 0 | 0.7645 | 0.0022562 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.75547
| 6435000 | 6435000 | 2025-05-10 22:17 | 0.60356 | 0 | 0.75772 | 0.0022594 | 6435 | 0.67758 | 0.75547 | 0.75546 | 0.75546 | 0.75546 | 0.76676
| 1000 | 1000 | 2025-05-10 22:18 | 0.89872 | 0 | 0.7645 | 0.0022771 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.75539
| 6450000 | 6450000 | 2025-05-10 22:21 | 0.59023 | 0 | 0.75766 | 0.0022784 | 6450 | 0.66264 | 0.75538 | 0.75538 | 0.75538 | 0.75538 | 0.76677
| 1000 | 1000 | 2025-05-10 22:22 | 0.89873 | 0 | 0.76447 | 0.0023068 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.75525
| 6465000 | 6465000 | 2025-05-10 22:25 | 0.57669 | 0 | 0.75754 | 0.0022946 | 6465 | 0.64752 | 0.75525 | 0.75524 | 0.75524 | 0.75524 | 0.76672
| 1000 | 1000 | 2025-05-10 22:26 | 0.89874 | 0 | 0.76444 | 0.002297 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.75525
| 6480000 | 6480000 | 2025-05-10 22:29 | 0.56291 | 0 | 0.75754 | 0.0022949 | 6480 | 0.63203 | 0.75525 | 0.75524 | 0.75524 | 0.75523 | 0.76672
| 1000 | 1000 | 2025-05-10 22:30 | 0.89869 | 0 | 0.7644 | 0.0023234 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.7551
| 6495000 | 6495000 | 2025-05-10 22:33 | 0.54892 | 0 | 0.75743 | 0.0023296 | 6495 | 0.61645 | 0.7551 | 0.7551 | 0.7551 | 0.75509 | 0.76675
| 1000 | 1000 | 2025-05-10 22:34 | 0.89866 | 0 | 0.76442 | 0.0023413 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.75505
| 6510000 | 6510000 | 2025-05-10 22:37 | 0.53475 | 0 | 0.75739 | 0.0023462 | 6510 | 0.60058 | 0.75505 | 0.75504 | 0.75504 | 0.75504 | 0.76677
| 1000 | 1000 | 2025-05-10 22:38 | 0.89864 | 0 | 0.76441 | 0.0023705 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.75493
| 6525000 | 6525000 | 2025-05-10 22:41 | 0.52049 | 0 | 0.75729 | 0.0023679 | 6525 | 0.58458 | 0.75493 | 0.75492 | 0.75492 | 0.75491 | 0.76676
| 1000 | 1000 | 2025-05-10 22:42 | 0.8987 | 0 | 0.76438 | 0.0023878 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.75483
| 6540000 | 6540000 | 2025-05-10 22:45 | 0.50607 | 0 | 0.7572 | 0.0023783 | 6540 | 0.5684 | 0.75483 | 0.75482 | 0.75482 | 0.75481 | 0.76671
| 1000 | 1000 | 2025-05-10 22:46 | 0.8987 | 0 | 0.76431 | 0.0024064 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.75469
| 6555000 | 6555000 | 2025-05-10 22:49 | 0.49145 | 0 | 0.75708 | 0.0024134 | 6555 | 0.55207 | 0.75468 | 0.75467 | 0.75467 | 0.75467 | 0.76674
| 1000 | 1000 | 2025-05-10 22:50 | 0.8987 | 0 | 0.76432 | 0.0024166 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.75466
| 6570000 | 6570000 | 2025-05-10 22:53 | 0.47677 | 0 | 0.75708 | 0.0024243 | 6570 | 0.53556 | 0.75465 | 0.75465 | 0.75465 | 0.75465 | 0.76677
| 1000 | 1000 | 2025-05-10 22:54 | 0.89874 | 0 | 0.76432 | 0.0024564 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.75449
| 6585000 | 6585000 | 2025-05-10 22:57 | 0.46203 | 0 | 0.75692 | 0.0024398 | 6585 | 0.51903 | 0.75448 | 0.75448 | 0.75448 | 0.75447 | 0.76668
| 1000 | 1000 | 2025-05-10 22:59 | 0.89876 | 0 | 0.76423 | 0.0024612 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.75438
| 6600000 | 6600000 | 2025-05-10 23:01 | 0.4472 | 0 | 0.75685 | 0.0024705 | 6600 | 0.50242 | 0.75437 | 0.75438 | 0.75438 | 0.75438 | 0.76673
| 1000 | 1000 | 2025-05-10 23:03 | 0.8988 | 0 | 0.76425 | 0.0024831 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76673 | 0.76674 | 0.75432
| 6615000 | 6615000 | 2025-05-10 23:05 | 0.43227 | 0 | 0.75678 | 0.0024793 | 6615 | 0.48571 | 0.75431 | 0.75431 | 0.7543 | 0.75429 | 0.7667
| 1000 | 1000 | 2025-05-10 23:07 | 0.89872 | 0 | 0.76421 | 0.0024959 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.75423
| 6630000 | 6630000 | 2025-05-10 23:09 | 0.4173 | 0 | 0.75673 | 0.0025024 | 6630 | 0.46896 | 0.75422 | 0.75423 | 0.75422 | 0.75422 | 0.76674
| 1000 | 1000 | 2025-05-10 23:11 | 0.89866 | 0 | 0.76422 | 0.0025266 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.75411
| 6645000 | 6645000 | 2025-05-10 23:13 | 0.40232 | 0 | 0.75664 | 0.0025305 | 6645 | 0.45218 | 0.75412 | 0.75411 | 0.75411 | 0.75411 | 0.76677
| 1000 | 1000 | 2025-05-10 23:15 | 0.89865 | 0 | 0.76423 | 0.0025473 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.75404
| 6660000 | 6660000 | 2025-05-10 23:17 | 0.38735 | 0 | 0.75658 | 0.0025457 | 6660 | 0.43538 | 0.75404 | 0.75403 | 0.75403 | 0.75402 | 0.76676
| 1000 | 1000 | 2025-05-10 23:19 | 0.89865 | 0 | 0.76421 | 0.0025595 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.75397
| 6675000 | 6675000 | 2025-05-10 23:21 | 0.37242 | 0 | 0.75651 | 0.0025607 | 6675 | 0.41859 | 0.75396 | 0.75395 | 0.75395 | 0.75395 | 0.76676
| 1000 | 1000 | 2025-05-10 23:23 | 0.89871 | 0 | 0.76419 | 0.0025716 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.75391
| 6690000 | 6690000 | 2025-05-10 23:25 | 0.35743 | 0 | 0.75648 | 0.0025781 | 6690 | 0.4018 | 0.7539 | 0.7539 | 0.7539 | 0.7539 | 0.76679
| 1000 | 1000 | 2025-05-10 23:27 | 0.89862 | 0 | 0.76418 | 0.002617 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.75371
| 6705000 | 6705000 | 2025-05-10 23:29 | 0.34255 | 0 | 0.75632 | 0.0026143 | 6705 | 0.3851 | 0.7537 | 0.7537 | 0.7537 | 0.7537 | 0.76677
| 1000 | 1000 | 2025-05-10 23:31 | 0.8987 | 0 | 0.76415 | 0.0026361 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.76679 | 0.7536
| 6720000 | 6720000 | 2025-05-10 23:33 | 0.32776 | 0 | 0.75623 | 0.0026335 | 6720 | 0.36849 | 0.7536 | 0.7536 | 0.75359 | 0.75359 | 0.76676
| 1000 | 1000 | 2025-05-10 23:35 | 0.89874 | 0 | 0.76411 | 0.0026617 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.75347
| 6735000 | 6735000 | 2025-05-10 23:37 | 0.31299 | 0 | 0.75611 | 0.002646 | 6735 | 0.35195 | 0.75347 | 0.75346 | 0.75346 | 0.75345 | 0.76669
| 1000 | 1000 | 2025-05-10 23:39 | 0.89865 | 0 | 0.76405 | 0.0026505 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.75345
| 6750000 | 6750000 | 2025-05-10 23:41 | 0.29832 | 0 | 0.7561 | 0.0026591 | 6750 | 0.33546 | 0.75344 | 0.75344 | 0.75344 | 0.75343 | 0.76673
| 1000 | 1000 | 2025-05-10 23:43 | 0.89862 | 0 | 0.76405 | 0.0026874 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.7533
| 6765000 | 6765000 | 2025-05-10 23:45 | 0.28379 | 0 | 0.75599 | 0.0026932 | 6765 | 0.31916 | 0.75331 | 0.7533 | 0.7533 | 0.7533 | 0.76677
| 1000 | 1000 | 2025-05-10 23:47 | 0.89866 | 0 | 0.76406 | 0.0027153 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76677 | 0.7532
| 6780000 | 6780000 | 2025-05-10 23:49 | 0.26939 | 0 | 0.75591 | 0.002723 | 6780 | 0.30297 | 0.75319 | 0.75319 | 0.75318 | 0.75318 | 0.7668
| 1000 | 1000 | 2025-05-10 23:51 | 0.89872 | 0 | 0.76406 | 0.0027441 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.76681 | 0.75308
| 6795000 | 6795000 | 2025-05-10 23:53 | 0.25511 | 0 | 0.7558 | 0.0027323 | 6795 | 0.28694 | 0.75307 | 0.75307 | 0.75306 | 0.75306 | 0.76673
| 1000 | 1000 | 2025-05-10 23:55 | 0.8987 | 0 | 0.764 | 0.0027357 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.75306
| 6810000 | 6810000 | 2025-05-10 23:57 | 0.24101 | 0 | 0.75579 | 0.0027419 | 6810 | 0.27109 | 0.75305 | 0.75305 | 0.75305 | 0.75304 | 0.76676
| 1000 | 1000 | 2025-05-10 23:59 | 0.89868 | 0 | 0.76399 | 0.0027717 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.7529
| 6825000 | 6825000 | 2025-05-11 00:01 | 0.22708 | 0 | 0.75566 | 0.0027668 | 6825 | 0.25544 | 0.7529 | 0.7529 | 0.75289 | 0.75288 | 0.76673
| 1000 | 1000 | 2025-05-11 00:03 | 0.89869 | 0 | 0.76394 | 0.0027945 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76673 | 0.75276
| 6840000 | 6840000 | 2025-05-11 00:05 | 0.21337 | 0 | 0.75555 | 0.0027976 | 6840 | 0.24005 | 0.75275 | 0.75275 | 0.75275 | 0.75274 | 0.76674
| 1000 | 1000 | 2025-05-11 00:07 | 0.89872 | 0 | 0.76393 | 0.0028224 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.75264
| 6855000 | 6855000 | 2025-05-11 00:09 | 0.19982 | 0 | 0.75544 | 0.0028174 | 6855 | 0.22483 | 0.75263 | 0.75263 | 0.75262 | 0.75262 | 0.76671
| 1000 | 1000 | 2025-05-11 00:11 | 0.8987 | 0 | 0.7639 | 0.002817 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.75264
| 6870000 | 6870000 | 2025-05-11 00:13 | 0.1865 | 0 | 0.75546 | 0.0028303 | 6870 | 0.20983 | 0.75263 | 0.75263 | 0.75263 | 0.75263 | 0.76678
| 1000 | 1000 | 2025-05-11 00:15 | 0.89875 | 0 | 0.76392 | 0.0028716 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.75243
| 6885000 | 6885000 | 2025-05-11 00:17 | 0.17341 | 0 | 0.75528 | 0.0028609 | 6885 | 0.19516 | 0.75242 | 0.75242 | 0.75242 | 0.75242 | 0.76672
| 1000 | 1000 | 2025-05-11 00:19 | 0.89866 | 0 | 0.76385 | 0.0028759 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76673 | 0.75235
| 6900000 | 6900000 | 2025-05-11 00:21 | 0.16056 | 0 | 0.7552 | 0.0028667 | 6900 | 0.18068 | 0.75234 | 0.75234 | 0.75233 | 0.75232 | 0.76667
| 1000 | 1000 | 2025-05-11 00:23 | 0.89875 | 0 | 0.76378 | 0.0028893 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76667 | 0.75223
| 6915000 | 6915000 | 2025-05-11 00:25 | 0.14798 | 0 | 0.75512 | 0.0029057 | 6915 | 0.16656 | 0.75222 | 0.75222 | 0.75221 | 0.75221 | 0.76674
| 1000 | 1000 | 2025-05-11 00:27 | 0.89868 | 0 | 0.76383 | 0.0029229 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.75214
| 6930000 | 6930000 | 2025-05-11 00:29 | 0.13566 | 0 | 0.75505 | 0.0029273 | 6930 | 0.15269 | 0.75213 | 0.75213 | 0.75212 | 0.75212 | 0.76676
| 1000 | 1000 | 2025-05-11 00:31 | 0.89875 | 0 | 0.76382 | 0.0029477 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.75203
| 6945000 | 6945000 | 2025-05-11 00:33 | 0.12363 | 0 | 0.75496 | 0.0029406 | 6945 | 0.13918 | 0.75203 | 0.75202 | 0.75202 | 0.75201 | 0.76672
| 1000 | 1000 | 2025-05-11 00:35 | 0.89865 | 0 | 0.76378 | 0.00295 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.75198
| 6960000 | 6960000 | 2025-05-11 00:37 | 0.11186 | 0 | 0.75493 | 0.0029562 | 6960 | 0.12592 | 0.75198 | 0.75198 | 0.75197 | 0.75196 | 0.76675
| 1000 | 1000 | 2025-05-11 00:39 | 0.89874 | 0 | 0.76376 | 0.0029935 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.75179
| 6975000 | 6975000 | 2025-05-11 00:41 | 0.10041 | 0 | 0.75478 | 0.0029979 | 6975 | 0.11306 | 0.75178 | 0.75178 | 0.75178 | 0.75178 | 0.76677
| 1000 | 1000 | 2025-05-11 00:43 | 0.89869 | 0 | 0.76377 | 0.0030016 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76677 | 0.75177
| 6990000 | 6990000 | 2025-05-11 00:45 | 0.089257 | 0 | 0.75476 | 0.0030047 | 6990 | 0.1005 | 0.75176 | 0.75176 | 0.75176 | 0.75175 | 0.76678
| 1000 | 1000 | 2025-05-11 00:47 | 0.8987 | 0 | 0.76375 | 0.003032 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.76679 | 0.75163
| 7005000 | 7005000 | 2025-05-11 00:49 | 0.078416 | 0 | 0.75465 | 0.003031 | 7005 | 0.088307 | 0.75162 | 0.75162 | 0.75162 | 0.75162 | 0.76677
| 1000 | 1000 | 2025-05-11 00:51 | 0.89867 | 0 | 0.76372 | 0.0030613 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.76679 | 0.75148
| 7020000 | 7020000 | 2025-05-11 00:53 | 0.067904 | 0 | 0.75453 | 0.0030478 | 7020 | 0.076475 | 0.75148 | 0.75148 | 0.75148 | 0.75148 | 0.76672
| 1000 | 1000 | 2025-05-11 00:55 | 0.89867 | 0 | 0.76366 | 0.0030655 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76673 | 0.7514
| 7035000 | 7035000 | 2025-05-11 00:57 | 0.05772 | 0 | 0.75445 | 0.0030616 | 7035 | 0.065008 | 0.75139 | 0.75139 | 0.75139 | 0.75138 | 0.7667
| 1000 | 1000 | 2025-05-11 00:59 | 0.8987 | 0 | 0.76364 | 0.0030735 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.75134
| 7050000 | 7050000 | 2025-05-11 01:02 | 0.047875 | 0 | 0.75443 | 0.0030903 | 7050 | 0.053929 | 0.75135 | 0.75134 | 0.75134 | 0.75134 | 0.7668
| 1000 | 1000 | 2025-05-11 01:03 | 0.89865 | 0 | 0.76369 | 0.0031045 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.75128
| 7065000 | 7065000 | 2025-05-11 01:06 | 0.03838 | 0 | 0.75438 | 0.0031114 | 7065 | 0.043234 | 0.75127 | 0.75127 | 0.75127 | 0.75127 | 0.76683
| 1000 | 1000 | 2025-05-11 01:07 | 0.89868 | 0 | 0.76369 | 0.0031416 | #NAME? | #NAME? | 0.76683 | 0.76682 | 0.76683 | 0.76683 | 0.75112
| 7080000 | 7080000 | 2025-05-11 01:10 | 0.029238 | 0 | 0.75424 | 0.0031295 | 7080 | 0.032938 | 0.75112 | 0.75111 | 0.75112 | 0.75111 | 0.76676
| 1000 | 1000 | 2025-05-11 01:11 | 0.89868 | 0 | 0.76361 | 0.0031628 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76678 | 0.75096
| 7095000 | 7095000 | 2025-05-11 01:14 | 0.020479 | 0 | 0.75412 | 0.0031595 | 7095 | 0.023073 | 0.75096 | 0.75096 | 0.75096 | 0.75096 | 0.76676
| 1000 | 1000 | 2025-05-11 01:15 | 0.89871 | 0 | 0.76358 | 0.0031809 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.75086
| 7110000 | 7110000 | 2025-05-11 01:18 | 0.012128 | 0 | 0.75403 | 0.00318 | 7110 | 0.013666 | 0.75086 | 0.75085 | 0.75085 | 0.75084 | 0.76675
| 1000 | 1000 | 2025-05-11 01:19 | 0.89871 | 0 | 0.76355 | 0.0032089 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76677 | 0.75072
| 7125000 | 7125000 | 2025-05-11 01:22 | 0.0048014 | 0 | 0.75391 | 0.0032062 | 7125 | 0.0054106 | 0.75071 | 0.7507 | 0.7507 | 0.75069 | 0.76673
| 1000 | 1000 | 2025-05-11 01:23 | 0.89872 | 0 | 0.76352 | 0.0032223 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.75063
| 7140000 | 7140000 | 2025-05-11 01:26 | 0.0052062 | 0 | 0.75384 | 0.0032221 | 7140 | 0.0058671 | 0.75062 | 0.75062 | 0.75062 | 0.75061 | 0.76673
| 1000 | 1000 | 2025-05-11 01:27 | 0.8987 | 0 | 0.76348 | 0.003263 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.75043
| 7155000 | 7155000 | 2025-05-11 01:30 | 0.011948 | 0 | 0.75368 | 0.0032654 | 7155 | 0.013467 | 0.75042 | 0.75041 | 0.75041 | 0.75042 | 0.76674
| 1000 | 1000 | 2025-05-11 01:32 | 0.89872 | 0 | 0.76348 | 0.0032652 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.75042
| 7170000 | 7170000 | 2025-05-11 01:34 | 0.018896 | 0 | 0.75368 | 0.0032692 | 7170 | 0.021299 | 0.75041 | 0.75041 | 0.75041 | 0.75041 | 0.76676
| 1000 | 1000 | 2025-05-11 01:36 | 0.89869 | 0 | 0.76347 | 0.0032962 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.75028
| 7185000 | 7185000 | 2025-05-11 01:38 | 0.025534 | 0 | 0.75357 | 0.0032964 | 7185 | 0.028787 | 0.75028 | 0.75028 | 0.75028 | 0.75027 | 0.76676
| 1000 | 1000 | 2025-05-11 01:40 | 0.89861 | 0 | 0.76345 | 0.0033161 | #NAME? | #NAME? | 0.76675 | 0.76677 | 0.76678 | 0.76678 | 0.75019
| 7200000 | 7200000 | 2025-05-11 01:42 | 0.031867 | 0 | 0.75351 | 0.0033259 | 7200 | 0.035927 | 0.75019 | 0.75019 | 0.75018 | 0.75018 | 0.76681
| 1000 | 1000 | 2025-05-11 01:44 | 0.89871 | 0 | 0.76346 | 0.0033552 | #NAME? | #NAME? | 0.76681 | 0.76682 | 0.76682 | 0.76682 | 0.75004
| 7215000 | 7215000 | 2025-05-11 01:46 | 0.03786 | 0 | 0.75337 | 0.0033359 | 7215 | 0.042686 | 0.75004 | 0.75004 | 0.75004 | 0.75003 | 0.76672
| 1000 | 1000 | 2025-05-11 01:48 | 0.89872 | 0 | 0.76337 | 0.0033527 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.74996
| 7230000 | 7230000 | 2025-05-11 01:50 | 0.043498 | 0 | 0.75331 | 0.0033609 | 7230 | 0.049048 | 0.74995 | 0.74996 | 0.74995 | 0.74995 | 0.76676
| 1000 | 1000 | 2025-05-11 01:52 | 0.89869 | 0 | 0.76338 | 0.0033849 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.74984
| 7245000 | 7245000 | 2025-05-11 01:54 | 0.048812 | 0 | 0.75321 | 0.0033879 | 7245 | 0.055047 | 0.74983 | 0.74983 | 0.74982 | 0.74982 | 0.76677
| 1000 | 1000 | 2025-05-11 01:56 | 0.89868 | 0 | 0.76336 | 0.0034081 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.74973
| 7260000 | 7260000 | 2025-05-11 01:58 | 0.053772 | 0 | 0.75313 | 0.0034115 | 7260 | 0.060651 | 0.74972 | 0.74972 | 0.74971 | 0.74971 | 0.76677
| 1000 | 1000 | 2025-05-11 02:00 | 0.8986 | 0 | 0.76334 | 0.0034426 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.76679 | 0.74957
| 7275000 | 7275000 | 2025-05-11 02:02 | 0.058391 | 0 | 0.753 | 0.0034353 | 7275 | 0.065868 | 0.74957 | 0.74957 | 0.74955 | 0.74955 | 0.76674
| 1000 | 1000 | 2025-05-11 02:04 | 0.89862 | 0 | 0.7633 | 0.0034498 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.7495
| 7290000 | 7290000 | 2025-05-11 02:06 | 0.062693 | 0 | 0.75295 | 0.0034553 | 7290 | 0.070723 | 0.7495 | 0.7495 | 0.74949 | 0.74949 | 0.76677
| 1000 | 1000 | 2025-05-11 02:08 | 0.89864 | 0 | 0.7633 | 0.0034806 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.74937
| 7305000 | 7305000 | 2025-05-11 02:10 | 0.066658 | 0 | 0.75284 | 0.0034784 | 7305 | 0.075199 | 0.74937 | 0.74936 | 0.74936 | 0.74935 | 0.76675
| 1000 | 1000 | 2025-05-11 02:12 | 0.89869 | 0 | 0.76326 | 0.0035038 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.74924
| 7320000 | 7320000 | 2025-05-11 02:14 | 0.070314 | 0 | 0.75273 | 0.003498 | 7320 | 0.079319 | 0.74925 | 0.74924 | 0.74924 | 0.74923 | 0.76673
| 1000 | 1000 | 2025-05-11 02:16 | 0.89881 | 0 | 0.76321 | 0.0035222 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.74913
| 7335000 | 7335000 | 2025-05-11 02:18 | 0.073631 | 0 | 0.75264 | 0.003526 | 7335 | 0.083081 | 0.74912 | 0.74912 | 0.74911 | 0.74911 | 0.76675
| 1000 | 1000 | 2025-05-11 02:20 | 0.89869 | 0 | 0.7632 | 0.0035512 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76675 | 0.749
| 7350000 | 7350000 | 2025-05-11 02:22 | 0.076651 | 0 | 0.75254 | 0.0035478 | 7350 | 0.086494 | 0.749 | 0.74899 | 0.74898 | 0.74898 | 0.76673
| 1000 | 1000 | 2025-05-11 02:24 | 0.8987 | 0 | 0.76314 | 0.0035851 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76672 | 0.76673 | 0.7488
| 7365000 | 7365000 | 2025-05-11 02:26 | 0.079367 | 0 | 0.75238 | 0.0035843 | 7365 | 0.089583 | 0.7488 | 0.74879 | 0.74879 | 0.74879 | 0.76671
| 1000 | 1000 | 2025-05-11 02:28 | 0.89862 | 0 | 0.76313 | 0.0035938 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.74875
| 7380000 | 7380000 | 2025-05-11 02:30 | 0.081793 | 0 | 0.75234 | 0.0035968 | 7380 | 0.092324 | 0.74875 | 0.74874 | 0.74874 | 0.74874 | 0.76673
| 1000 | 1000 | 2025-05-11 02:32 | 0.89862 | 0 | 0.76311 | 0.0036197 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76674 | 0.74863
| 7395000 | 7395000 | 2025-05-11 02:34 | 0.083926 | 0 | 0.75224 | 0.0036245 | 7395 | 0.094739 | 0.74862 | 0.74861 | 0.74861 | 0.74861 | 0.76674
| 1000 | 1000 | 2025-05-11 02:36 | 0.89866 | 0 | 0.76311 | 0.0036436 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.74853
| 7410000 | 7410000 | 2025-05-11 02:38 | 0.085785 | 0 | 0.75217 | 0.0036385 | 7410 | 0.09684 | 0.74853 | 0.74853 | 0.74853 | 0.74852 | 0.76672
| 1000 | 1000 | 2025-05-11 02:40 | 0.89868 | 0 | 0.76306 | 0.0036679 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.74839
| 7425000 | 7425000 | 2025-05-11 02:42 | 0.087371 | 0 | 0.75206 | 0.0036766 | 7425 | 0.09864 | 0.74838 | 0.74838 | 0.74838 | 0.74838 | 0.76676
| 1000 | 1000 | 2025-05-11 02:44 | 0.89875 | 0 | 0.76308 | 0.0036897 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.74832
| 7440000 | 7440000 | 2025-05-11 02:46 | 0.088703 | 0 | 0.75201 | 0.0036924 | 7440 | 0.10015 | 0.74832 | 0.74832 | 0.74832 | 0.74831 | 0.76678
| 1000 | 1000 | 2025-05-11 02:48 | 0.89874 | 0 | 0.76305 | 0.003731 | #NAME? | #NAME? | 0.76678 | 0.76677 | 0.76678 | 0.76678 | 0.74812
| 7455000 | 7455000 | 2025-05-11 02:50 | 0.089776 | 0 | 0.75185 | 0.0037366 | 7455 | 0.10138 | 0.74812 | 0.74812 | 0.74812 | 0.74811 | 0.7668
| 1000 | 1000 | 2025-05-11 02:52 | 0.89871 | 0 | 0.76305 | 0.0037477 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.74806
| 7470000 | 7470000 | 2025-05-11 02:54 | 0.090594 | 0 | 0.7518 | 0.0037425 | 7470 | 0.10231 | 0.74806 | 0.74806 | 0.74806 | 0.74806 | 0.76677
| 1000 | 1000 | 2025-05-11 02:56 | 0.89863 | 0 | 0.76299 | 0.003779 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.74788
| 7485000 | 7485000 | 2025-05-11 02:58 | 0.091201 | 0 | 0.75164 | 0.0037696 | 7485 | 0.10301 | 0.74788 | 0.74787 | 0.74787 | 0.74786 | 0.76672
| 1000 | 1000 | 2025-05-11 03:00 | 0.89866 | 0 | 0.76295 | 0.0037813 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.74782
| 7500000 | 7500000 | 2025-05-11 03:02 | 0.091556 | 0 | 0.7516 | 0.0037808 | 7500 | 0.10342 | 0.74782 | 0.74782 | 0.74781 | 0.74782 | 0.76672
STOP | 1000 | 1000 | 2025-05-11 03:04 | 0.89865 | 0 | 0.76292 | 0.0038032 | 1 | #DIV/0! | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.74771
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
Note on D2: labina: Real sine amplitude, i.e. A*sin(w*t). It is NOT rms or peak-to-peak value.
Note on D7: labina: The algorithm must generate output quantity 'rms'. That is e.g. TWM-WFFT or TWM-WRMS
Note on D8: labina: Set to 1 to measure simultaneously both Digitizer amplitude and DC multimeter voltage. Set to 0 or keep empty to measure first AC and DC after.
Note on C12: Martin Šíra: The record length should be such to get coherent sampling. If the number of periods is the same as the signal frequency, coherent sampling should be always met, however the data processing can take a lot of time.