## List1
Frequency response measurement template |  |  | measurement properties |  |  | hardware data |  |  |  | Actual number of readings: |  |  |  | 666
Software will read settings from cells E2 to E7 and |  |  | AC voltage amplitude | 1.414 | V | AC source idn | TEKTRONIX,AFG3101C,C010192,SCPI:99.0 FV:1.0.2 |  |  | Crude estimate of measurement time (h): |  |  |  | 3.95
measurement points from collumns A-C (yellow cells). |  |  | Sampling frequency | 10000000 | Hz | DC meter idn | Keithley 2182
Results will be written into collumns D-H (and K+). Sheet must |  |  | AC-DC settle time | 20 | s | Digitizer idn | NI PXI-5922, sn. F22D5B, channel 1
contain keywords START and STOP in collumn A, values of |  |  | AC-DC warm up time | 900 | s
frequency in collumn B, and record length in C. |  |  | DC meter readings | 5
Place keyword STOP to the row with half of the sampling frequency. |  |  | Algorithm | TWM-WRMS
|  |  | Simultaneous meas.? | 1
inputs |  |  | results
Command | Signal frequency | Record length | TIme of DC readings | Digitizer amplitude rms |  | DC voltage |  | Calculation |  | DC meter readings
(START, STOP) | Hz | multiple of periods | (time of writing this line into excel) | value (V) | unc (V) | value (V) | unc (V) | for figure |  | ¦---------------->
START | 15000 | 15000 | 2026-02-25 08:44 | 0.90106 | 1e-06 | 0.76879 | 1.7131e-06 | 15 | 0.99998 | 0.7688 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2026-02-25 08:44 | 0.89877 | 1e-06 | 0.76682 | 1.4806e-06 | #NAME? | #NAME? | 0.76681 | 0.76682 | 0.76682 | 0.76682 | 0.76682
| 30000 | 30000 | 2026-02-25 08:44 | 0.89885 | 1e-06 | 0.76695 | 3.2364e-06 | 30 | 0.99991 | 0.76696 | 0.76696 | 0.76695 | 0.76695 | 0.76695
| 1000 | 1000 | 2026-02-25 08:45 | 0.89874 | 1e-06 | 0.7668 | 9.7198e-07 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.76679 | 0.7668 | 0.7668
| 45000 | 45000 | 2026-02-25 08:45 | 0.89836 | 1e-06 | 0.76663 | 2.0021e-06 | 45 | 0.99979 | 0.76664 | 0.76663 | 0.76663 | 0.76663 | 0.76663
| 1000 | 1000 | 2026-02-25 08:46 | 0.89871 | 1e-06 | 0.76676 | 1.5037e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 60000 | 60000 | 2026-02-25 08:46 | 0.89809 | 1e-06 | 0.76652 | 1.4141e-06 | 60 | 0.99964 | 0.76652 | 0.76651 | 0.76651 | 0.76651 | 0.76651
| 1000 | 1000 | 2026-02-25 08:46 | 0.8987 | 1e-06 | 0.76676 | 3.0228e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 75000 | 75000 | 2026-02-25 08:47 | 0.89783 | 1e-06 | 0.76645 | 2.091e-06 | 75 | 0.99945 | 0.76646 | 0.76645 | 0.76645 | 0.76645 | 0.76645
| 1000 | 1000 | 2026-02-25 08:47 | 0.89874 | 1e-06 | 0.76681 | 2.3081e-06 | #NAME? | #NAME? | 0.7668 | 0.76681 | 0.7668 | 0.76681 | 0.76681
| 90000 | 90000 | 2026-02-25 08:48 | 0.89762 | 1e-06 | 0.76644 | 1.7203e-06 | 90 | 0.99922 | 0.76644 | 0.76644 | 0.76644 | 0.76645 | 0.76644
| 1000 | 1000 | 2026-02-25 08:48 | 0.89877 | 1e-06 | 0.76683 | 6.5689e-07 | #NAME? | #NAME? | 0.76683 | 0.76683 | 0.76683 | 0.76683 | 0.76683
| 105000 | 105000 | 2026-02-25 08:48 | 0.89736 | 1e-06 | 0.76642 | 9.9065e-07 | 105 | 0.99896 | 0.76642 | 0.76642 | 0.76642 | 0.76642 | 0.76643
| 1000 | 1000 | 2026-02-25 08:49 | 0.8988 | 1e-06 | 0.76685 | 2.3089e-06 | #NAME? | #NAME? | 0.76685 | 0.76685 | 0.76685 | 0.76686 | 0.76686
| 120000 | 120000 | 2026-02-25 08:49 | 0.89703 | 1e-06 | 0.76638 | 1.4218e-06 | 120 | 0.99864 | 0.76639 | 0.76639 | 0.76638 | 0.76638 | 0.76638
| 1000 | 1000 | 2026-02-25 08:50 | 0.89875 | 1e-06 | 0.7668 | 7.6149e-07 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.76681 | 0.7668
| 135000 | 135000 | 2026-02-25 08:50 | 0.89661 | 1e-06 | 0.76627 | 3.0395e-06 | 135 | 0.99833 | 0.76627 | 0.76627 | 0.76626 | 0.76626 | 0.76626
| 1000 | 1000 | 2026-02-25 08:50 | 0.8987 | 1e-06 | 0.76677 | 1.8751e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 150000 | 150000 | 2026-02-25 08:51 | 0.89616 | 1e-06 | 0.76615 | 2.7917e-06 | 150 | 0.99797 | 0.76616 | 0.76615 | 0.76615 | 0.76615 | 0.76614
| 1000 | 1000 | 2026-02-25 08:51 | 0.89868 | 1e-06 | 0.76675 | 3.0159e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 165000 | 165000 | 2026-02-25 08:51 | 0.89581 | 1e-06 | 0.76612 | 7e-07 | 165 | 0.99763 | 0.76612 | 0.76612 | 0.76612 | 0.76612 | 0.76611
| 1000 | 1000 | 2026-02-25 08:52 | 0.89868 | 1e-06 | 0.76674 | 1.0949e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 180000 | 180000 | 2026-02-25 08:52 | 0.89547 | 1e-06 | 0.76611 | 2.4366e-06 | 180 | 0.99726 | 0.76612 | 0.76611 | 0.76611 | 0.7661 | 0.7661
| 1000 | 1000 | 2026-02-25 08:53 | 0.89866 | 1e-06 | 0.76673 | 1.6019e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 195000 | 195000 | 2026-02-25 08:53 | 0.89496 | 1e-06 | 0.76596 | 2.7655e-06 | 195 | 0.99688 | 0.76597 | 0.76596 | 0.76596 | 0.76595 | 0.76595
| 1000 | 1000 | 2026-02-25 08:54 | 0.89864 | 1e-06 | 0.76671 | 2.7814e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 210000 | 210000 | 2026-02-25 08:54 | 0.89456 | 1e-06 | 0.76589 | 7.9539e-07 | 210 | 0.99653 | 0.76589 | 0.76589 | 0.76589 | 0.76589 | 0.76589
| 1000 | 1000 | 2026-02-25 08:54 | 0.89863 | 1e-06 | 0.76671 | 1.3817e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.7667 | 0.76671 | 0.76671
| 225000 | 225000 | 2026-02-25 08:55 | 0.89418 | 1e-06 | 0.76583 | 3.0455e-06 | 225 | 0.99617 | 0.76584 | 0.76583 | 0.76583 | 0.76583 | 0.76582
| 1000 | 1000 | 2026-02-25 08:55 | 0.89869 | 1e-06 | 0.76675 | 3.09e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76676
| 240000 | 240000 | 2026-02-25 08:56 | 0.89389 | 1e-06 | 0.76584 | 1.7672e-06 | 240 | 0.99584 | 0.76584 | 0.76584 | 0.76584 | 0.76585 | 0.76585
| 1000 | 1000 | 2026-02-25 08:56 | 0.89874 | 1e-06 | 0.76679 | 6.2795e-07 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.76679
| 255000 | 255000 | 2026-02-25 08:56 | 0.89356 | 1e-06 | 0.76579 | 3.3e-06 | 255 | 0.99553 | 0.7658 | 0.7658 | 0.7658 | 0.76579 | 0.76578
| 1000 | 1000 | 2026-02-25 08:57 | 0.89873 | 1e-06 | 0.76679 | 2.403e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.76679 | 0.76679
| 270000 | 270000 | 2026-02-25 08:57 | 0.89318 | 1e-06 | 0.76568 | 1.7553e-06 | 270 | 0.99526 | 0.76569 | 0.76569 | 0.76568 | 0.76568 | 0.76568
| 1000 | 1000 | 2026-02-25 08:58 | 0.89868 | 1e-06 | 0.76675 | 1.2338e-06 | #NAME? | #NAME? | 0.76675 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 285000 | 285000 | 2026-02-25 08:58 | 0.89284 | 1e-06 | 0.76558 | 4.5247e-07 | 285 | 0.99501 | 0.76558 | 0.76558 | 0.76558 | 0.76558 | 0.76558
| 1000 | 1000 | 2026-02-25 08:58 | 0.89872 | 1e-06 | 0.76678 | 1.6678e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.76678 | 0.76679
| 300000 | 300000 | 2026-02-25 08:59 | 0.89251 | 1e-06 | 0.76546 | 1.4182e-06 | 300 | 0.9948 | 0.76546 | 0.76546 | 0.76546 | 0.76546 | 0.76545
| 1000 | 1000 | 2026-02-25 08:59 | 0.89868 | 1e-06 | 0.76675 | 7.0925e-07 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 315000 | 315000 | 2026-02-25 09:00 | 0.89225 | 1e-06 | 0.76536 | 2.46e-06 | 315 | 0.99465 | 0.76537 | 0.76536 | 0.76536 | 0.76536 | 0.76536
| 1000 | 1000 | 2026-02-25 09:00 | 0.89863 | 1e-06 | 0.76671 | 1.2641e-06 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 330000 | 330000 | 2026-02-25 09:00 | 0.89202 | 1e-06 | 0.76525 | 1.6763e-06 | 330 | 0.99452 | 0.76526 | 0.76526 | 0.76525 | 0.76525 | 0.76525
| 1000 | 1000 | 2026-02-25 09:01 | 0.89867 | 1e-06 | 0.76673 | 9.3616e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 345000 | 345000 | 2026-02-25 09:01 | 0.892 | 1e-06 | 0.76528 | 7.4508e-07 | 345 | 0.99447 | 0.76528 | 0.76528 | 0.76528 | 0.76528 | 0.76528
| 1000 | 1000 | 2026-02-25 09:01 | 0.89872 | 1e-06 | 0.76677 | 2.5688e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 360000 | 360000 | 2026-02-25 09:02 | 0.89192 | 1e-06 | 0.76522 | 2.0493e-06 | 360 | 0.99447 | 0.76522 | 0.76523 | 0.76522 | 0.76522 | 0.76521
| 1000 | 1000 | 2026-02-25 09:02 | 0.89868 | 1e-06 | 0.76675 | 1.1076e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 375000 | 375000 | 2026-02-25 09:03 | 0.89185 | 1e-06 | 0.76512 | 2.0104e-06 | 375 | 0.99452 | 0.76513 | 0.76512 | 0.76512 | 0.76512 | 0.76512
| 1000 | 1000 | 2026-02-25 09:03 | 0.89869 | 1e-06 | 0.76676 | 1.6496e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 390000 | 390000 | 2026-02-25 09:03 | 0.89186 | 1e-06 | 0.76505 | 9.0966e-07 | 390 | 0.99461 | 0.76505 | 0.76505 | 0.76505 | 0.76505 | 0.76505
| 1000 | 1000 | 2026-02-25 09:04 | 0.89864 | 1e-06 | 0.76671 | 1.7071e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 405000 | 405000 | 2026-02-25 09:04 | 0.89188 | 1e-06 | 0.76493 | 1.3608e-06 | 405 | 0.99479 | 0.76494 | 0.76493 | 0.76493 | 0.76493 | 0.76493
| 1000 | 1000 | 2026-02-25 09:05 | 0.89861 | 1e-06 | 0.76669 | 2.2932e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 420000 | 420000 | 2026-02-25 09:05 | 0.89209 | 1e-06 | 0.76494 | 1.5282e-06 | 420 | 0.995 | 0.76494 | 0.76493 | 0.76494 | 0.76494 | 0.76494
| 1000 | 1000 | 2026-02-25 09:05 | 0.89874 | 1e-06 | 0.76679 | 3.5743e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.7668 | 0.7668
| 435000 | 435000 | 2026-02-25 09:06 | 0.89234 | 1e-06 | 0.76494 | 1.0536e-06 | 435 | 0.99528 | 0.76494 | 0.76494 | 0.76494 | 0.76494 | 0.76495
| 1000 | 1000 | 2026-02-25 09:06 | 0.89876 | 1e-06 | 0.76681 | 3.9866e-06 | #NAME? | #NAME? | 0.7668 | 0.76681 | 0.76681 | 0.76682 | 0.76682
| 450000 | 450000 | 2026-02-25 09:07 | 0.89262 | 1e-06 | 0.76493 | 2.5373e-06 | 450 | 0.99562 | 0.76494 | 0.76493 | 0.76493 | 0.76493 | 0.76492
| 1000 | 1000 | 2026-02-25 09:07 | 0.89872 | 1e-06 | 0.76678 | 1.3421e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 465000 | 465000 | 2026-02-25 09:07 | 0.89278 | 1e-06 | 0.76477 | 3.1282e-06 | 465 | 0.99601 | 0.76477 | 0.76477 | 0.76476 | 0.76476 | 0.76476
| 1000 | 1000 | 2026-02-25 09:08 | 0.89866 | 1e-06 | 0.76673 | 1.7629e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76672 | 0.76673 | 0.76673
| 480000 | 480000 | 2026-02-25 09:08 | 0.89315 | 1e-06 | 0.76475 | 2.1649e-06 | 480 | 0.99643 | 0.76475 | 0.76475 | 0.76475 | 0.76474 | 0.76474
| 1000 | 1000 | 2026-02-25 09:09 | 0.89871 | 1e-06 | 0.76677 | 2.7843e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76678
| 495000 | 495000 | 2026-02-25 09:09 | 0.8935 | 1e-06 | 0.76467 | 2.8313e-06 | 495 | 0.99693 | 0.76467 | 0.76467 | 0.76467 | 0.76467 | 0.76466
| 1000 | 1000 | 2026-02-25 09:09 | 0.89868 | 1e-06 | 0.76675 | 1.1089e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 510000 | 510000 | 2026-02-25 09:10 | 0.89389 | 1e-06 | 0.76461 | 1.1526e-06 | 510 | 0.99744 | 0.76462 | 0.76462 | 0.76461 | 0.76461 | 0.76461
| 1000 | 1000 | 2026-02-25 09:10 | 0.89868 | 1e-06 | 0.76674 | 2.1934e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 525000 | 525000 | 2026-02-25 09:11 | 0.89436 | 1e-06 | 0.7646 | 9.7021e-07 | 525 | 0.99798 | 0.76461 | 0.7646 | 0.7646 | 0.7646 | 0.7646
| 1000 | 1000 | 2026-02-25 09:11 | 0.89867 | 1e-06 | 0.76673 | 1.7831e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 540000 | 540000 | 2026-02-25 09:11 | 0.89485 | 1e-06 | 0.76456 | 1.103e-06 | 540 | 0.99858 | 0.76457 | 0.76456 | 0.76456 | 0.76456 | 0.76456
| 1000 | 1000 | 2026-02-25 09:12 | 0.89866 | 1e-06 | 0.76673 | 1.8082e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 555000 | 555000 | 2026-02-25 09:12 | 0.89531 | 1e-06 | 0.7645 | 7.6093e-07 | 555 | 0.99917 | 0.7645 | 0.7645 | 0.76451 | 0.7645 | 0.7645
| 1000 | 1000 | 2026-02-25 09:13 | 0.89875 | 1e-06 | 0.76681 | 3.9365e-06 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.76681 | 0.76681 | 0.76682
| 570000 | 570000 | 2026-02-25 09:13 | 0.89581 | 1e-06 | 0.76445 | 3.9398e-06 | 570 | 0.9998 | 0.76446 | 0.76445 | 0.76445 | 0.76444 | 0.76444
| 1000 | 1000 | 2026-02-25 09:13 | 0.8987 | 1e-06 | 0.76676 | 4.4844e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76677 | 0.76677
| 585000 | 585000 | 2026-02-25 09:14 | 0.89638 | 1e-06 | 0.76446 | 3.2324e-06 | 585 | 1.0004 | 0.76447 | 0.76447 | 0.76446 | 0.76446 | 0.76445
| 1000 | 1000 | 2026-02-25 09:14 | 0.89869 | 1e-06 | 0.76676 | 1.0356e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 600000 | 600000 | 2026-02-25 09:15 | 0.8968 | 1e-06 | 0.76434 | 1.9007e-06 | 600 | 1.001 | 0.76434 | 0.76434 | 0.76434 | 0.76433 | 0.76433
| 1000 | 1000 | 2026-02-25 09:15 | 0.89865 | 1e-06 | 0.76672 | 1.8863e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 615000 | 615000 | 2026-02-25 09:15 | 0.89725 | 1e-06 | 0.76425 | 2.4482e-06 | 615 | 1.0017 | 0.76426 | 0.76426 | 0.76425 | 0.76425 | 0.76425
| 1000 | 1000 | 2026-02-25 09:16 | 0.89862 | 1e-06 | 0.7667 | 1.6689e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 630000 | 630000 | 2026-02-25 09:16 | 0.89772 | 1e-06 | 0.7642 | 2.7908e-06 | 630 | 1.0023 | 0.7642 | 0.7642 | 0.7642 | 0.76419 | 0.76419
| 1000 | 1000 | 2026-02-25 09:17 | 0.89863 | 1e-06 | 0.76671 | 2.182e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 645000 | 645000 | 2026-02-25 09:17 | 0.89824 | 1e-06 | 0.76418 | 2.5493e-06 | 645 | 1.0029 | 0.76419 | 0.76419 | 0.76418 | 0.76418 | 0.76417
| 1000 | 1000 | 2026-02-25 09:17 | 0.89862 | 1e-06 | 0.76669 | 9.07e-07 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 660000 | 660000 | 2026-02-25 09:18 | 0.8987 | 1e-06 | 0.76415 | 8.5532e-07 | 660 | 1.0034 | 0.76415 | 0.76415 | 0.76415 | 0.76415 | 0.76415
| 1000 | 1000 | 2026-02-25 09:18 | 0.89868 | 1e-06 | 0.76675 | 6.861e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 675000 | 675000 | 2026-02-25 09:19 | 0.8991 | 1e-06 | 0.76409 | 2.6944e-06 | 675 | 1.0039 | 0.7641 | 0.76409 | 0.76409 | 0.76409 | 0.76408
| 1000 | 1000 | 2026-02-25 09:19 | 0.89863 | 1e-06 | 0.7667 | 1.3233e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.7667 | 0.7667
| 690000 | 690000 | 2026-02-25 09:19 | 0.8995 | 1e-06 | 0.76406 | 2.0823e-06 | 690 | 1.0044 | 0.76407 | 0.76407 | 0.76406 | 0.76406 | 0.76406
| 1000 | 1000 | 2026-02-25 09:20 | 0.89863 | 1e-06 | 0.7667 | 1.3658e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.7667
| 705000 | 705000 | 2026-02-25 09:20 | 0.89979 | 1e-06 | 0.76398 | 2.7004e-06 | 705 | 1.0048 | 0.76399 | 0.76398 | 0.76398 | 0.76398 | 0.76398
| 1000 | 1000 | 2026-02-25 09:21 | 0.89862 | 1e-06 | 0.76669 | 1.498e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 720000 | 720000 | 2026-02-25 09:21 | 0.90018 | 1e-06 | 0.76402 | 1.116e-06 | 720 | 1.0052 | 0.76402 | 0.76402 | 0.76402 | 0.76402 | 0.76402
| 1000 | 1000 | 2026-02-25 09:21 | 0.8987 | 1e-06 | 0.76676 | 3.7297e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 735000 | 735000 | 2026-02-25 09:22 | 0.90045 | 1e-06 | 0.764 | 3.9787e-06 | 735 | 1.0056 | 0.76401 | 0.76401 | 0.764 | 0.764 | 0.76399
| 1000 | 1000 | 2026-02-25 09:22 | 0.89867 | 1e-06 | 0.76673 | 1.5247e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 750000 | 750000 | 2026-02-25 09:23 | 0.90069 | 1e-06 | 0.76401 | 5.7084e-07 | 750 | 1.0058 | 0.76401 | 0.76401 | 0.76401 | 0.76401 | 0.76401
| 1000 | 1000 | 2026-02-25 09:23 | 0.89871 | 1e-06 | 0.76678 | 4.7147e-07 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76677 | 0.76678 | 0.76678
| 765000 | 765000 | 2026-02-25 09:23 | 0.90072 | 1e-06 | 0.76388 | 3.3254e-06 | 765 | 1.006 | 0.76389 | 0.76389 | 0.76388 | 0.76388 | 0.76387
| 1000 | 1000 | 2026-02-25 09:24 | 0.89865 | 1e-06 | 0.76673 | 3.4162e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76674
| 780000 | 780000 | 2026-02-25 09:24 | 0.90083 | 1e-06 | 0.76387 | 2.5342e-06 | 780 | 1.0062 | 0.76387 | 0.76387 | 0.76387 | 0.76387 | 0.76386
| 1000 | 1000 | 2026-02-25 09:25 | 0.89865 | 1e-06 | 0.76672 | 9.9432e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 795000 | 795000 | 2026-02-25 09:25 | 0.90078 | 1e-06 | 0.76378 | 2.1599e-06 | 795 | 1.0062 | 0.76378 | 0.76378 | 0.76378 | 0.76378 | 0.76377
| 1000 | 1000 | 2026-02-25 09:25 | 0.89862 | 1e-06 | 0.7667 | 2.0199e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 810000 | 810000 | 2026-02-25 09:26 | 0.90073 | 1e-06 | 0.76375 | 2.6679e-06 | 810 | 1.0062 | 0.76376 | 0.76375 | 0.76374 | 0.76374 | 0.76374
| 1000 | 1000 | 2026-02-25 09:26 | 0.89861 | 1e-06 | 0.76669 | 1.6798e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 825000 | 825000 | 2026-02-25 09:27 | 0.90073 | 1e-06 | 0.7638 | 1.8203e-06 | 825 | 1.0062 | 0.7638 | 0.7638 | 0.76381 | 0.7638 | 0.7638
| 1000 | 1000 | 2026-02-25 09:27 | 0.89868 | 1e-06 | 0.76675 | 4.7633e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 840000 | 840000 | 2026-02-25 09:27 | 0.90059 | 1e-06 | 0.7638 | 1.3849e-06 | 840 | 1.006 | 0.76379 | 0.76379 | 0.7638 | 0.7638 | 0.7638
| 1000 | 1000 | 2026-02-25 09:28 | 0.89871 | 1e-06 | 0.76678 | 8.2943e-07 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 855000 | 855000 | 2026-02-25 09:28 | 0.90029 | 1e-06 | 0.76371 | 1.9632e-06 | 855 | 1.0058 | 0.76371 | 0.76371 | 0.76371 | 0.7637 | 0.7637
| 1000 | 1000 | 2026-02-25 09:29 | 0.89869 | 1e-06 | 0.76675 | 1.3256e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76675
| 870000 | 870000 | 2026-02-25 09:29 | 0.90003 | 1e-06 | 0.76369 | 2.8328e-06 | 870 | 1.0055 | 0.7637 | 0.76369 | 0.76369 | 0.76369 | 0.76368
| 1000 | 1000 | 2026-02-25 09:29 | 0.89869 | 1e-06 | 0.76675 | 2.3968e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76676
| 885000 | 885000 | 2026-02-25 09:30 | 0.89964 | 1e-06 | 0.76362 | 4.1622e-06 | 885 | 1.0052 | 0.76363 | 0.76363 | 0.76362 | 0.76361 | 0.76361
| 1000 | 1000 | 2026-02-25 09:30 | 0.89866 | 1e-06 | 0.76673 | 1.7354e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 900000 | 900000 | 2026-02-25 09:31 | 0.89931 | 1e-06 | 0.76364 | 1.7221e-06 | 900 | 1.0048 | 0.76364 | 0.76364 | 0.76363 | 0.76363 | 0.76363
| 1000 | 1000 | 2026-02-25 09:31 | 0.89878 | 1e-06 | 0.76683 | 1.7844e-06 | #NAME? | #NAME? | 0.76683 | 0.76683 | 0.76684 | 0.76683 | 0.76683
| 915000 | 915000 | 2026-02-25 09:31 | 0.89892 | 1e-06 | 0.76365 | 2.4064e-06 | 915 | 1.0043 | 0.76366 | 0.76365 | 0.76365 | 0.76365 | 0.76364
| 1000 | 1000 | 2026-02-25 09:32 | 0.89872 | 1e-06 | 0.76678 | 1.424e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 930000 | 930000 | 2026-02-25 09:32 | 0.89844 | 1e-06 | 0.76361 | 2.2806e-06 | 930 | 1.0038 | 0.76362 | 0.76361 | 0.76361 | 0.76361 | 0.76361
| 1000 | 1000 | 2026-02-25 09:33 | 0.89872 | 1e-06 | 0.76678 | 9.3775e-07 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 945000 | 945000 | 2026-02-25 09:33 | 0.89782 | 1e-06 | 0.76349 | 2.0295e-06 | 945 | 1.0033 | 0.7635 | 0.76349 | 0.76349 | 0.76349 | 0.76348
| 1000 | 1000 | 2026-02-25 09:33 | 0.89866 | 1e-06 | 0.76674 | 8.4802e-07 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 960000 | 960000 | 2026-02-25 09:34 | 0.89733 | 1e-06 | 0.7635 | 2.3787e-06 | 960 | 1.0028 | 0.76351 | 0.7635 | 0.7635 | 0.7635 | 0.7635
| 1000 | 1000 | 2026-02-25 09:34 | 0.89869 | 1e-06 | 0.76676 | 2.5339e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 975000 | 975000 | 2026-02-25 09:35 | 0.89668 | 1e-06 | 0.76341 | 2.1877e-06 | 975 | 1.0021 | 0.76342 | 0.76342 | 0.76341 | 0.76341 | 0.76341
| 1000 | 1000 | 2026-02-25 09:35 | 0.89866 | 1e-06 | 0.76673 | 2.6145e-06 | #NAME? | #NAME? | 0.76673 | 0.76672 | 0.76673 | 0.76674 | 0.76674
| 990000 | 990000 | 2026-02-25 09:35 | 0.89607 | 1e-06 | 0.76335 | 1.2322e-06 | 990 | 1.0015 | 0.76335 | 0.76335 | 0.76335 | 0.76335 | 0.76334
| 1000 | 1000 | 2026-02-25 09:36 | 0.89865 | 1e-06 | 0.76673 | 1.5859e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 1005000 | 1005000 | 2026-02-25 09:36 | 0.89554 | 1e-06 | 0.76336 | 1.9985e-06 | 1005 | 1.0009 | 0.76337 | 0.76336 | 0.76335 | 0.76336 | 0.76336
| 1000 | 1000 | 2026-02-25 09:36 | 0.89871 | 1e-06 | 0.76676 | 2.1079e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 1020000 | 1020000 | 2026-02-25 09:37 | 0.89489 | 1e-06 | 0.76329 | 2.3455e-06 | 1020 | 1.0003 | 0.7633 | 0.76329 | 0.76329 | 0.76329 | 0.76328
| 1000 | 1000 | 2026-02-25 09:37 | 0.89865 | 1e-06 | 0.76672 | 2.0459e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1035000 | 1035000 | 2026-02-25 09:38 | 0.89426 | 1e-06 | 0.76323 | 1.9641e-06 | 1035 | 0.99967 | 0.76324 | 0.76323 | 0.76323 | 0.76323 | 0.76323
| 1000 | 1000 | 2026-02-25 09:38 | 0.89864 | 1e-06 | 0.76671 | 1.949e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 1050000 | 1050000 | 2026-02-25 09:38 | 0.89378 | 1e-06 | 0.76328 | 3.1796e-06 | 1050 | 0.99907 | 0.76329 | 0.76328 | 0.76328 | 0.76327 | 0.76327
| 1000 | 1000 | 2026-02-25 09:39 | 0.89871 | 1e-06 | 0.76677 | 1.4322e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76677
| 1065000 | 1065000 | 2026-02-25 09:39 | 0.89317 | 1e-06 | 0.7632 | 3.1759e-06 | 1065 | 0.99849 | 0.76321 | 0.7632 | 0.7632 | 0.76319 | 0.76319
| 1000 | 1000 | 2026-02-25 09:40 | 0.89868 | 1e-06 | 0.76674 | 2.3739e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76675 | 0.76674
| 1080000 | 1080000 | 2026-02-25 09:40 | 0.89267 | 1e-06 | 0.7632 | 2.4084e-06 | 1080 | 0.99793 | 0.7632 | 0.7632 | 0.7632 | 0.76319 | 0.76319
| 1000 | 1000 | 2026-02-25 09:40 | 0.89866 | 1e-06 | 0.76674 | 1.8474e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 1095000 | 1095000 | 2026-02-25 09:41 | 0.89214 | 1e-06 | 0.76313 | 2.1839e-06 | 1095 | 0.99741 | 0.76314 | 0.76313 | 0.76314 | 0.76313 | 0.76313
| 1000 | 1000 | 2026-02-25 09:41 | 0.89867 | 1e-06 | 0.76674 | 6.8113e-07 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 1110000 | 1110000 | 2026-02-25 09:42 | 0.89165 | 1e-06 | 0.76308 | 1.2748e-06 | 1110 | 0.99695 | 0.76308 | 0.76308 | 0.76308 | 0.76308 | 0.76307
| 1000 | 1000 | 2026-02-25 09:42 | 0.89866 | 1e-06 | 0.76673 | 2.3327e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 1125000 | 1125000 | 2026-02-25 09:42 | 0.89117 | 1e-06 | 0.763 | 2.8404e-06 | 1125 | 0.99651 | 0.76301 | 0.763 | 0.763 | 0.76299 | 0.763
| 1000 | 1000 | 2026-02-25 09:43 | 0.89864 | 1e-06 | 0.76672 | 1.5289e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 1140000 | 1140000 | 2026-02-25 09:43 | 0.89079 | 1e-06 | 0.76296 | 1.9063e-06 | 1140 | 0.99615 | 0.76296 | 0.76296 | 0.76296 | 0.76296 | 0.76295
| 1000 | 1000 | 2026-02-25 09:44 | 0.89864 | 1e-06 | 0.76671 | 2.0956e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 1155000 | 1155000 | 2026-02-25 09:44 | 0.89044 | 1e-06 | 0.76289 | 1.6223e-06 | 1155 | 0.99584 | 0.7629 | 0.76289 | 0.76289 | 0.76289 | 0.76289
| 1000 | 1000 | 2026-02-25 09:44 | 0.89861 | 1e-06 | 0.76669 | 3.1919e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 1170000 | 1170000 | 2026-02-25 09:45 | 0.89024 | 1e-06 | 0.7629 | 3.1251e-06 | 1170 | 0.99558 | 0.76291 | 0.76291 | 0.7629 | 0.7629 | 0.7629
| 1000 | 1000 | 2026-02-25 09:45 | 0.89866 | 1e-06 | 0.76672 | 1.6928e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 1185000 | 1185000 | 2026-02-25 09:46 | 0.89009 | 1e-06 | 0.76292 | 2.4949e-06 | 1185 | 0.99541 | 0.76292 | 0.76292 | 0.76291 | 0.76291 | 0.76291
| 1000 | 1000 | 2026-02-25 09:46 | 0.89867 | 1e-06 | 0.76673 | 2.3802e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 1200000 | 1200000 | 2026-02-25 09:46 | 0.88989 | 1e-06 | 0.76283 | 3.1354e-06 | 1200 | 0.9953 | 0.76284 | 0.76283 | 0.76282 | 0.76282 | 0.76282
| 1000 | 1000 | 2026-02-25 09:47 | 0.89865 | 1e-06 | 0.76672 | 2.3376e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1215000 | 1215000 | 2026-02-25 09:47 | 0.88978 | 1e-06 | 0.76277 | 2.0654e-06 | 1215 | 0.99526 | 0.76278 | 0.76277 | 0.76277 | 0.76277 | 0.76277
| 1000 | 1000 | 2026-02-25 09:48 | 0.89866 | 1e-06 | 0.76673 | 2.7623e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 1230000 | 1230000 | 2026-02-25 09:48 | 0.88979 | 1e-06 | 0.76274 | 2.5483e-06 | 1230 | 0.99531 | 0.76275 | 0.76275 | 0.76274 | 0.76274 | 0.76274
| 1000 | 1000 | 2026-02-25 09:48 | 0.89863 | 1e-06 | 0.76671 | 1.3443e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1245000 | 1245000 | 2026-02-25 09:49 | 0.88988 | 1e-06 | 0.76274 | 1.5607e-06 | 1245 | 0.99541 | 0.76274 | 0.76274 | 0.76274 | 0.76274 | 0.76273
| 1000 | 1000 | 2026-02-25 09:49 | 0.89862 | 1e-06 | 0.7667 | 2.4476e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 1260000 | 1260000 | 2026-02-25 09:50 | 0.88993 | 1e-06 | 0.76264 | 1.9601e-06 | 1260 | 0.99559 | 0.76265 | 0.76264 | 0.76264 | 0.76264 | 0.76264
| 1000 | 1000 | 2026-02-25 09:50 | 0.89864 | 1e-06 | 0.76671 | 2.1689e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 1275000 | 1275000 | 2026-02-25 09:50 | 0.89022 | 1e-06 | 0.7627 | 3.2223e-06 | 1275 | 0.99583 | 0.76269 | 0.7627 | 0.7627 | 0.76271 | 0.76271
| 1000 | 1000 | 2026-02-25 09:51 | 0.89869 | 1e-06 | 0.76675 | 1.2677e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 1290000 | 1290000 | 2026-02-25 09:51 | 0.89042 | 1e-06 | 0.76263 | 1.9541e-06 | 1290 | 0.99615 | 0.76264 | 0.76263 | 0.76263 | 0.76263 | 0.76263
| 1000 | 1000 | 2026-02-25 09:52 | 0.8987 | 1e-06 | 0.76677 | 1.3193e-06 | #NAME? | #NAME? | 0.76677 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 1305000 | 1305000 | 2026-02-25 09:52 | 0.89076 | 1e-06 | 0.76264 | 2.3035e-06 | 1305 | 0.99651 | 0.76265 | 0.76264 | 0.76264 | 0.76264 | 0.76264
| 1000 | 1000 | 2026-02-25 09:52 | 0.8987 | 1e-06 | 0.76676 | 2.121e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 1320000 | 1320000 | 2026-02-25 09:53 | 0.89106 | 1e-06 | 0.76257 | 5.3117e-07 | 1320 | 0.99695 | 0.76257 | 0.76257 | 0.76257 | 0.76257 | 0.76257
| 1000 | 1000 | 2026-02-25 09:53 | 0.89869 | 1e-06 | 0.76675 | 1.8555e-06 | #NAME? | #NAME? | 0.76675 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 1335000 | 1335000 | 2026-02-25 09:54 | 0.89144 | 1e-06 | 0.76252 | 2.8201e-06 | 1335 | 0.99744 | 0.76253 | 0.76252 | 0.76252 | 0.76252 | 0.76251
| 1000 | 1000 | 2026-02-25 09:54 | 0.89867 | 1e-06 | 0.76674 | 3.966e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 1350000 | 1350000 | 2026-02-25 09:54 | 0.89186 | 1e-06 | 0.76247 | 2.6561e-06 | 1350 | 0.99797 | 0.76248 | 0.76248 | 0.76247 | 0.76247 | 0.76247
| 1000 | 1000 | 2026-02-25 09:55 | 0.89865 | 1e-06 | 0.76672 | 1.0724e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1365000 | 1365000 | 2026-02-25 09:55 | 0.89232 | 1e-06 | 0.76243 | 2.3061e-06 | 1365 | 0.99854 | 0.76243 | 0.76243 | 0.76242 | 0.76242 | 0.76242
| 1000 | 1000 | 2026-02-25 09:56 | 0.89871 | 1e-06 | 0.76676 | 3.0368e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 1380000 | 1380000 | 2026-02-25 09:56 | 0.89287 | 1e-06 | 0.76243 | 4.4716e-06 | 1380 | 0.99915 | 0.76245 | 0.76244 | 0.76243 | 0.76242 | 0.76243
| 1000 | 1000 | 2026-02-25 09:56 | 0.89868 | 1e-06 | 0.76674 | 1.6551e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 1395000 | 1395000 | 2026-02-25 09:57 | 0.89332 | 1e-06 | 0.76234 | 2.836e-06 | 1395 | 0.99979 | 0.76234 | 0.76234 | 0.76234 | 0.76233 | 0.76233
| 1000 | 1000 | 2026-02-25 09:57 | 0.89871 | 1e-06 | 0.76678 | 1.7746e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 1410000 | 1410000 | 2026-02-25 09:58 | 0.89389 | 1e-06 | 0.76232 | 6.1894e-07 | 1410 | 1.0004 | 0.76232 | 0.76232 | 0.76233 | 0.76232 | 0.76232
| 1000 | 1000 | 2026-02-25 09:58 | 0.89869 | 1e-06 | 0.76676 | 1.5172e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76675 | 0.76676 | 0.76676
| 1425000 | 1425000 | 2026-02-25 09:58 | 0.89438 | 1e-06 | 0.76226 | 4.408e-06 | 1425 | 1.0011 | 0.76227 | 0.76226 | 0.76226 | 0.76225 | 0.76225
| 1000 | 1000 | 2026-02-25 09:59 | 0.89864 | 1e-06 | 0.76671 | 1.1867e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 1440000 | 1440000 | 2026-02-25 09:59 | 0.89495 | 1e-06 | 0.76223 | 3.068e-06 | 1440 | 1.0017 | 0.76224 | 0.76224 | 0.76223 | 0.76223 | 0.76222
| 1000 | 1000 | 2026-02-25 10:00 | 0.89869 | 1e-06 | 0.76675 | 3.9054e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76677
| 1455000 | 1455000 | 2026-02-25 10:00 | 0.89547 | 1e-06 | 0.76219 | 4.5044e-06 | 1455 | 1.0024 | 0.7622 | 0.76219 | 0.76219 | 0.76218 | 0.76218
| 1000 | 1000 | 2026-02-25 10:00 | 0.89865 | 1e-06 | 0.76672 | 3.6982e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76673 | 0.76673
| 1470000 | 1470000 | 2026-02-25 10:01 | 0.89594 | 1e-06 | 0.7621 | 2.4553e-06 | 1470 | 1.003 | 0.7621 | 0.7621 | 0.7621 | 0.7621 | 0.76209
| 1000 | 1000 | 2026-02-25 10:01 | 0.89864 | 1e-06 | 0.7667 | 2.4082e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 1485000 | 1485000 | 2026-02-25 10:02 | 0.89649 | 1e-06 | 0.76211 | 2.4447e-06 | 1485 | 1.0036 | 0.76212 | 0.76211 | 0.76211 | 0.76211 | 0.76211
| 1000 | 1000 | 2026-02-25 10:02 | 0.89864 | 1e-06 | 0.76671 | 1.5097e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1500000 | 1500000 | 2026-02-25 10:02 | 0.89694 | 1e-06 | 0.76205 | 3.2866e-06 | 1500 | 1.0042 | 0.76206 | 0.76205 | 0.76205 | 0.76204 | 0.76204
| 1000 | 1000 | 2026-02-25 10:03 | 0.89863 | 1e-06 | 0.7667 | 2.1443e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 1515000 | 1515000 | 2026-02-25 10:03 | 0.89732 | 1e-06 | 0.76196 | 1.9983e-06 | 1515 | 1.0048 | 0.76197 | 0.76196 | 0.76196 | 0.76197 | 0.76196
| 1000 | 1000 | 2026-02-25 10:04 | 0.89861 | 1e-06 | 0.76669 | 1.5977e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 1530000 | 1530000 | 2026-02-25 10:04 | 0.89768 | 1e-06 | 0.7619 | 2.1318e-06 | 1530 | 1.0052 | 0.76191 | 0.7619 | 0.7619 | 0.7619 | 0.76189
| 1000 | 1000 | 2026-02-25 10:04 | 0.89864 | 1e-06 | 0.7667 | 3.4446e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 1545000 | 1545000 | 2026-02-25 10:05 | 0.89808 | 1e-06 | 0.76191 | 1.1212e-06 | 1545 | 1.0057 | 0.76191 | 0.76191 | 0.76191 | 0.7619 | 0.7619
| 1000 | 1000 | 2026-02-25 10:05 | 0.89869 | 1e-06 | 0.76675 | 2.811e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76676
| 1560000 | 1560000 | 2026-02-25 10:06 | 0.89843 | 1e-06 | 0.76191 | 8.0453e-07 | 1560 | 1.0061 | 0.76191 | 0.76191 | 0.7619 | 0.7619 | 0.76191
| 1000 | 1000 | 2026-02-25 10:06 | 0.89867 | 1e-06 | 0.76673 | 8.3407e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 1575000 | 1575000 | 2026-02-25 10:06 | 0.89857 | 1e-06 | 0.76178 | 1.518e-06 | 1575 | 1.0064 | 0.76178 | 0.76178 | 0.76178 | 0.76178 | 0.76177
| 1000 | 1000 | 2026-02-25 10:07 | 0.89863 | 1e-06 | 0.76669 | 1.8371e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.76669
| 1590000 | 1590000 | 2026-02-25 10:07 | 0.89874 | 1e-06 | 0.76174 | 2.4028e-06 | 1590 | 1.0066 | 0.76174 | 0.76174 | 0.76174 | 0.76174 | 0.76173
| 1000 | 1000 | 2026-02-25 10:08 | 0.89861 | 1e-06 | 0.76669 | 1.86e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 1605000 | 1605000 | 2026-02-25 10:08 | 0.89886 | 1e-06 | 0.76171 | 2.8184e-06 | 1605 | 1.0068 | 0.76172 | 0.76171 | 0.7617 | 0.7617 | 0.7617
| 1000 | 1000 | 2026-02-25 10:08 | 0.89864 | 1e-06 | 0.76671 | 3.6239e-06 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.76671 | 0.76672 | 0.76672
| 1620000 | 1620000 | 2026-02-25 10:09 | 0.89901 | 1e-06 | 0.76174 | 3.1745e-06 | 1620 | 1.0069 | 0.76175 | 0.76175 | 0.76175 | 0.76174 | 0.76174
| 1000 | 1000 | 2026-02-25 10:09 | 0.89868 | 1e-06 | 0.76674 | 1.7362e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 1635000 | 1635000 | 2026-02-25 10:10 | 0.89894 | 1e-06 | 0.76166 | 1.6042e-06 | 1635 | 1.007 | 0.76166 | 0.76165 | 0.76166 | 0.76165 | 0.76166
| 1000 | 1000 | 2026-02-25 10:10 | 0.89863 | 1e-06 | 0.7667 | 2.2256e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 1650000 | 1650000 | 2026-02-25 10:10 | 0.89888 | 1e-06 | 0.76164 | 1.75e-06 | 1650 | 1.0069 | 0.76164 | 0.76164 | 0.76164 | 0.76163 | 0.76163
| 1000 | 1000 | 2026-02-25 10:11 | 0.89865 | 1e-06 | 0.76672 | 1.2995e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1665000 | 1665000 | 2026-02-25 10:11 | 0.89866 | 1e-06 | 0.76154 | 2.1197e-06 | 1665 | 1.0068 | 0.76154 | 0.76154 | 0.76153 | 0.76153 | 0.76153
| 1000 | 1000 | 2026-02-25 10:12 | 0.89868 | 1e-06 | 0.7667 | 2.649e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 1680000 | 1680000 | 2026-02-25 10:12 | 0.89848 | 1e-06 | 0.76152 | 9.4227e-07 | 1680 | 1.0066 | 0.76152 | 0.76152 | 0.76152 | 0.76152 | 0.76152
| 1000 | 1000 | 2026-02-25 10:13 | 0.89866 | 1e-06 | 0.76672 | 8.9284e-07 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76673 | 0.76672
| 1695000 | 1695000 | 2026-02-25 10:13 | 0.89813 | 1e-06 | 0.76141 | 3.4619e-06 | 1695 | 1.0064 | 0.76142 | 0.76142 | 0.76141 | 0.7614 | 0.7614
| 1000 | 1000 | 2026-02-25 10:13 | 0.89864 | 1e-06 | 0.7667 | 2.4791e-06 | #NAME? | #NAME? | 0.7667 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 1710000 | 1710000 | 2026-02-25 10:14 | 0.89786 | 1e-06 | 0.76143 | 1.1761e-06 | 1710 | 1.0061 | 0.76143 | 0.76143 | 0.76143 | 0.76143 | 0.76144
| 1000 | 1000 | 2026-02-25 10:14 | 0.89869 | 1e-06 | 0.76675 | 1.162e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76675 | 0.76675 | 0.76675
| 1725000 | 1725000 | 2026-02-25 10:15 | 0.89745 | 1e-06 | 0.76136 | 3.1483e-06 | 1725 | 1.0057 | 0.76137 | 0.76136 | 0.76136 | 0.76136 | 0.76135
| 1000 | 1000 | 2026-02-25 10:15 | 0.89866 | 1e-06 | 0.76673 | 1.1661e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 1740000 | 1740000 | 2026-02-25 10:15 | 0.89699 | 1e-06 | 0.76131 | 2.4895e-06 | 1740 | 1.0052 | 0.76132 | 0.76131 | 0.7613 | 0.76131 | 0.7613
| 1000 | 1000 | 2026-02-25 10:16 | 0.89864 | 1e-06 | 0.76671 | 6.4673e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1755000 | 1755000 | 2026-02-25 10:16 | 0.89645 | 1e-06 | 0.76122 | 2.9521e-06 | 1755 | 1.0048 | 0.76122 | 0.76122 | 0.76121 | 0.76121 | 0.76121
| 1000 | 1000 | 2026-02-25 10:17 | 0.89862 | 1e-06 | 0.7667 | 1.5663e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76669 | 0.7667 | 0.7667
| 1770000 | 1770000 | 2026-02-25 10:17 | 0.89592 | 1e-06 | 0.76116 | 1.9027e-06 | 1770 | 1.0043 | 0.76116 | 0.76116 | 0.76116 | 0.76115 | 0.76115
| 1000 | 1000 | 2026-02-25 10:17 | 0.89863 | 1e-06 | 0.76671 | 3.8867e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 1785000 | 1785000 | 2026-02-25 10:18 | 0.89534 | 1e-06 | 0.7611 | 1.7556e-06 | 1785 | 1.0037 | 0.7611 | 0.7611 | 0.7611 | 0.7611 | 0.76109
| 1000 | 1000 | 2026-02-25 10:18 | 0.89871 | 1e-06 | 0.76678 | 3.7208e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76679
| 1800000 | 1800000 | 2026-02-25 10:19 | 0.89485 | 1e-06 | 0.76114 | 2.9812e-06 | 1800 | 1.0031 | 0.76115 | 0.76115 | 0.76114 | 0.76114 | 0.76113
| 1000 | 1000 | 2026-02-25 10:19 | 0.89868 | 1e-06 | 0.76675 | 1.8667e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76675 | 0.76675
| 1815000 | 1815000 | 2026-02-25 10:19 | 0.89408 | 1e-06 | 0.76096 | 2.1269e-06 | 1815 | 1.0025 | 0.76097 | 0.76096 | 0.76096 | 0.76095 | 0.76096
| 1000 | 1000 | 2026-02-25 10:20 | 0.89862 | 1e-06 | 0.7667 | 1.8588e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 1830000 | 1830000 | 2026-02-25 10:20 | 0.8936 | 1e-06 | 0.76104 | 9.1201e-07 | 1830 | 1.0018 | 0.76103 | 0.76104 | 0.76104 | 0.76104 | 0.76104
| 1000 | 1000 | 2026-02-25 10:21 | 0.89867 | 1e-06 | 0.76674 | 1.2522e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 1845000 | 1845000 | 2026-02-25 10:21 | 0.89286 | 1e-06 | 0.76089 | 2.4033e-06 | 1845 | 1.0012 | 0.76089 | 0.76089 | 0.76089 | 0.76088 | 0.76088
| 1000 | 1000 | 2026-02-25 10:21 | 0.89864 | 1e-06 | 0.76671 | 3.1332e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 1860000 | 1860000 | 2026-02-25 10:22 | 0.89231 | 1e-06 | 0.76091 | 6.4208e-07 | 1860 | 1.0005 | 0.76091 | 0.76091 | 0.76091 | 0.76091 | 0.76091
| 1000 | 1000 | 2026-02-25 10:22 | 0.89873 | 1e-06 | 0.76679 | 1.1728e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.76679
| 1875000 | 1875000 | 2026-02-25 10:23 | 0.89167 | 1e-06 | 0.76084 | 1.0366e-06 | 1875 | 0.99991 | 0.76084 | 0.76084 | 0.76084 | 0.76084 | 0.76083
| 1000 | 1000 | 2026-02-25 10:23 | 0.89866 | 1e-06 | 0.76673 | 2.0572e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 1890000 | 1890000 | 2026-02-25 10:23 | 0.89102 | 1e-06 | 0.76075 | 1.3776e-06 | 1890 | 0.9993 | 0.76075 | 0.76075 | 0.76075 | 0.76075 | 0.76074
| 1000 | 1000 | 2026-02-25 10:24 | 0.89866 | 1e-06 | 0.76674 | 7.8311e-07 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76673 | 0.76674
| 1905000 | 1905000 | 2026-02-25 10:24 | 0.8904 | 1e-06 | 0.76066 | 3.0216e-06 | 1905 | 0.99872 | 0.76066 | 0.76066 | 0.76066 | 0.76065 | 0.76065
| 1000 | 1000 | 2026-02-25 10:24 | 0.89865 | 1e-06 | 0.76672 | 1.3724e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 1920000 | 1920000 | 2026-02-25 10:25 | 0.88996 | 1e-06 | 0.76069 | 8.1496e-07 | 1920 | 0.99818 | 0.76069 | 0.76069 | 0.76069 | 0.76069 | 0.76069
| 1000 | 1000 | 2026-02-25 10:25 | 0.89868 | 1e-06 | 0.76675 | 9.5934e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 1935000 | 1935000 | 2026-02-25 10:26 | 0.88938 | 1e-06 | 0.76059 | 1.6009e-06 | 1935 | 0.99767 | 0.76059 | 0.76059 | 0.76059 | 0.76059 | 0.76058
| 1000 | 1000 | 2026-02-25 10:26 | 0.89867 | 1e-06 | 0.76674 | 6.395e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 1950000 | 1950000 | 2026-02-25 10:26 | 0.88896 | 1e-06 | 0.76056 | 1.2778e-06 | 1950 | 0.99723 | 0.76056 | 0.76056 | 0.76056 | 0.76056 | 0.76056
| 1000 | 1000 | 2026-02-25 10:27 | 0.89869 | 1e-06 | 0.76675 | 1.7794e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 1965000 | 1965000 | 2026-02-25 10:27 | 0.88861 | 1e-06 | 0.76057 | 2.3366e-06 | 1965 | 0.99683 | 0.76057 | 0.76057 | 0.76056 | 0.76056 | 0.76056
| 1000 | 1000 | 2026-02-25 10:28 | 0.89869 | 1e-06 | 0.76675 | 1.0847e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76675
| 1980000 | 1980000 | 2026-02-25 10:28 | 0.88816 | 1e-06 | 0.76044 | 3.8034e-06 | 1980 | 0.9965 | 0.76045 | 0.76044 | 0.76043 | 0.76043 | 0.76043
| 1000 | 1000 | 2026-02-25 10:28 | 0.89867 | 1e-06 | 0.76674 | 3.3565e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76675 | 0.76675
| 1995000 | 1995000 | 2026-02-25 10:29 | 0.88788 | 1e-06 | 0.7604 | 1.2412e-06 | 1995 | 0.99623 | 0.7604 | 0.7604 | 0.7604 | 0.7604 | 0.7604
| 1000 | 1000 | 2026-02-25 10:29 | 0.89865 | 1e-06 | 0.76672 | 1.5801e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2010000 | 2010000 | 2026-02-25 10:30 | 0.88762 | 1e-06 | 0.76033 | 8.4719e-07 | 2010 | 0.99603 | 0.76033 | 0.76033 | 0.76033 | 0.76032 | 0.76032
| 1000 | 1000 | 2026-02-25 10:30 | 0.89864 | 1e-06 | 0.76671 | 1.7958e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 2025000 | 2025000 | 2026-02-25 10:30 | 0.88749 | 1e-06 | 0.76031 | 2.4028e-06 | 2025 | 0.99592 | 0.76031 | 0.76031 | 0.76031 | 0.7603 | 0.7603
| 1000 | 1000 | 2026-02-25 10:31 | 0.89869 | 1e-06 | 0.76676 | 1.4481e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76675 | 0.76675
| 2040000 | 2040000 | 2026-02-25 10:31 | 0.88734 | 1e-06 | 0.76022 | 3.1225e-06 | 2040 | 0.99585 | 0.76023 | 0.76022 | 0.76022 | 0.76022 | 0.76021
| 1000 | 1000 | 2026-02-25 10:32 | 0.89865 | 1e-06 | 0.76672 | 1.4128e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 2055000 | 2055000 | 2026-02-25 10:32 | 0.88724 | 1e-06 | 0.76011 | 2.3328e-06 | 2055 | 0.99588 | 0.76012 | 0.76012 | 0.76011 | 0.76011 | 0.7601
| 1000 | 1000 | 2026-02-25 10:32 | 0.89863 | 1e-06 | 0.7667 | 1.4611e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 2070000 | 2070000 | 2026-02-25 10:33 | 0.88725 | 1e-06 | 0.76004 | 9.9637e-07 | 2070 | 0.99599 | 0.76004 | 0.76004 | 0.76004 | 0.76004 | 0.76004
| 1000 | 1000 | 2026-02-25 10:33 | 0.89865 | 1e-06 | 0.76672 | 1.6968e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2085000 | 2085000 | 2026-02-25 10:34 | 0.88736 | 1e-06 | 0.76 | 1.5915e-06 | 2085 | 0.99617 | 0.76001 | 0.76001 | 0.76001 | 0.76 | 0.76
| 1000 | 1000 | 2026-02-25 10:34 | 0.89866 | 1e-06 | 0.76673 | 1.2162e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 2100000 | 2100000 | 2026-02-25 10:34 | 0.8876 | 1e-06 | 0.76001 | 2.387e-06 | 2100 | 0.99641 | 0.76002 | 0.76001 | 0.76002 | 0.76001 | 0.76001
| 1000 | 1000 | 2026-02-25 10:35 | 0.89873 | 1e-06 | 0.76679 | 3.933e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76679 | 0.76679 | 0.7668
| 2115000 | 2115000 | 2026-02-25 10:35 | 0.88778 | 1e-06 | 0.75994 | 2.5966e-06 | 2115 | 0.99673 | 0.75995 | 0.75994 | 0.75994 | 0.75994 | 0.75993
| 1000 | 1000 | 2026-02-25 10:36 | 0.8987 | 1e-06 | 0.76677 | 8.5421e-07 | #NAME? | #NAME? | 0.76677 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 2130000 | 2130000 | 2026-02-25 10:36 | 0.88803 | 1e-06 | 0.75986 | 2.9572e-06 | 2130 | 0.99711 | 0.75987 | 0.75986 | 0.75986 | 0.75986 | 0.75985
| 1000 | 1000 | 2026-02-25 10:36 | 0.89866 | 1e-06 | 0.76673 | 1.2636e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76672 | 0.76673
| 2145000 | 2145000 | 2026-02-25 10:37 | 0.88838 | 1e-06 | 0.75983 | 1.3809e-06 | 2145 | 0.99754 | 0.75984 | 0.75983 | 0.75983 | 0.75983 | 0.75983
| 1000 | 1000 | 2026-02-25 10:37 | 0.89869 | 1e-06 | 0.76675 | 2.9929e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 2160000 | 2160000 | 2026-02-25 10:38 | 0.88878 | 1e-06 | 0.75979 | 3.6028e-06 | 2160 | 0.99803 | 0.7598 | 0.7598 | 0.75979 | 0.75978 | 0.75978
| 1000 | 1000 | 2026-02-25 10:38 | 0.89867 | 1e-06 | 0.76673 | 4.5627e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2175000 | 2175000 | 2026-02-25 10:38 | 0.88907 | 1e-06 | 0.75963 | 2.6018e-06 | 2175 | 0.99856 | 0.75964 | 0.75963 | 0.75963 | 0.75963 | 0.75962
| 1000 | 1000 | 2026-02-25 10:39 | 0.89865 | 1e-06 | 0.76671 | 3.1357e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 2190000 | 2190000 | 2026-02-25 10:39 | 0.88952 | 1e-06 | 0.75957 | 1.3589e-06 | 2190 | 0.99916 | 0.75957 | 0.75957 | 0.75957 | 0.75957 | 0.75957
| 1000 | 1000 | 2026-02-25 10:40 | 0.89866 | 1e-06 | 0.76673 | 2.5963e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 2205000 | 2205000 | 2026-02-25 10:40 | 0.89003 | 1e-06 | 0.75954 | 3.0305e-06 | 2205 | 0.99976 | 0.75955 | 0.75954 | 0.75954 | 0.75953 | 0.75953
| 1000 | 1000 | 2026-02-25 10:40 | 0.89865 | 1e-06 | 0.76672 | 8.9919e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2220000 | 2220000 | 2026-02-25 10:41 | 0.89051 | 1e-06 | 0.75945 | 3.1235e-06 | 2220 | 1.0004 | 0.75946 | 0.75946 | 0.75946 | 0.75945 | 0.75944
| 1000 | 1000 | 2026-02-25 10:41 | 0.89866 | 1e-06 | 0.76672 | 3.5964e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76672 | 0.76673
| 2235000 | 2235000 | 2026-02-25 10:42 | 0.89101 | 1e-06 | 0.7594 | 1.221e-06 | 2235 | 1.0011 | 0.7594 | 0.7594 | 0.75939 | 0.75939 | 0.75939
| 1000 | 1000 | 2026-02-25 10:42 | 0.89868 | 1e-06 | 0.76674 | 2.1495e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76674 | 0.76675
| 2250000 | 2250000 | 2026-02-25 10:42 | 0.89153 | 1e-06 | 0.75932 | 1.6537e-06 | 2250 | 1.0017 | 0.75933 | 0.75933 | 0.75932 | 0.75932 | 0.75932
| 1000 | 1000 | 2026-02-25 10:43 | 0.89866 | 1e-06 | 0.76673 | 2.5456e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 2265000 | 2265000 | 2026-02-25 10:43 | 0.89202 | 1e-06 | 0.75923 | 1.8636e-06 | 2265 | 1.0024 | 0.75923 | 0.75923 | 0.75923 | 0.75923 | 0.75922
| 1000 | 1000 | 2026-02-25 10:44 | 0.89865 | 1e-06 | 0.76672 | 1.594e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 2280000 | 2280000 | 2026-02-25 10:44 | 0.89249 | 1e-06 | 0.75914 | 2.0541e-06 | 2280 | 1.0031 | 0.75914 | 0.75914 | 0.75914 | 0.75914 | 0.75913
| 1000 | 1000 | 2026-02-25 10:44 | 0.89864 | 1e-06 | 0.76672 | 1.8115e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2295000 | 2295000 | 2026-02-25 10:45 | 0.8929 | 1e-06 | 0.759 | 3.1022e-06 | 2295 | 1.0037 | 0.75901 | 0.75901 | 0.759 | 0.759 | 0.75899
| 1000 | 1000 | 2026-02-25 10:45 | 0.89863 | 1e-06 | 0.7667 | 1.5806e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 2310000 | 2310000 | 2026-02-25 10:46 | 0.89342 | 1e-06 | 0.75897 | 4.0206e-06 | 2310 | 1.0043 | 0.75898 | 0.75898 | 0.75898 | 0.75897 | 0.75896
| 1000 | 1000 | 2026-02-25 10:46 | 0.89863 | 1e-06 | 0.7667 | 5.6814e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.7667 | 0.76671 | 0.76672
| 2325000 | 2325000 | 2026-02-25 10:46 | 0.8938 | 1e-06 | 0.75886 | 7.774e-07 | 2325 | 1.0049 | 0.75886 | 0.75886 | 0.75886 | 0.75886 | 0.75886
| 1000 | 1000 | 2026-02-25 10:47 | 0.89861 | 1e-06 | 0.76669 | 3.9822e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.7667 | 0.7667
| 2340000 | 2340000 | 2026-02-25 10:47 | 0.89427 | 1e-06 | 0.75885 | 1.4385e-06 | 2340 | 1.0054 | 0.75886 | 0.75886 | 0.75885 | 0.75885 | 0.75885
| 1000 | 1000 | 2026-02-25 10:48 | 0.89864 | 1e-06 | 0.76671 | 2.5555e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 2355000 | 2355000 | 2026-02-25 10:48 | 0.89462 | 1e-06 | 0.75879 | 2.7752e-06 | 2355 | 1.0059 | 0.75879 | 0.75879 | 0.75878 | 0.75878 | 0.75878
| 1000 | 1000 | 2026-02-25 10:48 | 0.89862 | 1e-06 | 0.76669 | 1.8309e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 2370000 | 2370000 | 2026-02-25 10:49 | 0.89493 | 1e-06 | 0.75872 | 2.3546e-06 | 2370 | 1.0064 | 0.75873 | 0.75873 | 0.75872 | 0.75871 | 0.75872
| 1000 | 1000 | 2026-02-25 10:49 | 0.89863 | 1e-06 | 0.7667 | 1.0666e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 2385000 | 2385000 | 2026-02-25 10:50 | 0.89512 | 1e-06 | 0.75861 | 2.348e-06 | 2385 | 1.0067 | 0.75862 | 0.75861 | 0.75861 | 0.75861 | 0.7586
| 1000 | 1000 | 2026-02-25 10:50 | 0.8986 | 1e-06 | 0.76668 | 1.7913e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76668
| 2400000 | 2400000 | 2026-02-25 10:50 | 0.89533 | 1e-06 | 0.75855 | 3.0136e-06 | 2400 | 1.007 | 0.75856 | 0.75855 | 0.75855 | 0.75854 | 0.75854
| 1000 | 1000 | 2026-02-25 10:51 | 0.89868 | 1e-06 | 0.76675 | 3.9554e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 2415000 | 2415000 | 2026-02-25 10:51 | 0.89552 | 1e-06 | 0.75853 | 4.601e-06 | 2415 | 1.0073 | 0.75854 | 0.75853 | 0.75853 | 0.75852 | 0.75851
| 1000 | 1000 | 2026-02-25 10:52 | 0.89862 | 1e-06 | 0.7667 | 2.1826e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 2430000 | 2430000 | 2026-02-25 10:52 | 0.89547 | 1e-06 | 0.75837 | 3.0057e-06 | 2430 | 1.0074 | 0.75838 | 0.75838 | 0.75837 | 0.75837 | 0.75836
| 1000 | 1000 | 2026-02-25 10:52 | 0.8986 | 1e-06 | 0.76667 | 2.2016e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76668
| 2445000 | 2445000 | 2026-02-25 10:53 | 0.89551 | 1e-06 | 0.75833 | 1.7794e-06 | 2445 | 1.0075 | 0.75834 | 0.75833 | 0.75833 | 0.75833 | 0.75833
| 1000 | 1000 | 2026-02-25 10:53 | 0.89861 | 1e-06 | 0.76669 | 2.5083e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.7667 | 0.76669
| 2460000 | 2460000 | 2026-02-25 10:54 | 0.89545 | 1e-06 | 0.75827 | 9.3218e-07 | 2460 | 1.0075 | 0.75827 | 0.75827 | 0.75827 | 0.75827 | 0.75827
| 1000 | 1000 | 2026-02-25 10:54 | 0.89861 | 1e-06 | 0.76668 | 2.4735e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 2475000 | 2475000 | 2026-02-25 10:54 | 0.8954 | 1e-06 | 0.75828 | 3.2939e-06 | 2475 | 1.0075 | 0.75829 | 0.75828 | 0.75828 | 0.75828 | 0.75827
| 1000 | 1000 | 2026-02-25 10:55 | 0.89864 | 1e-06 | 0.7667 | 4.9833e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.76671 | 0.76672
| 2490000 | 2490000 | 2026-02-25 10:55 | 0.8952 | 1e-06 | 0.75822 | 1.2348e-06 | 2490 | 1.0073 | 0.75822 | 0.75822 | 0.75822 | 0.75822 | 0.75822
| 1000 | 1000 | 2026-02-25 10:56 | 0.89872 | 1e-06 | 0.76678 | 1.71e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76679 | 0.76678
| 2505000 | 2505000 | 2026-02-25 10:56 | 0.89486 | 1e-06 | 0.7581 | 2.8201e-06 | 2505 | 1.0071 | 0.7581 | 0.7581 | 0.7581 | 0.75809 | 0.75809
| 1000 | 1000 | 2026-02-25 10:56 | 0.89865 | 1e-06 | 0.76672 | 1.6493e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2520000 | 2520000 | 2026-02-25 10:57 | 0.89452 | 1e-06 | 0.75802 | 1.1682e-06 | 2520 | 1.0068 | 0.75803 | 0.75802 | 0.75802 | 0.75802 | 0.75803
| 1000 | 1000 | 2026-02-25 10:57 | 0.89863 | 1e-06 | 0.76671 | 3.6015e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 2535000 | 2535000 | 2026-02-25 10:58 | 0.89411 | 1e-06 | 0.75794 | 7.5765e-07 | 2535 | 1.0065 | 0.75795 | 0.75794 | 0.75795 | 0.75794 | 0.75794
| 1000 | 1000 | 2026-02-25 10:58 | 0.89864 | 1e-06 | 0.76671 | 3.0431e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 2550000 | 2550000 | 2026-02-25 10:58 | 0.8937 | 1e-06 | 0.7579 | 1.0143e-06 | 2550 | 1.0061 | 0.75791 | 0.7579 | 0.7579 | 0.7579 | 0.7579
| 1000 | 1000 | 2026-02-25 10:59 | 0.89873 | 1e-06 | 0.76678 | 3.9655e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76679 | 0.76679
| 2565000 | 2565000 | 2026-02-25 10:59 | 0.89323 | 1e-06 | 0.75786 | 5.2903e-06 | 2565 | 1.0056 | 0.75787 | 0.75786 | 0.75786 | 0.75785 | 0.75784
| 1000 | 1000 | 2026-02-25 10:59 | 0.89865 | 1e-06 | 0.76672 | 2.042e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 2580000 | 2580000 | 2026-02-25 11:00 | 0.89273 | 1e-06 | 0.75784 | 7.6361e-07 | 2580 | 1.0051 | 0.75783 | 0.75784 | 0.75784 | 0.75784 | 0.75784
| 1000 | 1000 | 2026-02-25 11:00 | 0.89866 | 1e-06 | 0.76673 | 3.013e-06 | #NAME? | #NAME? | 0.76671 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2595000 | 2595000 | 2026-02-25 11:01 | 0.8921 | 1e-06 | 0.75772 | 2.273e-06 | 2595 | 1.0045 | 0.75772 | 0.75772 | 0.75772 | 0.75772 | 0.75771
| 1000 | 1000 | 2026-02-25 11:01 | 0.89865 | 1e-06 | 0.76671 | 1.8581e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 2610000 | 2610000 | 2026-02-25 11:01 | 0.89153 | 1e-06 | 0.7577 | 1.3042e-06 | 2610 | 1.0039 | 0.7577 | 0.7577 | 0.7577 | 0.7577 | 0.7577
| 1000 | 1000 | 2026-02-25 11:02 | 0.89865 | 1e-06 | 0.76672 | 2.6263e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76671 | 0.76672 | 0.76672
| 2625000 | 2625000 | 2026-02-25 11:02 | 0.89085 | 1e-06 | 0.75759 | 1.1785e-06 | 2625 | 1.0033 | 0.75759 | 0.75759 | 0.75759 | 0.75759 | 0.75759
| 1000 | 1000 | 2026-02-25 11:03 | 0.89862 | 1e-06 | 0.7667 | 1.6367e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76669 | 0.7667 | 0.7667
| 2640000 | 2640000 | 2026-02-25 11:03 | 0.89026 | 1e-06 | 0.75758 | 1.0711e-06 | 2640 | 1.0026 | 0.75757 | 0.75758 | 0.75758 | 0.75758 | 0.75758
| 1000 | 1000 | 2026-02-25 11:03 | 0.89865 | 1e-06 | 0.76672 | 2.0259e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2655000 | 2655000 | 2026-02-25 11:04 | 0.88954 | 1e-06 | 0.75746 | 2.8688e-06 | 2655 | 1.002 | 0.75747 | 0.75747 | 0.75746 | 0.75746 | 0.75746
| 1000 | 1000 | 2026-02-25 11:04 | 0.8986 | 1e-06 | 0.76667 | 2.7248e-06 | #NAME? | #NAME? | 0.76666 | 0.76667 | 0.76668 | 0.76668 | 0.76668
| 2670000 | 2670000 | 2026-02-25 11:05 | 0.88889 | 1e-06 | 0.75739 | 1.9976e-06 | 2670 | 1.0013 | 0.7574 | 0.7574 | 0.75739 | 0.75739 | 0.75739
| 1000 | 1000 | 2026-02-25 11:05 | 0.89859 | 1e-06 | 0.76667 | 2.0132e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76668
| 2685000 | 2685000 | 2026-02-25 11:05 | 0.88832 | 1e-06 | 0.7574 | 3.1341e-06 | 2685 | 1.0007 | 0.75741 | 0.7574 | 0.7574 | 0.7574 | 0.75739
| 1000 | 1000 | 2026-02-25 11:06 | 0.89862 | 1e-06 | 0.76669 | 2.4648e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 2700000 | 2700000 | 2026-02-25 11:06 | 0.88768 | 1e-06 | 0.75733 | 2.4894e-06 | 2700 | 1 | 0.75734 | 0.75733 | 0.75733 | 0.75733 | 0.75733
| 1000 | 1000 | 2026-02-25 11:07 | 0.89865 | 1e-06 | 0.76673 | 1.17e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2715000 | 2715000 | 2026-02-25 11:07 | 0.8872 | 1e-06 | 0.75737 | 1.9344e-06 | 2715 | 0.99944 | 0.75737 | 0.75738 | 0.75737 | 0.75737 | 0.75737
| 1000 | 1000 | 2026-02-25 11:07 | 0.89864 | 1e-06 | 0.76671 | 2.914e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2730000 | 2730000 | 2026-02-25 11:08 | 0.88663 | 1e-06 | 0.75731 | 2.5932e-06 | 2730 | 0.99888 | 0.75732 | 0.75731 | 0.75731 | 0.75731 | 0.75731
| 1000 | 1000 | 2026-02-25 11:08 | 0.89866 | 1e-06 | 0.76673 | 2.5688e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 2745000 | 2745000 | 2026-02-25 11:09 | 0.88623 | 1e-06 | 0.75736 | 2.6292e-06 | 2745 | 0.99837 | 0.75737 | 0.75736 | 0.75736 | 0.75736 | 0.75735
| 1000 | 1000 | 2026-02-25 11:09 | 0.89868 | 1e-06 | 0.76674 | 2.0997e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 2760000 | 2760000 | 2026-02-25 11:09 | 0.88566 | 1e-06 | 0.75723 | 1.1675e-06 | 2760 | 0.99789 | 0.75724 | 0.75724 | 0.75723 | 0.75723 | 0.75723
| 1000 | 1000 | 2026-02-25 11:10 | 0.89866 | 1e-06 | 0.76673 | 2.3846e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 2775000 | 2775000 | 2026-02-25 11:10 | 0.88519 | 1e-06 | 0.75715 | 4.0277e-06 | 2775 | 0.99748 | 0.75716 | 0.75716 | 0.75715 | 0.75714 | 0.75714
| 1000 | 1000 | 2026-02-25 11:11 | 0.89865 | 1e-06 | 0.76672 | 3.8335e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76673 | 0.76673
| 2790000 | 2790000 | 2026-02-25 11:11 | 0.88488 | 1e-06 | 0.75713 | 3.8681e-06 | 2790 | 0.99714 | 0.75714 | 0.75714 | 0.75713 | 0.75713 | 0.75712
| 1000 | 1000 | 2026-02-25 11:11 | 0.89864 | 1e-06 | 0.76671 | 3.5205e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2805000 | 2805000 | 2026-02-25 11:12 | 0.88462 | 1e-06 | 0.75712 | 9.4583e-07 | 2805 | 0.99687 | 0.75711 | 0.75712 | 0.75712 | 0.75712 | 0.75712
| 1000 | 1000 | 2026-02-25 11:12 | 0.89867 | 1e-06 | 0.76674 | 2.2974e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 2820000 | 2820000 | 2026-02-25 11:13 | 0.88441 | 1e-06 | 0.7571 | 1.3498e-06 | 2820 | 0.99667 | 0.7571 | 0.7571 | 0.75709 | 0.7571 | 0.75709
| 1000 | 1000 | 2026-02-25 11:13 | 0.89865 | 1e-06 | 0.76673 | 1.7772e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2835000 | 2835000 | 2026-02-25 11:13 | 0.88422 | 1e-06 | 0.75702 | 1.4927e-06 | 2835 | 0.99654 | 0.75703 | 0.75702 | 0.75702 | 0.75702 | 0.75703
| 1000 | 1000 | 2026-02-25 11:14 | 0.89867 | 1e-06 | 0.76673 | 1.4744e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2850000 | 2850000 | 2026-02-25 11:14 | 0.88409 | 1e-06 | 0.75694 | 3.9157e-07 | 2850 | 0.9965 | 0.75694 | 0.75695 | 0.75694 | 0.75694 | 0.75694
| 1000 | 1000 | 2026-02-25 11:15 | 0.89862 | 1e-06 | 0.76669 | 1.3171e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 2865000 | 2865000 | 2026-02-25 11:15 | 0.88407 | 1e-06 | 0.75689 | 1.3103e-06 | 2865 | 0.99654 | 0.75689 | 0.75689 | 0.75689 | 0.7569 | 0.7569
| 1000 | 1000 | 2026-02-25 11:15 | 0.89869 | 1e-06 | 0.76674 | 1.5477e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 2880000 | 2880000 | 2026-02-25 11:16 | 0.88429 | 1e-06 | 0.757 | 1.0075e-06 | 2880 | 0.99666 | 0.75699 | 0.757 | 0.757 | 0.75699 | 0.75699
| 1000 | 1000 | 2026-02-25 11:16 | 0.8987 | 1e-06 | 0.76677 | 1.0107e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 2895000 | 2895000 | 2026-02-25 11:17 | 0.88428 | 1e-06 | 0.75684 | 2.7657e-06 | 2895 | 0.99685 | 0.75685 | 0.75684 | 0.75684 | 0.75684 | 0.75684
| 1000 | 1000 | 2026-02-25 11:17 | 0.89863 | 1e-06 | 0.7667 | 2.5166e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 2910000 | 2910000 | 2026-02-25 11:17 | 0.88449 | 1e-06 | 0.75682 | 2.7679e-06 | 2910 | 0.99712 | 0.75682 | 0.75682 | 0.75681 | 0.75681 | 0.75681
| 1000 | 1000 | 2026-02-25 11:18 | 0.89865 | 1e-06 | 0.76672 | 2.2115e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2925000 | 2925000 | 2026-02-25 11:18 | 0.88479 | 1e-06 | 0.7568 | 1.1807e-06 | 2925 | 0.99748 | 0.7568 | 0.7568 | 0.7568 | 0.7568 | 0.7568
| 1000 | 1000 | 2026-02-25 11:19 | 0.89866 | 1e-06 | 0.76673 | 1.3116e-06 | #NAME? | #NAME? | 0.76673 | 0.76672 | 0.76673 | 0.76672 | 0.76673
| 2940000 | 2940000 | 2026-02-25 11:19 | 0.88506 | 1e-06 | 0.75672 | 2.8557e-06 | 2940 | 0.99788 | 0.75673 | 0.75673 | 0.75672 | 0.75672 | 0.75672
| 1000 | 1000 | 2026-02-25 11:19 | 0.89865 | 1e-06 | 0.76672 | 1.6862e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 2955000 | 2955000 | 2026-02-25 11:20 | 0.88554 | 1e-06 | 0.75676 | 8.9698e-07 | 2955 | 0.99837 | 0.75677 | 0.75677 | 0.75677 | 0.75676 | 0.75676
| 1000 | 1000 | 2026-02-25 11:20 | 0.89872 | 1e-06 | 0.76678 | 2.7884e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76679
| 2970000 | 2970000 | 2026-02-25 11:21 | 0.8859 | 1e-06 | 0.75667 | 3.8099e-06 | 2970 | 0.99891 | 0.75668 | 0.75668 | 0.75667 | 0.75667 | 0.75666
| 1000 | 1000 | 2026-02-25 11:21 | 0.89867 | 1e-06 | 0.76674 | 2.379e-06 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 2985000 | 2985000 | 2026-02-25 11:21 | 0.88635 | 1e-06 | 0.7566 | 2.8985e-06 | 2985 | 0.9995 | 0.75661 | 0.7566 | 0.7566 | 0.7566 | 0.7566
| 1000 | 1000 | 2026-02-25 11:22 | 0.89867 | 1e-06 | 0.76674 | 1.0053e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 3000000 | 3000000 | 2026-02-25 11:22 | 0.88689 | 1e-06 | 0.75659 | 2.2113e-06 | 3000 | 1.0001 | 0.75659 | 0.75659 | 0.75659 | 0.75658 | 0.75658
| 1000 | 1000 | 2026-02-25 11:23 | 0.89865 | 1e-06 | 0.76673 | 6.4871e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 3015000 | 3015000 | 2026-02-25 11:23 | 0.88744 | 1e-06 | 0.75656 | 3.8138e-06 | 3015 | 1.0008 | 0.75657 | 0.75656 | 0.75656 | 0.75655 | 0.75654
| 1000 | 1000 | 2026-02-25 11:23 | 0.89867 | 1e-06 | 0.76674 | 7.7331e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76674
| 3030000 | 3030000 | 2026-02-25 11:24 | 0.88793 | 1e-06 | 0.75646 | 1.2769e-06 | 3030 | 1.0015 | 0.75646 | 0.75646 | 0.75646 | 0.75646 | 0.75645
| 1000 | 1000 | 2026-02-25 11:24 | 0.89868 | 1e-06 | 0.76674 | 2.8192e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 3045000 | 3045000 | 2026-02-25 11:25 | 0.88841 | 1e-06 | 0.75634 | 2.5995e-06 | 3045 | 1.0022 | 0.75635 | 0.75635 | 0.75634 | 0.75634 | 0.75634
| 1000 | 1000 | 2026-02-25 11:25 | 0.89864 | 1e-06 | 0.76672 | 1.7186e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76671 | 0.76672 | 0.76672
| 3060000 | 3060000 | 2026-02-25 11:25 | 0.88897 | 1e-06 | 0.75628 | 3.2388e-06 | 3060 | 1.0029 | 0.75629 | 0.75629 | 0.75629 | 0.75628 | 0.75627
| 1000 | 1000 | 2026-02-25 11:26 | 0.89862 | 1e-06 | 0.7667 | 1.0818e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 3075000 | 3075000 | 2026-02-25 11:26 | 0.88952 | 1e-06 | 0.75621 | 1.0214e-06 | 3075 | 1.0036 | 0.75622 | 0.75621 | 0.75622 | 0.75622 | 0.75621
| 1000 | 1000 | 2026-02-25 11:27 | 0.89865 | 1e-06 | 0.76672 | 2.8988e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76673 | 0.76672
| 3090000 | 3090000 | 2026-02-25 11:27 | 0.88999 | 1e-06 | 0.75611 | 3.7432e-06 | 3090 | 1.0043 | 0.75612 | 0.75611 | 0.7561 | 0.7561 | 0.7561
| 1000 | 1000 | 2026-02-25 11:27 | 0.89865 | 1e-06 | 0.76671 | 4.6715e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76672 | 0.76673
| 3105000 | 3105000 | 2026-02-25 11:28 | 0.89063 | 1e-06 | 0.75615 | 1.9082e-06 | 3105 | 1.0049 | 0.75615 | 0.75615 | 0.75615 | 0.75615 | 0.75616
| 1000 | 1000 | 2026-02-25 11:28 | 0.89872 | 1e-06 | 0.76678 | 2.8075e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76679 | 0.76679
| 3120000 | 3120000 | 2026-02-25 11:29 | 0.8911 | 1e-06 | 0.75607 | 3.3484e-06 | 3120 | 1.0056 | 0.75608 | 0.75608 | 0.75607 | 0.75607 | 0.75607
| 1000 | 1000 | 2026-02-25 11:29 | 0.89867 | 1e-06 | 0.76673 | 3.9188e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 3135000 | 3135000 | 2026-02-25 11:29 | 0.89151 | 1e-06 | 0.75599 | 2.756e-06 | 3135 | 1.0061 | 0.75599 | 0.75599 | 0.75599 | 0.75598 | 0.75598
| 1000 | 1000 | 2026-02-25 11:30 | 0.89865 | 1e-06 | 0.76671 | 1.6768e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 3150000 | 3150000 | 2026-02-25 11:30 | 0.89182 | 1e-06 | 0.75585 | 9.843e-07 | 3150 | 1.0067 | 0.75586 | 0.75585 | 0.75585 | 0.75585 | 0.75585
| 1000 | 1000 | 2026-02-25 11:31 | 0.89862 | 1e-06 | 0.7667 | 2.9248e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 3165000 | 3165000 | 2026-02-25 11:31 | 0.89213 | 1e-06 | 0.75577 | 3.3304e-06 | 3165 | 1.0071 | 0.75578 | 0.75578 | 0.75577 | 0.75577 | 0.75576
| 1000 | 1000 | 2026-02-25 11:31 | 0.89866 | 1e-06 | 0.76672 | 2.7917e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 3180000 | 3180000 | 2026-02-25 11:32 | 0.89242 | 1e-06 | 0.75572 | 4.0944e-06 | 3180 | 1.0075 | 0.75573 | 0.75573 | 0.75573 | 0.75572 | 0.75571
| 1000 | 1000 | 2026-02-25 11:32 | 0.89863 | 1e-06 | 0.76669 | 2.1849e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 3195000 | 3195000 | 2026-02-25 11:33 | 0.89261 | 1e-06 | 0.75564 | 3.2696e-06 | 3195 | 1.0078 | 0.75565 | 0.75565 | 0.75564 | 0.75564 | 0.75563
| 1000 | 1000 | 2026-02-25 11:33 | 0.89869 | 1e-06 | 0.76675 | 9.6828e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76675
| 3210000 | 3210000 | 2026-02-25 11:33 | 0.89276 | 1e-06 | 0.7556 | 3.1702e-06 | 3210 | 1.0081 | 0.75561 | 0.7556 | 0.7556 | 0.75559 | 0.75559
| 1000 | 1000 | 2026-02-25 11:34 | 0.89861 | 1e-06 | 0.76669 | 1.8705e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 3225000 | 3225000 | 2026-02-25 11:34 | 0.8928 | 1e-06 | 0.75551 | 2.4011e-06 | 3225 | 1.0082 | 0.75552 | 0.75551 | 0.75551 | 0.75551 | 0.7555
| 1000 | 1000 | 2026-02-25 11:35 | 0.8986 | 1e-06 | 0.76667 | 3.274e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76668
| 3240000 | 3240000 | 2026-02-25 11:35 | 0.89276 | 1e-06 | 0.75543 | 8.5462e-07 | 3240 | 1.0083 | 0.75543 | 0.75543 | 0.75543 | 0.75543 | 0.75543
| 1000 | 1000 | 2026-02-25 11:35 | 0.89862 | 1e-06 | 0.76669 | 5.2152e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.7667 | 0.7667
| 3255000 | 3255000 | 2026-02-25 11:36 | 0.8926 | 1e-06 | 0.75532 | 2.0417e-06 | 3255 | 1.0083 | 0.75532 | 0.75532 | 0.75532 | 0.75531 | 0.75531
| 1000 | 1000 | 2026-02-25 11:36 | 0.89863 | 1e-06 | 0.76671 | 2.0435e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 3270000 | 3270000 | 2026-02-25 11:37 | 0.89247 | 1e-06 | 0.75528 | 2.0769e-06 | 3270 | 1.0082 | 0.75529 | 0.75529 | 0.75528 | 0.75528 | 0.75528
| 1000 | 1000 | 2026-02-25 11:37 | 0.8987 | 1e-06 | 0.76677 | 7.9294e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 3285000 | 3285000 | 2026-02-25 11:37 | 0.89219 | 1e-06 | 0.7552 | 3.1746e-06 | 3285 | 1.0079 | 0.75521 | 0.75521 | 0.7552 | 0.7552 | 0.7552
| 1000 | 1000 | 2026-02-25 11:38 | 0.89866 | 1e-06 | 0.76672 | 1.1381e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 3300000 | 3300000 | 2026-02-25 11:38 | 0.89186 | 1e-06 | 0.75513 | 1.8246e-06 | 3300 | 1.0077 | 0.75514 | 0.75514 | 0.75514 | 0.75513 | 0.75513
| 1000 | 1000 | 2026-02-25 11:38 | 0.89863 | 1e-06 | 0.76671 | 2.5887e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.7667 | 0.76672
| 3315000 | 3315000 | 2026-02-25 11:39 | 0.89146 | 1e-06 | 0.75507 | 9.3165e-07 | 3315 | 1.0073 | 0.75508 | 0.75507 | 0.75507 | 0.75507 | 0.75508
| 1000 | 1000 | 2026-02-25 11:39 | 0.89875 | 1e-06 | 0.76681 | 4.7038e-07 | #NAME? | #NAME? | 0.76682 | 0.76681 | 0.76682 | 0.76682 | 0.76681
| 3330000 | 3330000 | 2026-02-25 11:40 | 0.891 | 1e-06 | 0.75501 | 3.6519e-06 | 3330 | 1.0068 | 0.75503 | 0.75501 | 0.75501 | 0.75501 | 0.75501
| 1000 | 1000 | 2026-02-25 11:40 | 0.89875 | 1e-06 | 0.7668 | 5.7402e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.7668 | 0.76681 | 0.76682
| 3345000 | 3345000 | 2026-02-25 11:40 | 0.89041 | 1e-06 | 0.75489 | 1.7167e-06 | 3345 | 1.0064 | 0.7549 | 0.75489 | 0.75489 | 0.75489 | 0.75489
| 1000 | 1000 | 2026-02-25 11:41 | 0.89868 | 1e-06 | 0.76675 | 1.9732e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76674 | 0.76675
| 3360000 | 3360000 | 2026-02-25 11:41 | 0.88975 | 1e-06 | 0.75477 | 3.4545e-06 | 3360 | 1.0058 | 0.75478 | 0.75477 | 0.75477 | 0.75477 | 0.75476
| 1000 | 1000 | 2026-02-25 11:42 | 0.89865 | 1e-06 | 0.76673 | 1.2162e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76672 | 0.76672 | 0.76672
| 3375000 | 3375000 | 2026-02-25 11:42 | 0.88901 | 1e-06 | 0.75461 | 2.4656e-06 | 3375 | 1.0051 | 0.75462 | 0.75462 | 0.75461 | 0.75461 | 0.75461
| 1000 | 1000 | 2026-02-25 11:42 | 0.89864 | 1e-06 | 0.76671 | 4.225e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 3390000 | 3390000 | 2026-02-25 11:43 | 0.88836 | 1e-06 | 0.75457 | 3.1455e-06 | 3390 | 1.0045 | 0.75457 | 0.75457 | 0.75457 | 0.75456 | 0.75456
| 1000 | 1000 | 2026-02-25 11:43 | 0.89865 | 1e-06 | 0.76672 | 1.5516e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76672
| 3405000 | 3405000 | 2026-02-25 11:44 | 0.88763 | 1e-06 | 0.75447 | 1.732e-06 | 3405 | 1.0038 | 0.75447 | 0.75447 | 0.75446 | 0.75446 | 0.75446
| 1000 | 1000 | 2026-02-25 11:44 | 0.8987 | 1e-06 | 0.76676 | 2.5786e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76677 | 0.76676
| 3420000 | 3420000 | 2026-02-25 11:44 | 0.88683 | 1e-06 | 0.75434 | 1.8865e-06 | 3420 | 1.003 | 0.75434 | 0.75434 | 0.75434 | 0.75434 | 0.75433
| 1000 | 1000 | 2026-02-25 11:45 | 0.89864 | 1e-06 | 0.76671 | 2.1841e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3435000 | 3435000 | 2026-02-25 11:45 | 0.88603 | 1e-06 | 0.75421 | 1.5942e-06 | 3435 | 1.0023 | 0.75421 | 0.75421 | 0.75421 | 0.7542 | 0.7542
| 1000 | 1000 | 2026-02-25 11:46 | 0.89863 | 1e-06 | 0.76671 | 2.0336e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3450000 | 3450000 | 2026-02-25 11:46 | 0.88535 | 1e-06 | 0.75417 | 1.3255e-06 | 3450 | 1.0016 | 0.75417 | 0.75417 | 0.75417 | 0.75418 | 0.75417
| 1000 | 1000 | 2026-02-25 11:46 | 0.89867 | 1e-06 | 0.76674 | 1.7934e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 3465000 | 3465000 | 2026-02-25 11:47 | 0.88459 | 1e-06 | 0.75406 | 2.1801e-06 | 3465 | 1.0009 | 0.75406 | 0.75406 | 0.75406 | 0.75405 | 0.75405
| 1000 | 1000 | 2026-02-25 11:47 | 0.89864 | 1e-06 | 0.76671 | 2.79e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3480000 | 3480000 | 2026-02-25 11:48 | 0.88393 | 1e-06 | 0.75401 | 2.1938e-06 | 3480 | 1.0002 | 0.75402 | 0.75401 | 0.75401 | 0.75401 | 0.75401
| 1000 | 1000 | 2026-02-25 11:48 | 0.89864 | 1e-06 | 0.76672 | 1.7969e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 3495000 | 3495000 | 2026-02-25 11:48 | 0.88314 | 1e-06 | 0.75382 | 1.4357e-06 | 3495 | 0.99956 | 0.75382 | 0.75382 | 0.75382 | 0.75382 | 0.75382
| 1000 | 1000 | 2026-02-25 11:49 | 0.89863 | 1e-06 | 0.7667 | 9.1073e-07 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.7667 | 0.7667
| 3510000 | 3510000 | 2026-02-25 11:49 | 0.88252 | 1e-06 | 0.75373 | 2.7041e-06 | 3510 | 0.99898 | 0.75374 | 0.75373 | 0.75373 | 0.75373 | 0.75372
| 1000 | 1000 | 2026-02-25 11:50 | 0.8986 | 1e-06 | 0.76668 | 3.5136e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 3525000 | 3525000 | 2026-02-25 11:50 | 0.88189 | 1e-06 | 0.75358 | 7.788e-07 | 3525 | 0.99846 | 0.75359 | 0.75358 | 0.75358 | 0.75358 | 0.75359
| 1000 | 1000 | 2026-02-25 11:50 | 0.89862 | 1e-06 | 0.76671 | 3.0228e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3540000 | 3540000 | 2026-02-25 11:51 | 0.88147 | 1e-06 | 0.75356 | 1.8182e-06 | 3540 | 0.99803 | 0.75356 | 0.75356 | 0.75356 | 0.75356 | 0.75355
| 1000 | 1000 | 2026-02-25 11:51 | 0.89863 | 1e-06 | 0.76671 | 2.0329e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76671 | 0.76672 | 0.76672
| 3555000 | 3555000 | 2026-02-25 11:52 | 0.88099 | 1e-06 | 0.75341 | 2.6336e-06 | 3555 | 0.99767 | 0.75341 | 0.75341 | 0.7534 | 0.7534 | 0.7534
| 1000 | 1000 | 2026-02-25 11:52 | 0.89866 | 1e-06 | 0.76673 | 1.2812e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 3570000 | 3570000 | 2026-02-25 11:52 | 0.88065 | 1e-06 | 0.75331 | 2.4208e-06 | 3570 | 0.99742 | 0.75331 | 0.75331 | 0.75331 | 0.7533 | 0.7533
| 1000 | 1000 | 2026-02-25 11:53 | 0.89866 | 1e-06 | 0.76673 | 3.8899e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76674
| 3585000 | 3585000 | 2026-02-25 11:53 | 0.88041 | 1e-06 | 0.75321 | 3.1043e-06 | 3585 | 0.99729 | 0.75322 | 0.75321 | 0.75321 | 0.75321 | 0.7532
| 1000 | 1000 | 2026-02-25 11:54 | 0.89861 | 1e-06 | 0.76669 | 2.1393e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 3600000 | 3600000 | 2026-02-25 11:54 | 0.88026 | 1e-06 | 0.75311 | 2.3608e-06 | 3600 | 0.99724 | 0.75312 | 0.75311 | 0.75312 | 0.75311 | 0.75311
| 1000 | 1000 | 2026-02-25 11:54 | 0.89861 | 1e-06 | 0.76668 | 2.6878e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 3615000 | 3615000 | 2026-02-25 11:55 | 0.88016 | 1e-06 | 0.75296 | 1.6388e-06 | 3615 | 0.99732 | 0.75296 | 0.75296 | 0.75296 | 0.75296 | 0.75296
| 1000 | 1000 | 2026-02-25 11:55 | 0.89864 | 1e-06 | 0.76671 | 1.9705e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76671
| 3630000 | 3630000 | 2026-02-25 11:56 | 0.88023 | 1e-06 | 0.75288 | 1.7058e-06 | 3630 | 0.99751 | 0.75288 | 0.75288 | 0.75288 | 0.75287 | 0.75288
| 1000 | 1000 | 2026-02-25 11:56 | 0.89865 | 1e-06 | 0.76672 | 1.6785e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 3645000 | 3645000 | 2026-02-25 11:56 | 0.88041 | 1e-06 | 0.75279 | 3.562e-06 | 3645 | 0.99783 | 0.7528 | 0.7528 | 0.7528 | 0.75279 | 0.75278
| 1000 | 1000 | 2026-02-25 11:57 | 0.89867 | 1e-06 | 0.76674 | 3.0721e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 3660000 | 3660000 | 2026-02-25 11:57 | 0.88079 | 1e-06 | 0.7528 | 1.5473e-06 | 3660 | 0.99825 | 0.7528 | 0.7528 | 0.75279 | 0.75279 | 0.75279
| 1000 | 1000 | 2026-02-25 11:58 | 0.89865 | 1e-06 | 0.76673 | 2.2659e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 3675000 | 3675000 | 2026-02-25 11:58 | 0.88099 | 1e-06 | 0.75256 | 2.5345e-06 | 3675 | 0.99879 | 0.75257 | 0.75256 | 0.75256 | 0.75256 | 0.75255
| 1000 | 1000 | 2026-02-25 11:58 | 0.89867 | 1e-06 | 0.76674 | 4.6065e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76675 | 0.76675
| 3690000 | 3690000 | 2026-02-25 11:59 | 0.88144 | 1e-06 | 0.75247 | 3.4714e-06 | 3690 | 0.99943 | 0.75248 | 0.75247 | 0.75246 | 0.75246 | 0.75246
| 1000 | 1000 | 2026-02-25 11:59 | 0.89864 | 1e-06 | 0.76671 | 2.3309e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 3705000 | 3705000 | 2026-02-25 12:00 | 0.88195 | 1e-06 | 0.75235 | 1.8774e-06 | 3705 | 1.0001 | 0.75236 | 0.75236 | 0.75236 | 0.75235 | 0.75235
| 1000 | 1000 | 2026-02-25 12:00 | 0.89866 | 1e-06 | 0.76672 | 1.4509e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76672
| 3720000 | 3720000 | 2026-02-25 12:00 | 0.88267 | 1e-06 | 0.75236 | 3.5731e-06 | 3720 | 1.001 | 0.75237 | 0.75237 | 0.75237 | 0.75235 | 0.75235
| 1000 | 1000 | 2026-02-25 12:01 | 0.8987 | 1e-06 | 0.76677 | 9.2764e-07 | #NAME? | #NAME? | 0.76677 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 3735000 | 3735000 | 2026-02-25 12:01 | 0.8832 | 1e-06 | 0.75217 | 2.4587e-06 | 3735 | 1.0018 | 0.75217 | 0.75217 | 0.75217 | 0.75216 | 0.75216
| 1000 | 1000 | 2026-02-25 12:02 | 0.89861 | 1e-06 | 0.76669 | 8.5117e-07 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 3750000 | 3750000 | 2026-02-25 12:02 | 0.88391 | 1e-06 | 0.75208 | 4.0415e-06 | 3750 | 1.0028 | 0.75209 | 0.75209 | 0.75208 | 0.75208 | 0.75207
| 1000 | 1000 | 2026-02-25 12:02 | 0.89859 | 1e-06 | 0.76668 | 2.7637e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76668 | 0.76668 | 0.76669
| 3765000 | 3765000 | 2026-02-25 12:03 | 0.88455 | 1e-06 | 0.75192 | 3.2141e-06 | 3765 | 1.0037 | 0.75192 | 0.75192 | 0.75192 | 0.75191 | 0.7519
| 1000 | 1000 | 2026-02-25 12:03 | 0.89862 | 1e-06 | 0.7667 | 6.2872e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.76671 | 0.76672
| 3780000 | 3780000 | 2026-02-25 12:04 | 0.88536 | 1e-06 | 0.75191 | 2.4302e-06 | 3780 | 1.0046 | 0.75191 | 0.75191 | 0.7519 | 0.7519 | 0.7519
| 1000 | 1000 | 2026-02-25 12:04 | 0.89864 | 1e-06 | 0.76672 | 1.9284e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 3795000 | 3795000 | 2026-02-25 12:04 | 0.88603 | 1e-06 | 0.75179 | 2.4795e-06 | 3795 | 1.0056 | 0.7518 | 0.75179 | 0.75179 | 0.75179 | 0.75179
| 1000 | 1000 | 2026-02-25 12:05 | 0.89868 | 1e-06 | 0.76676 | 1.5812e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 3810000 | 3810000 | 2026-02-25 12:05 | 0.88665 | 1e-06 | 0.75167 | 1.5792e-06 | 3810 | 1.0064 | 0.75167 | 0.75167 | 0.75167 | 0.75167 | 0.75166
| 1000 | 1000 | 2026-02-25 12:06 | 0.89867 | 1e-06 | 0.76674 | 1.9419e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 3825000 | 3825000 | 2026-02-25 12:06 | 0.88719 | 1e-06 | 0.75154 | 2.3879e-06 | 3825 | 1.0072 | 0.75155 | 0.75154 | 0.75154 | 0.75154 | 0.75154
| 1000 | 1000 | 2026-02-25 12:06 | 0.89863 | 1e-06 | 0.7667 | 1.8799e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 3840000 | 3840000 | 2026-02-25 12:07 | 0.88773 | 1e-06 | 0.75151 | 2.1719e-06 | 3840 | 1.0078 | 0.75152 | 0.75151 | 0.75151 | 0.75151 | 0.7515
| 1000 | 1000 | 2026-02-25 12:07 | 0.89863 | 1e-06 | 0.76671 | 3.1969e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.7667 | 0.76671 | 0.76672
| 3855000 | 3855000 | 2026-02-25 12:08 | 0.88802 | 1e-06 | 0.75139 | 1.5584e-06 | 3855 | 1.0083 | 0.75139 | 0.75139 | 0.75139 | 0.75139 | 0.75138
| 1000 | 1000 | 2026-02-25 12:08 | 0.8987 | 1e-06 | 0.76677 | 3.3272e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 3870000 | 3870000 | 2026-02-25 12:08 | 0.88815 | 1e-06 | 0.75125 | 2.9644e-06 | 3870 | 1.0087 | 0.75126 | 0.75126 | 0.75125 | 0.75125 | 0.75125
| 1000 | 1000 | 2026-02-25 12:09 | 0.89867 | 1e-06 | 0.76673 | 2.4352e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 3885000 | 3885000 | 2026-02-25 12:09 | 0.88809 | 1e-06 | 0.75114 | 1.2664e-06 | 3885 | 1.0087 | 0.75115 | 0.75114 | 0.75114 | 0.75114 | 0.75114
| 1000 | 1000 | 2026-02-25 12:10 | 0.89864 | 1e-06 | 0.76671 | 1.6849e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3900000 | 3900000 | 2026-02-25 12:10 | 0.88773 | 1e-06 | 0.75097 | 1.6334e-06 | 3900 | 1.0086 | 0.75097 | 0.75097 | 0.75097 | 0.75097 | 0.75097
| 1000 | 1000 | 2026-02-25 12:10 | 0.89868 | 1e-06 | 0.76675 | 8.8812e-07 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 3915000 | 3915000 | 2026-02-25 12:11 | 0.88721 | 1e-06 | 0.7509 | 1.6889e-06 | 3915 | 1.0081 | 0.75089 | 0.75089 | 0.7509 | 0.7509 | 0.7509
| 1000 | 1000 | 2026-02-25 12:11 | 0.89869 | 1e-06 | 0.76675 | 2.3622e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 3930000 | 3930000 | 2026-02-25 12:12 | 0.88636 | 1e-06 | 0.75082 | 2.1701e-06 | 3930 | 1.0072 | 0.75083 | 0.75082 | 0.75082 | 0.75082 | 0.75082
| 1000 | 1000 | 2026-02-25 12:12 | 0.89864 | 1e-06 | 0.76671 | 9.6497e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3945000 | 3945000 | 2026-02-25 12:12 | 0.88508 | 1e-06 | 0.75066 | 4.0393e-06 | 3945 | 1.006 | 0.75067 | 0.75066 | 0.75066 | 0.75065 | 0.75065
| 1000 | 1000 | 2026-02-25 12:13 | 0.89867 | 1e-06 | 0.76673 | 4.4599e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76673 | 0.76674
| 3960000 | 3960000 | 2026-02-25 12:13 | 0.88364 | 1e-06 | 0.75069 | 5.5712e-07 | 3960 | 1.0043 | 0.75069 | 0.75069 | 0.75069 | 0.75069 | 0.75069
| 1000 | 1000 | 2026-02-25 12:14 | 0.89869 | 1e-06 | 0.76675 | 1.5348e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 3975000 | 3975000 | 2026-02-25 12:14 | 0.88153 | 1e-06 | 0.75051 | 2.1773e-06 | 3975 | 1.0021 | 0.75052 | 0.75052 | 0.75051 | 0.75051 | 0.75051
| 1000 | 1000 | 2026-02-25 12:14 | 0.89867 | 1e-06 | 0.76674 | 7.8883e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 3990000 | 3990000 | 2026-02-25 12:15 | 0.87904 | 1e-06 | 0.7504 | 1.2574e-06 | 3990 | 0.99946 | 0.7504 | 0.75039 | 0.7504 | 0.75039 | 0.75039
| 1000 | 1000 | 2026-02-25 12:15 | 0.89867 | 1e-06 | 0.76674 | 1.5571e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 4005000 | 4005000 | 2026-02-25 12:16 | 0.87602 | 1e-06 | 0.75026 | 3.5043e-06 | 4005 | 0.9962 | 0.75027 | 0.75026 | 0.75026 | 0.75025 | 0.75025
| 1000 | 1000 | 2026-02-25 12:16 | 0.89866 | 1e-06 | 0.76673 | 4.5762e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76674 | 0.76674
| 4020000 | 4020000 | 2026-02-25 12:16 | 0.87257 | 1e-06 | 0.75019 | 8.8729e-07 | 4020 | 0.99237 | 0.75019 | 0.7502 | 0.75019 | 0.75019 | 0.75019
| 1000 | 1000 | 2026-02-25 12:17 | 0.89867 | 1e-06 | 0.76673 | 2.9694e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 4035000 | 4035000 | 2026-02-25 12:17 | 0.86845 | 1e-06 | 0.75004 | 3.6093e-06 | 4035 | 0.98789 | 0.75005 | 0.75005 | 0.75004 | 0.75004 | 0.75003
| 1000 | 1000 | 2026-02-25 12:17 | 0.89862 | 1e-06 | 0.7667 | 9.1146e-07 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 4050000 | 4050000 | 2026-02-25 12:18 | 0.86373 | 1e-06 | 0.74989 | 3.3512e-06 | 4050 | 0.98273 | 0.7499 | 0.74989 | 0.74988 | 0.74988 | 0.74988
| 1000 | 1000 | 2026-02-25 12:18 | 0.89862 | 1e-06 | 0.7667 | 2.7704e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 4065000 | 4065000 | 2026-02-25 12:19 | 0.85843 | 1e-06 | 0.74976 | 3.2923e-06 | 4065 | 0.97686 | 0.74976 | 0.74976 | 0.74976 | 0.74975 | 0.74974
| 1000 | 1000 | 2026-02-25 12:19 | 0.89861 | 1e-06 | 0.76669 | 2.6415e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.7667
| 4080000 | 4080000 | 2026-02-25 12:19 | 0.85252 | 1e-06 | 0.74967 | 9.5915e-07 | 4080 | 0.97026 | 0.74966 | 0.74967 | 0.74967 | 0.74966 | 0.74966
| 1000 | 1000 | 2026-02-25 12:20 | 0.89863 | 1e-06 | 0.76671 | 1.3561e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 4095000 | 4095000 | 2026-02-25 12:20 | 0.84594 | 1e-06 | 0.74958 | 9.2026e-07 | 4095 | 0.96287 | 0.74959 | 0.74958 | 0.74959 | 0.74958 | 0.74958
| 1000 | 1000 | 2026-02-25 12:21 | 0.89866 | 1e-06 | 0.76673 | 1.8772e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 4110000 | 4110000 | 2026-02-25 12:21 | 0.83863 | 1e-06 | 0.74947 | 2.4157e-06 | 4110 | 0.95469 | 0.74948 | 0.74948 | 0.74947 | 0.74947 | 0.74947
| 1000 | 1000 | 2026-02-25 12:21 | 0.89864 | 1e-06 | 0.76672 | 2.2351e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 4125000 | 4125000 | 2026-02-25 12:22 | 0.83057 | 1e-06 | 0.74932 | 2.06e-06 | 4125 | 0.9457 | 0.74933 | 0.74933 | 0.74933 | 0.74932 | 0.74932
| 1000 | 1000 | 2026-02-25 12:22 | 0.89862 | 1e-06 | 0.7667 | 3.228e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 4140000 | 4140000 | 2026-02-25 12:23 | 0.82177 | 1e-06 | 0.74918 | 3.5503e-06 | 4140 | 0.93585 | 0.74919 | 0.74919 | 0.74919 | 0.74918 | 0.74917
| 1000 | 1000 | 2026-02-25 12:23 | 0.89863 | 1e-06 | 0.7667 | 4.167e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 4155000 | 4155000 | 2026-02-25 12:23 | 0.81226 | 1e-06 | 0.74906 | 2.4779e-06 | 4155 | 0.92519 | 0.74906 | 0.74906 | 0.74906 | 0.74906 | 0.74905
| 1000 | 1000 | 2026-02-25 12:24 | 0.89864 | 1e-06 | 0.76671 | 2.5382e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 4170000 | 4170000 | 2026-02-25 12:24 | 0.80202 | 1e-06 | 0.74896 | 2.9767e-06 | 4170 | 0.91364 | 0.74897 | 0.74896 | 0.74896 | 0.74896 | 0.74895
| 1000 | 1000 | 2026-02-25 12:25 | 0.89862 | 1e-06 | 0.7667 | 9.9449e-07 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 4185000 | 4185000 | 2026-02-25 12:25 | 0.79097 | 1e-06 | 0.74882 | 1.5928e-06 | 4185 | 0.90123 | 0.74882 | 0.74882 | 0.74881 | 0.74881 | 0.74881
| 1000 | 1000 | 2026-02-25 12:25 | 0.89861 | 1e-06 | 0.76669 | 1.0923e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 4200000 | 4200000 | 2026-02-25 12:26 | 0.77932 | 1e-06 | 0.74882 | 2.9373e-06 | 4200 | 0.88794 | 0.74883 | 0.74883 | 0.74883 | 0.74882 | 0.74881
| 1000 | 1000 | 2026-02-25 12:26 | 0.89867 | 1e-06 | 0.76674 | 3.1314e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 4215000 | 4215000 | 2026-02-25 12:27 | 0.76669 | 1e-06 | 0.74862 | 2.8229e-06 | 4215 | 0.8738 | 0.74863 | 0.74862 | 0.74862 | 0.74862 | 0.74861
| 1000 | 1000 | 2026-02-25 12:27 | 0.89863 | 1e-06 | 0.76671 | 1.8347e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4230000 | 4230000 | 2026-02-25 12:27 | 0.75342 | 1e-06 | 0.74852 | 3.8287e-06 | 4230 | 0.85877 | 0.74853 | 0.74853 | 0.74852 | 0.74852 | 0.74851
| 1000 | 1000 | 2026-02-25 12:28 | 0.89862 | 1e-06 | 0.7667 | 1.6998e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.7667
| 4245000 | 4245000 | 2026-02-25 12:28 | 0.73938 | 1e-06 | 0.74841 | 3.3947e-06 | 4245 | 0.8429 | 0.74842 | 0.74842 | 0.74841 | 0.74841 | 0.7484
| 1000 | 1000 | 2026-02-25 12:29 | 0.89867 | 1e-06 | 0.76674 | 2.3922e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 4260000 | 4260000 | 2026-02-25 12:29 | 0.72463 | 1e-06 | 0.7483 | 2.1933e-06 | 4260 | 0.8262 | 0.74831 | 0.74831 | 0.7483 | 0.7483 | 0.7483
| 1000 | 1000 | 2026-02-25 12:29 | 0.89865 | 1e-06 | 0.76672 | 1.4619e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76672 | 0.76673
| 4275000 | 4275000 | 2026-02-25 12:30 | 0.70909 | 1e-06 | 0.74812 | 3.3019e-06 | 4275 | 0.80868 | 0.74813 | 0.74813 | 0.74812 | 0.74812 | 0.74811
| 1000 | 1000 | 2026-02-25 12:30 | 0.89862 | 1e-06 | 0.7667 | 1.0517e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 4290000 | 4290000 | 2026-02-25 12:31 | 0.69295 | 1e-06 | 0.74803 | 2.1251e-06 | 4290 | 0.79038 | 0.74803 | 0.74803 | 0.74803 | 0.74803 | 0.74802
| 1000 | 1000 | 2026-02-25 12:31 | 0.89859 | 1e-06 | 0.76668 | 1.9015e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76668
| 4305000 | 4305000 | 2026-02-25 12:31 | 0.67612 | 1e-06 | 0.74791 | 2.481e-06 | 4305 | 0.7713 | 0.74792 | 0.74792 | 0.74791 | 0.7479 | 0.7479
| 1000 | 1000 | 2026-02-25 12:32 | 0.89861 | 1e-06 | 0.76669 | 1.7009e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 4320000 | 4320000 | 2026-02-25 12:32 | 0.65873 | 1e-06 | 0.74786 | 7.3783e-07 | 4320 | 0.75151 | 0.74787 | 0.74786 | 0.74786 | 0.74786 | 0.74786
| 1000 | 1000 | 2026-02-25 12:33 | 0.89873 | 1e-06 | 0.76679 | 7.8944e-07 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.7668
| 4335000 | 4335000 | 2026-02-25 12:33 | 0.64069 | 1e-06 | 0.74777 | 3.8126e-06 | 4335 | 0.73102 | 0.74779 | 0.74778 | 0.74778 | 0.74777 | 0.74776
| 1000 | 1000 | 2026-02-25 12:33 | 0.89868 | 1e-06 | 0.76675 | 1.2914e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 4350000 | 4350000 | 2026-02-25 12:34 | 0.622 | 1e-06 | 0.74759 | 3.8157e-06 | 4350 | 0.70987 | 0.7476 | 0.74759 | 0.74759 | 0.74758 | 0.74758
| 1000 | 1000 | 2026-02-25 12:34 | 0.89863 | 1e-06 | 0.7667 | 3.3988e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76671 | 0.7667 | 0.76671
| 4365000 | 4365000 | 2026-02-25 12:35 | 0.60288 | 1e-06 | 0.74749 | 1.3651e-06 | 4365 | 0.68814 | 0.7475 | 0.74749 | 0.74749 | 0.74749 | 0.74749
| 1000 | 1000 | 2026-02-25 12:35 | 0.89872 | 1e-06 | 0.76678 | 2.3111e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76679
| 4380000 | 4380000 | 2026-02-25 12:35 | 0.58332 | 1e-06 | 0.74746 | 2.5695e-06 | 4380 | 0.66584 | 0.74747 | 0.74746 | 0.74746 | 0.74745 | 0.74745
| 1000 | 1000 | 2026-02-25 12:36 | 0.89874 | 1e-06 | 0.7668 | 1.5671e-06 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.7668
| 4395000 | 4395000 | 2026-02-25 12:36 | 0.56323 | 1e-06 | 0.7473 | 3.1228e-06 | 4395 | 0.64304 | 0.74731 | 0.7473 | 0.7473 | 0.74729 | 0.74729
| 1000 | 1000 | 2026-02-25 12:37 | 0.89871 | 1e-06 | 0.76678 | 1.5211e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 4410000 | 4410000 | 2026-02-25 12:37 | 0.54276 | 1e-06 | 0.74715 | 1.3847e-06 | 4410 | 0.61979 | 0.74716 | 0.74715 | 0.74715 | 0.74715 | 0.74715
| 1000 | 1000 | 2026-02-25 12:37 | 0.89866 | 1e-06 | 0.76673 | 2.4813e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 4425000 | 4425000 | 2026-02-25 12:38 | 0.5219 | 1e-06 | 0.74693 | 4.5304e-06 | 4425 | 0.59615 | 0.74695 | 0.74694 | 0.74693 | 0.74693 | 0.74692
| 1000 | 1000 | 2026-02-25 12:38 | 0.89864 | 1e-06 | 0.76672 | 1.7748e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 4440000 | 4440000 | 2026-02-25 12:39 | 0.50087 | 1e-06 | 0.74686 | 1.7902e-06 | 4440 | 0.57218 | 0.74687 | 0.74686 | 0.74686 | 0.74686 | 0.74686
| 1000 | 1000 | 2026-02-25 12:39 | 0.89863 | 1e-06 | 0.76671 | 3.4736e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76672
| 4455000 | 4455000 | 2026-02-25 12:39 | 0.47953 | 1e-06 | 0.74667 | 1.1081e-06 | 4455 | 0.54794 | 0.74667 | 0.74667 | 0.74668 | 0.74668 | 0.74667
| 1000 | 1000 | 2026-02-25 12:40 | 0.8986 | 1e-06 | 0.76668 | 2.1154e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 4470000 | 4470000 | 2026-02-25 12:40 | 0.45811 | 1e-06 | 0.74664 | 1.2228e-06 | 4470 | 0.5235 | 0.74663 | 0.74664 | 0.74664 | 0.74664 | 0.74664
| 1000 | 1000 | 2026-02-25 12:41 | 0.89872 | 1e-06 | 0.76679 | 3.4731e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.7668 | 0.7668
| 4485000 | 4485000 | 2026-02-25 12:41 | 0.43656 | 1e-06 | 0.74658 | 1.6261e-06 | 4485 | 0.4989 | 0.74658 | 0.74659 | 0.74658 | 0.74658 | 0.74658
| 1000 | 1000 | 2026-02-25 12:41 | 0.89875 | 1e-06 | 0.7668 | 1.8682e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.76681
| 4500000 | 4500000 | 2026-02-25 12:42 | 0.41487 | 1e-06 | 0.74639 | 2.2847e-06 | 4500 | 0.47423 | 0.7464 | 0.7464 | 0.7464 | 0.74639 | 0.74639
| 1000 | 1000 | 2026-02-25 12:42 | 0.89867 | 1e-06 | 0.76673 | 3.2606e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 4515000 | 4515000 | 2026-02-25 12:43 | 0.39318 | 1e-06 | 0.74621 | 2.3295e-06 | 4515 | 0.44955 | 0.74622 | 0.74622 | 0.74621 | 0.74621 | 0.74621
| 1000 | 1000 | 2026-02-25 12:43 | 0.89863 | 1e-06 | 0.76671 | 1.1349e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4530000 | 4530000 | 2026-02-25 12:43 | 0.37158 | 1e-06 | 0.7461 | 3.8705e-06 | 4530 | 0.42491 | 0.74611 | 0.7461 | 0.7461 | 0.74609 | 0.74609
| 1000 | 1000 | 2026-02-25 12:44 | 0.89869 | 1e-06 | 0.76676 | 9.913e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 4545000 | 4545000 | 2026-02-25 12:44 | 0.35009 | 1e-06 | 0.74598 | 1.1319e-06 | 4545 | 0.40041 | 0.74598 | 0.74598 | 0.74598 | 0.74597 | 0.74597
| 1000 | 1000 | 2026-02-25 12:45 | 0.89863 | 1e-06 | 0.76671 | 1.9772e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 4560000 | 4560000 | 2026-02-25 12:45 | 0.32879 | 1e-06 | 0.74592 | 2.3593e-06 | 4560 | 0.37607 | 0.74593 | 0.74593 | 0.74593 | 0.74592 | 0.74592
| 1000 | 1000 | 2026-02-25 12:45 | 0.89867 | 1e-06 | 0.76674 | 2.001e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 4575000 | 4575000 | 2026-02-25 12:46 | 0.30768 | 1e-06 | 0.74579 | 3.5641e-06 | 4575 | 0.35199 | 0.7458 | 0.74579 | 0.74579 | 0.74578 | 0.74578
| 1000 | 1000 | 2026-02-25 12:46 | 0.89864 | 1e-06 | 0.76671 | 1.4214e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4590000 | 4590000 | 2026-02-25 12:47 | 0.28682 | 1e-06 | 0.74558 | 2.4147e-06 | 4590 | 0.32822 | 0.74559 | 0.74558 | 0.74558 | 0.74558 | 0.74558
| 1000 | 1000 | 2026-02-25 12:47 | 0.89864 | 1e-06 | 0.76671 | 3.2241e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4605000 | 4605000 | 2026-02-25 12:47 | 0.26634 | 1e-06 | 0.74551 | 1.3541e-06 | 4605 | 0.3048 | 0.74551 | 0.74551 | 0.74551 | 0.74551 | 0.7455
| 1000 | 1000 | 2026-02-25 12:48 | 0.89865 | 1e-06 | 0.76672 | 2.9802e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 4620000 | 4620000 | 2026-02-25 12:48 | 0.24619 | 1e-06 | 0.74534 | 2.191e-06 | 4620 | 0.28182 | 0.74535 | 0.74534 | 0.74534 | 0.74534 | 0.74533
| 1000 | 1000 | 2026-02-25 12:49 | 0.89867 | 1e-06 | 0.76674 | 2.1043e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.76674
| 4635000 | 4635000 | 2026-02-25 12:49 | 0.22648 | 1e-06 | 0.74519 | 4.1971e-06 | 4635 | 0.25931 | 0.7452 | 0.74519 | 0.74519 | 0.74518 | 0.74517
| 1000 | 1000 | 2026-02-25 12:49 | 0.89864 | 1e-06 | 0.76672 | 1.9671e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 4650000 | 4650000 | 2026-02-25 12:50 | 0.20725 | 1e-06 | 0.74506 | 5.0601e-06 | 4650 | 0.23734 | 0.74507 | 0.74507 | 0.74505 | 0.74505 | 0.74504
| 1000 | 1000 | 2026-02-25 12:50 | 0.8986 | 1e-06 | 0.76668 | 3.5592e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 4665000 | 4665000 | 2026-02-25 12:50 | 0.18854 | 1e-06 | 0.74492 | 1.5007e-06 | 4665 | 0.21594 | 0.74493 | 0.74492 | 0.74492 | 0.74492 | 0.74492
| 1000 | 1000 | 2026-02-25 12:51 | 0.89859 | 1e-06 | 0.76668 | 2.3483e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76668
| 4680000 | 4680000 | 2026-02-25 12:51 | 0.17039 | 1e-06 | 0.74482 | 1.2565e-06 | 4680 | 0.19518 | 0.74482 | 0.74482 | 0.74482 | 0.74481 | 0.74482
| 1000 | 1000 | 2026-02-25 12:52 | 0.89859 | 1e-06 | 0.76667 | 1.1622e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76668 | 0.76668 | 0.76668
| 4695000 | 4695000 | 2026-02-25 12:52 | 0.15281 | 1e-06 | 0.74464 | 3.6904e-06 | 4695 | 0.17509 | 0.74465 | 0.74464 | 0.74464 | 0.74463 | 0.74463
| 1000 | 1000 | 2026-02-25 12:52 | 0.89862 | 1e-06 | 0.7667 | 1.9569e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76669 | 0.76669 | 0.7667
| 4710000 | 4710000 | 2026-02-25 12:53 | 0.13589 | 1e-06 | 0.74462 | 2.9659e-06 | 4710 | 0.1557 | 0.74463 | 0.74462 | 0.74461 | 0.74461 | 0.74461
| 1000 | 1000 | 2026-02-25 12:53 | 0.8986 | 1e-06 | 0.76668 | 1.9786e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 4725000 | 4725000 | 2026-02-25 12:54 | 0.1196 | 1e-06 | 0.74445 | 2.7025e-06 | 4725 | 0.13707 | 0.74446 | 0.74445 | 0.74445 | 0.74445 | 0.74444
| 1000 | 1000 | 2026-02-25 12:54 | 0.89861 | 1e-06 | 0.76669 | 3.2462e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 4740000 | 4740000 | 2026-02-25 12:54 | 0.104 | 1e-06 | 0.74434 | 1.8178e-06 | 4740 | 0.11921 | 0.74435 | 0.74434 | 0.74434 | 0.74434 | 0.74434
| 1000 | 1000 | 2026-02-25 12:55 | 0.89863 | 1e-06 | 0.7667 | 3.3611e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 4755000 | 4755000 | 2026-02-25 12:55 | 0.089111 | 1e-06 | 0.74422 | 7.6288e-07 | 4755 | 0.10216 | 0.74422 | 0.74422 | 0.74422 | 0.74423 | 0.74422
| 1000 | 1000 | 2026-02-25 12:56 | 0.89865 | 1e-06 | 0.76672 | 2.2508e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 4770000 | 4770000 | 2026-02-25 12:56 | 0.074945 | 1e-06 | 0.74405 | 1.4701e-06 | 4770 | 0.085938 | 0.74406 | 0.74406 | 0.74406 | 0.74405 | 0.74405
| 1000 | 1000 | 2026-02-25 12:56 | 0.8987 | 1e-06 | 0.76676 | 2.0999e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 4785000 | 4785000 | 2026-02-25 12:57 | Error | 1e-06 | 0.74392 | 3.3427e-06 | 4785 | #VALUE! | 0.74393 | 0.74392 | 0.74392 | 0.74391 | 0.74391
| 1000 | 1000 | 2026-02-25 12:58 | 0.89855 | 1e-06 | 0.76664 | 1.7366e-06 | #NAME? | #NAME? | 0.76664 | 0.76664 | 0.76664 | 0.76664 | 0.76664
| 4800000 | 4800000 | 2026-02-25 12:58 | Error | 1e-06 | 0.74381 | 2.171e-06 | 4800 | #VALUE! | 0.74382 | 0.74381 | 0.74381 | 0.74381 | 0.7438
| 1000 | 1000 | 2026-02-25 12:59 | 0.89859 | 1e-06 | 0.76667 | 1.1606e-06 | #NAME? | #NAME? | 0.76666 | 0.76666 | 0.76667 | 0.76667 | 0.76667
| 4815000 | 4815000 | 2026-02-25 12:59 | Error | 1e-06 | 0.74367 | 2.8228e-06 | 4815 | #VALUE! | 0.74368 | 0.74368 | 0.74368 | 0.74367 | 0.74367
| 1000 | 1000 | 2026-02-25 13:00 | 0.89855 | 1e-06 | 0.76663 | 3.042e-06 | #NAME? | #NAME? | 0.76662 | 0.76663 | 0.76662 | 0.76664 | 0.76664
| 4830000 | 4830000 | 2026-02-25 13:01 | Error | 1e-06 | 0.74348 | 6.3614e-07 | 4830 | #VALUE! | 0.74348 | 0.74348 | 0.74348 | 0.74348 | 0.74348
| 1000 | 1000 | 2026-02-25 13:01 | 0.89854 | 1e-06 | 0.76663 | 1.7644e-06 | #NAME? | #NAME? | 0.76662 | 0.76662 | 0.76663 | 0.76663 | 0.76663
| 4845000 | 4845000 | 2026-02-25 13:02 | Error | 1e-06 | 0.74352 | 2.6344e-06 | 4845 | #VALUE! | 0.74353 | 0.74352 | 0.74352 | 0.74352 | 0.74351
| 1000 | 1000 | 2026-02-25 13:02 | 0.89857 | 1e-06 | 0.76667 | 2.4892e-06 | #NAME? | #NAME? | 0.76666 | 0.76666 | 0.76667 | 0.76667 | 0.76667
| 4860000 | 4860000 | 2026-02-25 13:03 | Error | 1e-06 | 0.74325 | 1.9664e-06 | 4860 | #VALUE! | 0.74326 | 0.74325 | 0.74325 | 0.74325 | 0.74325
| 1000 | 1000 | 2026-02-25 13:03 | 0.89853 | 1e-06 | 0.76663 | 2.9326e-06 | #NAME? | #NAME? | 0.76662 | 0.76662 | 0.76663 | 0.76663 | 0.76664
| 4875000 | 4875000 | 2026-02-25 13:04 | Error | 1e-06 | 0.74311 | 2.9514e-06 | 4875 | #VALUE! | 0.74312 | 0.74311 | 0.74311 | 0.74311 | 0.7431
| 1000 | 1000 | 2026-02-25 13:04 | 0.89862 | 1e-06 | 0.76669 | 1.1486e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 4890000 | 4890000 | 2026-02-25 13:05 | Error | 1e-06 | 0.74302 | 2.9686e-06 | 4890 | #VALUE! | 0.74302 | 0.74302 | 0.74302 | 0.74302 | 0.74301
| 1000 | 1000 | 2026-02-25 13:05 | 0.89856 | 1e-06 | 0.76664 | 4.714e-06 | #NAME? | #NAME? | 0.76663 | 0.76664 | 0.76664 | 0.76665 | 0.76666
| 4905000 | 4905000 | 2026-02-25 13:06 | Error | 1e-06 | 0.74286 | 3.0297e-06 | 4905 | #VALUE! | 0.74287 | 0.74286 | 0.74286 | 0.74285 | 0.74285
| 1000 | 1000 | 2026-02-25 13:07 | 0.89861 | 1e-06 | 0.76668 | 3.1655e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 4920000 | 4920000 | 2026-02-25 13:07 | Error | 1e-06 | 0.74278 | 1.8305e-06 | 4920 | #VALUE! | 0.74278 | 0.74278 | 0.74278 | 0.74278 | 0.74277
| 1000 | 1000 | 2026-02-25 13:08 | 0.89853 | 1e-06 | 0.76663 | 2.5301e-06 | #NAME? | #NAME? | 0.76662 | 0.76662 | 0.76662 | 0.76663 | 0.76663
| 4935000 | 4935000 | 2026-02-25 13:08 | Error | 1e-06 | 0.74263 | 7.7948e-07 | 4935 | #VALUE! | 0.74263 | 0.74263 | 0.74263 | 0.74263 | 0.74263
| 1000 | 1000 | 2026-02-25 13:09 | 0.89856 | 1e-06 | 0.76666 | 2.1665e-06 | #NAME? | #NAME? | 0.76665 | 0.76665 | 0.76665 | 0.76666 | 0.76666
| 4950000 | 4950000 | 2026-02-25 13:10 | Error | 1e-06 | 0.74258 | 8.423e-07 | 4950 | #VALUE! | 0.74258 | 0.74259 | 0.74258 | 0.74258 | 0.74258
| 1000 | 1000 | 2026-02-25 13:10 | 0.89858 | 1e-06 | 0.76666 | 2.5383e-06 | #NAME? | #NAME? | 0.76665 | 0.76666 | 0.76666 | 0.76667 | 0.76667
| 4965000 | 4965000 | 2026-02-25 13:11 | Error | 1e-06 | 0.74243 | 1.4482e-06 | 4965 | #VALUE! | 0.74244 | 0.74244 | 0.74243 | 0.74243 | 0.74243
| 1000 | 1000 | 2026-02-25 13:11 | 0.8986 | 1e-06 | 0.76669 | 7.6563e-07 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 4980000 | 4980000 | 2026-02-25 13:12 | Error | 1e-06 | 0.7423 | 1.5602e-06 | 4980 | #VALUE! | 0.7423 | 0.7423 | 0.7423 | 0.7423 | 0.7423
| 1000 | 1000 | 2026-02-25 13:12 | 0.89868 | 1e-06 | 0.76674 | 2.5003e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 4995000 | 4995000 | 2026-02-25 13:13 | Error | 1e-06 | 0.74215 | 1.3397e-06 | 4995 | #VALUE! | 0.74215 | 0.74215 | 0.74214 | 0.74215 | 0.74214
STOP | 1000 | 1000 | 2026-02-25 13:13 | 0.89867 | 1e-06 | 0.76674 | 5.6439e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 5010000 | 5010000 | [5 empty] | 5010 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5025000 | 5025000 | [5 empty] | 5025 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5040000 | 5040000 | [5 empty] | 5040 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5055000 | 5055000 | [5 empty] | 5055 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5070000 | 5070000 | [5 empty] | 5070 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5085000 | 5085000 | [5 empty] | 5085 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5100000 | 5100000 | [5 empty] | 5100 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5115000 | 5115000 | [5 empty] | 5115 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5130000 | 5130000 | [5 empty] | 5130 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5145000 | 5145000 | [5 empty] | 5145 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5160000 | 5160000 | [5 empty] | 5160 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5175000 | 5175000 | [5 empty] | 5175 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5190000 | 5190000 | [5 empty] | 5190 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5205000 | 5205000 | [5 empty] | 5205 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5220000 | 5220000 | [5 empty] | 5220 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5235000 | 5235000 | [5 empty] | 5235 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5250000 | 5250000 | [5 empty] | 5250 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5265000 | 5265000 | [5 empty] | 5265 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5280000 | 5280000 | [5 empty] | 5280 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5295000 | 5295000 | [5 empty] | 5295 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5310000 | 5310000 | [5 empty] | 5310 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5325000 | 5325000 | [5 empty] | 5325 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5340000 | 5340000 | [5 empty] | 5340 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5355000 | 5355000 | [5 empty] | 5355 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5370000 | 5370000 | [5 empty] | 5370 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5385000 | 5385000 | [5 empty] | 5385 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5400000 | 5400000 | [5 empty] | 5400 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5415000 | 5415000 | [5 empty] | 5415 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5430000 | 5430000 | [5 empty] | 5430 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5445000 | 5445000 | [5 empty] | 5445 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5460000 | 5460000 | [5 empty] | 5460 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5475000 | 5475000 | [5 empty] | 5475 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5490000 | 5490000 | [5 empty] | 5490 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5505000 | 5505000 | [5 empty] | 5505 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5520000 | 5520000 | [5 empty] | 5520 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5535000 | 5535000 | [5 empty] | 5535 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5550000 | 5550000 | [5 empty] | 5550 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5565000 | 5565000 | [5 empty] | 5565 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5580000 | 5580000 | [5 empty] | 5580 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5595000 | 5595000 | [5 empty] | 5595 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5610000 | 5610000 | [5 empty] | 5610 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5625000 | 5625000 | [5 empty] | 5625 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5640000 | 5640000 | [5 empty] | 5640 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5655000 | 5655000 | [5 empty] | 5655 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5670000 | 5670000 | [5 empty] | 5670 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5685000 | 5685000 | [5 empty] | 5685 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5700000 | 5700000 | [5 empty] | 5700 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5715000 | 5715000 | [5 empty] | 5715 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5730000 | 5730000 | [5 empty] | 5730 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5745000 | 5745000 | [5 empty] | 5745 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5760000 | 5760000 | [5 empty] | 5760 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5775000 | 5775000 | [5 empty] | 5775 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5790000 | 5790000 | [5 empty] | 5790 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5805000 | 5805000 | [5 empty] | 5805 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5820000 | 5820000 | [5 empty] | 5820 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5835000 | 5835000 | [5 empty] | 5835 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5850000 | 5850000 | [5 empty] | 5850 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5865000 | 5865000 | [5 empty] | 5865 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5880000 | 5880000 | [5 empty] | 5880 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5895000 | 5895000 | [5 empty] | 5895 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5910000 | 5910000 | [5 empty] | 5910 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5925000 | 5925000 | [5 empty] | 5925 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5940000 | 5940000 | [5 empty] | 5940 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5955000 | 5955000 | [5 empty] | 5955 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5970000 | 5970000 | [5 empty] | 5970 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 5985000 | 5985000 | [5 empty] | 5985 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6000000 | 6000000 | [5 empty] | 6000 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6015000 | 6015000 | [5 empty] | 6015 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6030000 | 6030000 | [5 empty] | 6030 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6045000 | 6045000 | [5 empty] | 6045 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6060000 | 6060000 | [5 empty] | 6060 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6075000 | 6075000 | [5 empty] | 6075 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6090000 | 6090000 | [5 empty] | 6090 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6105000 | 6105000 | [5 empty] | 6105 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6120000 | 6120000 | [5 empty] | 6120 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6135000 | 6135000 | [5 empty] | 6135 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6150000 | 6150000 | [5 empty] | 6150 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6165000 | 6165000 | [5 empty] | 6165 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6180000 | 6180000 | [5 empty] | 6180 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6195000 | 6195000 | [5 empty] | 6195 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6210000 | 6210000 | [5 empty] | 6210 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6225000 | 6225000 | [5 empty] | 6225 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6240000 | 6240000 | [5 empty] | 6240 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6255000 | 6255000 | [5 empty] | 6255 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6270000 | 6270000 | [5 empty] | 6270 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6285000 | 6285000 | [5 empty] | 6285 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6300000 | 6300000 | [5 empty] | 6300 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6315000 | 6315000 | [5 empty] | 6315 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6330000 | 6330000 | [5 empty] | 6330 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6345000 | 6345000 | [5 empty] | 6345 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6360000 | 6360000 | [5 empty] | 6360 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6375000 | 6375000 | [5 empty] | 6375 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6390000 | 6390000 | [5 empty] | 6390 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6405000 | 6405000 | [5 empty] | 6405 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6420000 | 6420000 | [5 empty] | 6420 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6435000 | 6435000 | [5 empty] | 6435 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6450000 | 6450000 | [5 empty] | 6450 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6465000 | 6465000 | [5 empty] | 6465 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6480000 | 6480000 | [5 empty] | 6480 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6495000 | 6495000 | [5 empty] | 6495 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6510000 | 6510000 | [5 empty] | 6510 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6525000 | 6525000 | [5 empty] | 6525 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6540000 | 6540000 | [5 empty] | 6540 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6555000 | 6555000 | [5 empty] | 6555 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6570000 | 6570000 | [5 empty] | 6570 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6585000 | 6585000 | [5 empty] | 6585 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6600000 | 6600000 | [5 empty] | 6600 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6615000 | 6615000 | [5 empty] | 6615 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6630000 | 6630000 | [5 empty] | 6630 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6645000 | 6645000 | [5 empty] | 6645 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6660000 | 6660000 | [5 empty] | 6660 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6675000 | 6675000 | [5 empty] | 6675 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6690000 | 6690000 | [5 empty] | 6690 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6705000 | 6705000 | [5 empty] | 6705 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6720000 | 6720000 | [5 empty] | 6720 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6735000 | 6735000 | [5 empty] | 6735 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6750000 | 6750000 | [5 empty] | 6750 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6765000 | 6765000 | [5 empty] | 6765 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6780000 | 6780000 | [5 empty] | 6780 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6795000 | 6795000 | [5 empty] | 6795 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6810000 | 6810000 | [5 empty] | 6810 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6825000 | 6825000 | [5 empty] | 6825 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6840000 | 6840000 | [5 empty] | 6840 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6855000 | 6855000 | [5 empty] | 6855 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6870000 | 6870000 | [5 empty] | 6870 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6885000 | 6885000 | [5 empty] | 6885 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6900000 | 6900000 | [5 empty] | 6900 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6915000 | 6915000 | [5 empty] | 6915 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6930000 | 6930000 | [5 empty] | 6930 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6945000 | 6945000 | [5 empty] | 6945 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6960000 | 6960000 | [5 empty] | 6960 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6975000 | 6975000 | [5 empty] | 6975 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 6990000 | 6990000 | [5 empty] | 6990 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7005000 | 7005000 | [5 empty] | 7005 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7020000 | 7020000 | [5 empty] | 7020 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7035000 | 7035000 | [5 empty] | 7035 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7050000 | 7050000 | [5 empty] | 7050 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7065000 | 7065000 | [5 empty] | 7065 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7080000 | 7080000 | [5 empty] | 7080 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7095000 | 7095000 | [5 empty] | 7095 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7110000 | 7110000 | [5 empty] | 7110 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7125000 | 7125000 | [5 empty] | 7125 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7140000 | 7140000 | [5 empty] | 7140 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7155000 | 7155000 | [5 empty] | 7155 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7170000 | 7170000 | [5 empty] | 7170 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7185000 | 7185000 | [5 empty] | 7185 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7200000 | 7200000 | [5 empty] | 7200 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7215000 | 7215000 | [5 empty] | 7215 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7230000 | 7230000 | [5 empty] | 7230 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7245000 | 7245000 | [5 empty] | 7245 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7260000 | 7260000 | [5 empty] | 7260 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7275000 | 7275000 | [5 empty] | 7275 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7290000 | 7290000 | [5 empty] | 7290 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7305000 | 7305000 | [5 empty] | 7305 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7320000 | 7320000 | [5 empty] | 7320 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7335000 | 7335000 | [5 empty] | 7335 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7350000 | 7350000 | [5 empty] | 7350 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7365000 | 7365000 | [5 empty] | 7365 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7380000 | 7380000 | [5 empty] | 7380 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7395000 | 7395000 | [5 empty] | 7395 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7410000 | 7410000 | [5 empty] | 7410 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7425000 | 7425000 | [5 empty] | 7425 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7440000 | 7440000 | [5 empty] | 7440 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7455000 | 7455000 | [5 empty] | 7455 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7470000 | 7470000 | [5 empty] | 7470 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7485000 | 7485000 | [5 empty] | 7485 | #DIV/0!
| 1000 | 1000 | [5 empty] | #NAME? | #NAME?
| 7500000 | 7500000 | [5 empty] | 7500 | #DIV/0!
| 1000 | 1000 | [5 empty] | 1 | #DIV/0!
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
Note on C12: Martin Šíra: The record length should be such to get coherent sampling. If the number of periods is the same as the signal frequency, coherent sampling should be always met, however the data processing can take a lot of time.
Note on D2: labina: Real sine amplitude, i.e. A*sin(w*t). It is NOT rms or peak-to-peak value.
Note on D7: labina: The algorithm must generate output quantity 'rms'. That is e.g. TWM-WFFT or TWM-WRMS
Note on D8: labina: Set to 1 to measure simultaneously both Digitizer amplitude and DC multimeter voltage. Set to 0 or keep empty to measure first AC and DC after.

## PTB reference data
10 | 1.0002
30 | 1.0001
50 | 0.99997
70 | 0.9998
90 | 0.99956
110 | 0.99926
130 | 0.9989
150 | 0.99849
170 | 0.99804
190 | 0.99756
210 | 0.99708
230 | 0.9966
250 | 0.99614
270 | 0.99572
290 | 0.99536
310 | 0.99506
330 | 0.99485
350 | 0.99472
370 | 0.99468
390 | 0.99475
410 | 0.99492
430 | 0.99519
450 | 0.99556
470 | 0.99602
490 | 0.99657
510 | 0.9972
530 | 0.9979
550 | 0.99865
570 | 0.99945
590 | 1.0003
610 | 1.0011
630 | 1.0019
650 | 1.0027
670 | 1.0035
690 | 1.0041
710 | 1.0048
730 | 1.0053
750 | 1.0057
770 | 1.0061
790 | 1.0063
810 | 1.0064
830 | 1.0063
850 | 1.0062
870 | 1.0059
890 | 1.0055
910 | 1.005
930 | 1.0044
950 | 1.0037
970 | 1.003
990 | 1.0022
1010 | 1.0014
1030 | 1.0006
1050 | 0.99981
1070 | 0.99903
1090 | 0.9983
1110 | 0.99764
1130 | 0.99706
1150 | 0.99657
1170 | 0.99619
1190 | 0.99593
1210 | 0.99579
1230 | 0.99579
1250 | 0.99591
1270 | 0.99615
1290 | 0.99652
1310 | 0.997
1330 | 0.99759
1350 | 0.99827
1370 | 0.99902
1390 | 0.99983
1410 | 1.0007
1430 | 1.0015
1450 | 1.0024
1470 | 1.0033
1490 | 1.0041
1510 | 1.0048
1530 | 1.0055
1550 | 1.0061
1570 | 1.0066
1590 | 1.007
1610 | 1.0073
1630 | 1.0074
1650 | 1.0074
1670 | 1.0073
1690 | 1.007
1710 | 1.0066
1730 | 1.0062
1750 | 1.0056
1770 | 1.0049
1790 | 1.0042
1810 | 1.0034
1830 | 1.0026
1850 | 1.0018
1870 | 1.001
1890 | 1.0002
1910 | 0.99944
1930 | 0.99876
1950 | 0.99817
1970 | 0.99767
1990 | 0.99728
2010 | 0.99701
2030 | 0.99687
2050 | 0.99686
2070 | 0.99697
2090 | 0.99722
2110 | 0.99759
2130 | 0.99808
2150 | 0.99867
2170 | 0.99936
2190 | 1.0001
2210 | 1.0009
2230 | 1.0018
2250 | 1.0027
2270 | 1.0036
2290 | 1.0044
2310 | 1.0052
2330 | 1.006
2350 | 1.0067
2370 | 1.0073
2390 | 1.0078
2410 | 1.0081
2430 | 1.0084
2450 | 1.0085
2470 | 1.0084
2490 | 1.0083
2510 | 1.008
2530 | 1.0076
2550 | 1.007
2570 | 1.0064
2590 | 1.0057
2610 | 1.0049
2630 | 1.0041
2650 | 1.0033
2670 | 1.0024
2690 | 1.0016
2710 | 1.0008
2730 | 1.0001
2750 | 0.99944
2770 | 0.99889
2790 | 0.99847
2810 | 0.99816
2830 | 0.99799
2850 | 0.99797
2870 | 0.99808
2890 | 0.99834
2910 | 0.99873
2930 | 0.99925
2950 | 0.99988
2970 | 1.0006
2990 | 1.0014
3010 | 1.0023
3030 | 1.0032
3050 | 1.0042
3070 | 1.0051
3090 | 1.006
3110 | 1.0069
3130 | 1.0076
3150 | 1.0083
3170 | 1.0089
3190 | 1.0093
3210 | 1.0096
3230 | 1.0097
3250 | 1.0097
3270 | 1.0095
3290 | 1.0092
3310 | 1.0087
3330 | 1.0081
3350 | 1.0074
3370 | 1.0066
3390 | 1.0057
3410 | 1.0048
3430 | 1.0038
3450 | 1.0029
3470 | 1.002
3490 | 1.0012
3510 | 1.0005
3530 | 0.99991
3550 | 0.99948
3570 | 0.99922
3590 | 0.99915
3610 | 0.99928
3630 | 0.99962
3650 | 1.0002
3670 | 1.0009
3690 | 1.0018
3710 | 1.0029
3730 | 1.004
3750 | 1.0053
3770 | 1.0065
3790 | 1.0077
3810 | 1.0088
3830 | 1.0097
3850 | 1.0103
3870 | 1.0105
3890 | 1.0103
3910 | 1.0096
3930 | 1.0082
3950 | 1.0061
3970 | 1.0031
3990 | 0.9992
4010 | 0.99422
4030 | 0.98798
4050 | 0.98019
4070 | 0.97043
4090 | 0.95805