## List1
Frequency response measurement template |  |  | measurement properties |  |  | hardware data |  |  |  | Actual number of readings: |  |  |  | 1000
Software will read settings from cells E2 to E7 and |  |  | AC voltage amplitude | 1.414 | V | AC source idn | TEKTRONIX,AFG3101C,C010192,SCPI:99.0 FV:1.0.2 |  |  | Crude estimate of measurement time (h): |  |  |  | 5.8056
measurement points from collumns A-C (yellow cells). |  |  | Sampling frequency | 15000000 | Hz | DC meter idn | Keithley 2182
Results will be written into collumns D-H (and K+). Sheet must |  |  | AC-DC settle time | 20 | s | Digitizer idn | NI PXI-5922, sn. F22D5B, channel 0
contain keywords START and STOP in collumn A, values of |  |  | AC-DC warm up time | 900 | s
frequency in collumn B, and record length in C. |  |  | DC meter readings | 5
Place keyword STOP to the row with half of the sampling frequency. |  |  | Algorithm | TWM-WRMS
|  |  | Simultaneous meas.? | 1
inputs |  |  | results
Command | Signal frequency | Record length | TIme of DC readings | Digitizer amplitude rms |  | DC voltage |  | Calculation |  | DC meter readings
(START, STOP) | Hz | multiple of periods | (time of writing this line into excel) | value (V) | unc (V) | value (V) | unc (V) | for figure |  | ¦---------------->
START | 15000 | 15000 | 2026-02-21 11:38 | 0.90111 | 0 | 0.76879 | 8.2139e-07 | 15 | 1 | 0.76879 | 0.76879 | 0.76879 | 0.76879 | 0.76879
| 1000 | 1000 | 2026-02-21 11:38 | 0.8988 | 0 | 0.76683 | 9.6417e-07 | #NAME? | #NAME? | 0.76683 | 0.76683 | 0.76683 | 0.76683 | 0.76683
| 30000 | 30000 | 2026-02-21 11:38 | 0.89908 | 0 | 0.76702 | 2.0129e-06 | 30 | 1.0001 | 0.76702 | 0.76702 | 0.76701 | 0.76702 | 0.76701
| 1000 | 1000 | 2026-02-21 11:39 | 0.89878 | 0 | 0.76683 | 1.169e-06 | #NAME? | #NAME? | 0.76682 | 0.76683 | 0.76683 | 0.76683 | 0.76683
| 45000 | 45000 | 2026-02-21 11:39 | 0.89874 | 0 | 0.76665 | 2.6037e-06 | 45 | 1.0002 | 0.76666 | 0.76665 | 0.76665 | 0.76664 | 0.76665
| 1000 | 1000 | 2026-02-21 11:40 | 0.89873 | 0 | 0.76678 | 1.2863e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 60000 | 60000 | 2026-02-21 11:40 | 0.89877 | 0 | 0.76655 | 1.0402e-06 | 60 | 1.0003 | 0.76656 | 0.76656 | 0.76656 | 0.76655 | 0.76655
| 1000 | 1000 | 2026-02-21 11:40 | 0.89877 | 0 | 0.76682 | 1.3135e-06 | #NAME? | #NAME? | 0.76683 | 0.76682 | 0.76682 | 0.76682 | 0.76682
| 75000 | 75000 | 2026-02-21 11:41 | 0.8988 | 0 | 0.76645 | 1.746e-06 | 75 | 1.0005 | 0.76645 | 0.76645 | 0.76645 | 0.76644 | 0.76644
| 1000 | 1000 | 2026-02-21 11:41 | 0.89872 | 0 | 0.76677 | 1.3659e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 90000 | 90000 | 2026-02-21 11:42 | 0.89897 | 0 | 0.76642 | 1.325e-06 | 90 | 1.0007 | 0.76642 | 0.76642 | 0.76641 | 0.76642 | 0.76641
| 1000 | 1000 | 2026-02-21 11:42 | 0.89871 | 0 | 0.76676 | 1.5551e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 105000 | 105000 | 2026-02-21 11:43 | 0.89914 | 0 | 0.76636 | 6.3177e-07 | 105 | 1.001 | 0.76636 | 0.76636 | 0.76636 | 0.76636 | 0.76636
| 1000 | 1000 | 2026-02-21 11:43 | 0.8987 | 0 | 0.76676 | 8.5742e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 120000 | 120000 | 2026-02-21 11:43 | 0.89935 | 0 | 0.76633 | 1.4698e-06 | 120 | 1.0013 | 0.76634 | 0.76633 | 0.76633 | 0.76633 | 0.76633
| 1000 | 1000 | 2026-02-21 11:44 | 0.89871 | 0 | 0.76676 | 1.9222e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 135000 | 135000 | 2026-02-21 11:44 | 0.89958 | 0 | 0.76627 | 1.1818e-06 | 135 | 1.0016 | 0.76628 | 0.76628 | 0.76627 | 0.76627 | 0.76627
| 1000 | 1000 | 2026-02-21 11:45 | 0.89876 | 0 | 0.7668 | 1.4589e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.7668
| 150000 | 150000 | 2026-02-21 11:45 | 0.89985 | 0 | 0.76623 | 8.6562e-07 | 150 | 1.002 | 0.76623 | 0.76623 | 0.76623 | 0.76623 | 0.76623
| 1000 | 1000 | 2026-02-21 11:46 | 0.89879 | 0 | 0.76682 | 3.8638e-06 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76682 | 0.76683 | 0.76683
| 165000 | 165000 | 2026-02-21 11:46 | 0.90006 | 0 | 0.76612 | 2.602e-06 | 165 | 1.0023 | 0.76612 | 0.76612 | 0.76612 | 0.76611 | 0.76611
| 1000 | 1000 | 2026-02-21 11:46 | 0.89871 | 0 | 0.76676 | 2.1511e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 180000 | 180000 | 2026-02-21 11:47 | 0.90042 | 0 | 0.7661 | 1.7937e-06 | 180 | 1.0028 | 0.76611 | 0.76611 | 0.7661 | 0.7661 | 0.7661
| 1000 | 1000 | 2026-02-21 11:47 | 0.8987 | 0 | 0.76676 | 1.3075e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 195000 | 195000 | 2026-02-21 11:48 | 0.90068 | 0 | 0.76598 | 1.6741e-06 | 195 | 1.0032 | 0.76599 | 0.76599 | 0.76598 | 0.76598 | 0.76598
| 1000 | 1000 | 2026-02-21 11:48 | 0.89867 | 0 | 0.76673 | 2.4538e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76673 | 0.76674
| 210000 | 210000 | 2026-02-21 11:48 | 0.90105 | 0 | 0.76595 | 1.7022e-06 | 210 | 1.0037 | 0.76595 | 0.76596 | 0.76595 | 0.76595 | 0.76595
| 1000 | 1000 | 2026-02-21 11:49 | 0.89871 | 0 | 0.76677 | 3.1591e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 225000 | 225000 | 2026-02-21 11:49 | 0.90139 | 0 | 0.76587 | 9.7105e-07 | 225 | 1.0042 | 0.76587 | 0.76587 | 0.76588 | 0.76587 | 0.76587
| 1000 | 1000 | 2026-02-21 11:50 | 0.89872 | 0 | 0.76677 | 6.5737e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 240000 | 240000 | 2026-02-21 11:50 | 0.90183 | 0 | 0.76586 | 1.3101e-06 | 240 | 1.0047 | 0.76587 | 0.76586 | 0.76586 | 0.76586 | 0.76586
| 1000 | 1000 | 2026-02-21 11:51 | 0.89873 | 0 | 0.76678 | 1.3294e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 255000 | 255000 | 2026-02-21 11:51 | 0.90212 | 0 | 0.76572 | 2.6252e-06 | 255 | 1.0052 | 0.76573 | 0.76572 | 0.76572 | 0.76572 | 0.76571
| 1000 | 1000 | 2026-02-21 11:51 | 0.89873 | 0 | 0.76678 | 8.1867e-07 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 270000 | 270000 | 2026-02-21 11:52 | 0.90254 | 0 | 0.76568 | 1.672e-06 | 270 | 1.0057 | 0.76568 | 0.76568 | 0.76568 | 0.76568 | 0.76567
| 1000 | 1000 | 2026-02-21 11:52 | 0.89869 | 0 | 0.76675 | 1.6136e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 285000 | 285000 | 2026-02-21 11:53 | 0.90285 | 0 | 0.76553 | 1.468e-06 | 285 | 1.0062 | 0.76554 | 0.76554 | 0.76553 | 0.76553 | 0.76553
| 1000 | 1000 | 2026-02-21 11:53 | 0.89869 | 0 | 0.76675 | 1.6713e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 300000 | 300000 | 2026-02-21 11:54 | 0.90324 | 0 | 0.76546 | 1.2898e-06 | 300 | 1.0068 | 0.76547 | 0.76546 | 0.76546 | 0.76546 | 0.76546
| 1000 | 1000 | 2026-02-21 11:54 | 0.89866 | 0 | 0.76672 | 2.7928e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 315000 | 315000 | 2026-02-21 11:54 | 0.90373 | 0 | 0.76547 | 1.1404e-06 | 315 | 1.0073 | 0.76547 | 0.76547 | 0.76547 | 0.76547 | 0.76547
| 1000 | 1000 | 2026-02-21 11:55 | 0.89876 | 0 | 0.76681 | 9.6021e-07 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76681 | 0.76681 | 0.76681
| 330000 | 330000 | 2026-02-21 11:55 | 0.90408 | 0 | 0.76536 | 2.6623e-06 | 330 | 1.0078 | 0.76537 | 0.76537 | 0.76536 | 0.76536 | 0.76536
| 1000 | 1000 | 2026-02-21 11:56 | 0.89879 | 0 | 0.76683 | 1.1296e-06 | #NAME? | #NAME? | 0.76684 | 0.76684 | 0.76684 | 0.76683 | 0.76683
| 345000 | 345000 | 2026-02-21 11:56 | 0.90449 | 0 | 0.76531 | 3.752e-06 | 345 | 1.0083 | 0.76531 | 0.76531 | 0.76531 | 0.7653 | 0.76529
| 1000 | 1000 | 2026-02-21 11:56 | 0.8987 | 0 | 0.76675 | 1.6185e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 360000 | 360000 | 2026-02-21 11:57 | 0.90486 | 0 | 0.76522 | 2.5258e-06 | 360 | 1.0089 | 0.76523 | 0.76522 | 0.76522 | 0.76521 | 0.76521
| 1000 | 1000 | 2026-02-21 11:57 | 0.89868 | 0 | 0.76674 | 2.8856e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 375000 | 375000 | 2026-02-21 11:58 | 0.9052 | 0 | 0.76512 | 2.9012e-06 | 375 | 1.0094 | 0.76512 | 0.76513 | 0.76512 | 0.76512 | 0.76511
| 1000 | 1000 | 2026-02-21 11:58 | 0.89873 | 0 | 0.76677 | 5.3088e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76678 | 0.76679
| 390000 | 390000 | 2026-02-21 11:59 | 0.9056 | 0 | 0.76508 | 4.0947e-06 | 390 | 1.0099 | 0.7651 | 0.76509 | 0.76509 | 0.76508 | 0.76507
| 1000 | 1000 | 2026-02-21 11:59 | 0.8987 | 0 | 0.76676 | 2.4915e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 405000 | 405000 | 2026-02-21 11:59 | 0.90597 | 0 | 0.76504 | 3.1491e-06 | 405 | 1.0103 | 0.76504 | 0.76505 | 0.76504 | 0.76503 | 0.76503
| 1000 | 1000 | 2026-02-21 12:00 | 0.89875 | 0 | 0.76679 | 5.4697e-06 | #NAME? | #NAME? | 0.76677 | 0.76679 | 0.76679 | 0.7668 | 0.7668
| 420000 | 420000 | 2026-02-21 12:00 | 0.90635 | 0 | 0.76501 | 1.4908e-06 | 420 | 1.0108 | 0.76502 | 0.76501 | 0.76501 | 0.76501 | 0.76501
| 1000 | 1000 | 2026-02-21 12:01 | 0.8988 | 0 | 0.76684 | 1.9738e-06 | #NAME? | #NAME? | 0.76684 | 0.76684 | 0.76685 | 0.76685 | 0.76685
| 435000 | 435000 | 2026-02-21 12:01 | 0.90666 | 0 | 0.76496 | 1.5918e-06 | 435 | 1.0112 | 0.76496 | 0.76496 | 0.76496 | 0.76496 | 0.76495
| 1000 | 1000 | 2026-02-21 12:01 | 0.89872 | 0 | 0.76677 | 2.7114e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76678 | 0.76678
| 450000 | 450000 | 2026-02-21 12:02 | 0.90689 | 0 | 0.76485 | 2.3218e-06 | 450 | 1.0116 | 0.76485 | 0.76485 | 0.76485 | 0.76484 | 0.76484
| 1000 | 1000 | 2026-02-21 12:02 | 0.89869 | 0 | 0.76675 | 1.3031e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 465000 | 465000 | 2026-02-21 12:03 | 0.90719 | 0 | 0.7648 | 2.5354e-07 | 465 | 1.012 | 0.7648 | 0.7648 | 0.7648 | 0.7648 | 0.7648
| 1000 | 1000 | 2026-02-21 12:03 | 0.89872 | 0 | 0.76677 | 2.7262e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76676 | 0.76677 | 0.76678
| 480000 | 480000 | 2026-02-21 12:04 | 0.90746 | 0 | 0.76478 | 1.0006e-06 | 480 | 1.0124 | 0.76478 | 0.76478 | 0.76478 | 0.76477 | 0.76478
| 1000 | 1000 | 2026-02-21 12:04 | 0.89873 | 0 | 0.76678 | 1.0402e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 495000 | 495000 | 2026-02-21 12:04 | 0.90766 | 0 | 0.76472 | 1.5559e-06 | 495 | 1.0127 | 0.76472 | 0.76471 | 0.76472 | 0.76472 | 0.76472
| 1000 | 1000 | 2026-02-21 12:05 | 0.89878 | 0 | 0.76682 | 1.396e-06 | #NAME? | #NAME? | 0.76682 | 0.76682 | 0.76683 | 0.76682 | 0.76683
| 510000 | 510000 | 2026-02-21 12:05 | 0.90785 | 0 | 0.76468 | 1.6406e-06 | 510 | 1.0129 | 0.76468 | 0.76468 | 0.76467 | 0.76467 | 0.76467
| 1000 | 1000 | 2026-02-21 12:06 | 0.8987 | 0 | 0.76675 | 1.1135e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76675 | 0.76675
| 525000 | 525000 | 2026-02-21 12:06 | 0.90801 | 0 | 0.76463 | 1.5559e-06 | 525 | 1.0132 | 0.76464 | 0.76464 | 0.76464 | 0.76463 | 0.76463
| 1000 | 1000 | 2026-02-21 12:06 | 0.89868 | 0 | 0.76674 | 2.2079e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 540000 | 540000 | 2026-02-21 12:07 | 0.90812 | 0 | 0.76459 | 3.2328e-06 | 540 | 1.0134 | 0.7646 | 0.76459 | 0.76459 | 0.76458 | 0.76458
| 1000 | 1000 | 2026-02-21 12:07 | 0.89869 | 0 | 0.76675 | 9.198e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 555000 | 555000 | 2026-02-21 12:08 | 0.9082 | 0 | 0.76454 | 9.5891e-07 | 555 | 1.0135 | 0.76454 | 0.76454 | 0.76454 | 0.76454 | 0.76454
| 1000 | 1000 | 2026-02-21 12:08 | 0.89869 | 0 | 0.76675 | 1.658e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 570000 | 570000 | 2026-02-21 12:09 | 0.90819 | 0 | 0.76446 | 1.8603e-06 | 570 | 1.0136 | 0.76446 | 0.76446 | 0.76445 | 0.76446 | 0.76446
| 1000 | 1000 | 2026-02-21 12:09 | 0.89871 | 0 | 0.76677 | 1.4229e-06 | #NAME? | #NAME? | 0.76678 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 585000 | 585000 | 2026-02-21 12:09 | 0.90825 | 0 | 0.76446 | 1.0372e-06 | 585 | 1.0136 | 0.76446 | 0.76447 | 0.76446 | 0.76446 | 0.76447
| 1000 | 1000 | 2026-02-21 12:10 | 0.89872 | 0 | 0.76677 | 5.5217e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 600000 | 600000 | 2026-02-21 12:10 | 0.90812 | 0 | 0.76436 | 2.3222e-06 | 600 | 1.0136 | 0.76437 | 0.76437 | 0.76436 | 0.76436 | 0.76435
| 1000 | 1000 | 2026-02-21 12:11 | 0.89867 | 0 | 0.76672 | 2.3439e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 615000 | 615000 | 2026-02-21 12:11 | 0.90801 | 0 | 0.76429 | 7.8947e-07 | 615 | 1.0136 | 0.7643 | 0.76429 | 0.76429 | 0.76429 | 0.76429
| 1000 | 1000 | 2026-02-21 12:11 | 0.89865 | 0 | 0.76671 | 1.2445e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 630000 | 630000 | 2026-02-21 12:12 | 0.90785 | 0 | 0.76423 | 2.2033e-06 | 630 | 1.0135 | 0.76423 | 0.76423 | 0.76423 | 0.76422 | 0.76422
| 1000 | 1000 | 2026-02-21 12:12 | 0.89864 | 0 | 0.76671 | 1.3161e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 645000 | 645000 | 2026-02-21 12:13 | 0.90771 | 0 | 0.76421 | 1.5845e-06 | 645 | 1.0134 | 0.76422 | 0.76421 | 0.76421 | 0.76421 | 0.76421
| 1000 | 1000 | 2026-02-21 12:13 | 0.89864 | 0 | 0.76671 | 1.35e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 660000 | 660000 | 2026-02-21 12:14 | 0.90754 | 0 | 0.76421 | 1.0758e-06 | 660 | 1.0132 | 0.76421 | 0.76421 | 0.76421 | 0.76421 | 0.76421
| 1000 | 1000 | 2026-02-21 12:14 | 0.89871 | 0 | 0.76676 | 6.7579e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 675000 | 675000 | 2026-02-21 12:14 | 0.9072 | 0 | 0.76411 | 3.6192e-06 | 675 | 1.0129 | 0.76412 | 0.76411 | 0.7641 | 0.76411 | 0.7641
| 1000 | 1000 | 2026-02-21 12:15 | 0.89866 | 0 | 0.76672 | 2.238e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 690000 | 690000 | 2026-02-21 12:15 | 0.90692 | 0 | 0.76409 | 6.7304e-07 | 690 | 1.0127 | 0.76409 | 0.76409 | 0.76409 | 0.76409 | 0.76409
| 1000 | 1000 | 2026-02-21 12:16 | 0.89867 | 0 | 0.76673 | 1.8804e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 705000 | 705000 | 2026-02-21 12:16 | 0.90657 | 0 | 0.76404 | 2.078e-06 | 705 | 1.0123 | 0.76403 | 0.76404 | 0.76403 | 0.76404 | 0.76404
| 1000 | 1000 | 2026-02-21 12:16 | 0.89867 | 0 | 0.76673 | 2.5106e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 720000 | 720000 | 2026-02-21 12:17 | 0.90623 | 0 | 0.76404 | 2.0277e-06 | 720 | 1.012 | 0.76404 | 0.76404 | 0.76403 | 0.76404 | 0.76404
| 1000 | 1000 | 2026-02-21 12:17 | 0.89873 | 0 | 0.76678 | 1.0336e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76679
| 735000 | 735000 | 2026-02-21 12:18 | 0.90579 | 0 | 0.764 | 1.7221e-06 | 735 | 1.0115 | 0.76401 | 0.76401 | 0.764 | 0.764 | 0.764
| 1000 | 1000 | 2026-02-21 12:18 | 0.89868 | 0 | 0.76674 | 2.2036e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76673 | 0.76674 | 0.76675
| 750000 | 750000 | 2026-02-21 12:19 | 0.90534 | 0 | 0.76398 | 8.9298e-07 | 750 | 1.0111 | 0.76398 | 0.76398 | 0.76398 | 0.76397 | 0.76397
| 1000 | 1000 | 2026-02-21 12:19 | 0.89869 | 0 | 0.76675 | 1.7992e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 765000 | 765000 | 2026-02-21 12:19 | 0.90477 | 0 | 0.76388 | 2.1371e-06 | 765 | 1.0105 | 0.76389 | 0.76388 | 0.76388 | 0.76388 | 0.76388
| 1000 | 1000 | 2026-02-21 12:20 | 0.89865 | 0 | 0.76671 | 1.9288e-06 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 780000 | 780000 | 2026-02-21 12:20 | 0.90426 | 0 | 0.76387 | 7.4304e-07 | 780 | 1.01 | 0.76388 | 0.76388 | 0.76387 | 0.76388 | 0.76387
| 1000 | 1000 | 2026-02-21 12:21 | 0.89867 | 0 | 0.76673 | 1.7911e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 795000 | 795000 | 2026-02-21 12:21 | 0.90368 | 0 | 0.76384 | 2.0186e-06 | 795 | 1.0094 | 0.76385 | 0.76384 | 0.76384 | 0.76384 | 0.76384
| 1000 | 1000 | 2026-02-21 12:22 | 0.8987 | 0 | 0.76675 | 1.9328e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 810000 | 810000 | 2026-02-21 12:22 | 0.90307 | 0 | 0.7638 | 1.7218e-06 | 810 | 1.0088 | 0.76381 | 0.7638 | 0.7638 | 0.7638 | 0.7638
| 1000 | 1000 | 2026-02-21 12:22 | 0.89867 | 0 | 0.76673 | 8.4381e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 825000 | 825000 | 2026-02-21 12:23 | 0.90247 | 0 | 0.7638 | 3.1657e-06 | 825 | 1.0081 | 0.76381 | 0.7638 | 0.76379 | 0.76379 | 0.76379
| 1000 | 1000 | 2026-02-21 12:23 | 0.89869 | 0 | 0.76675 | 1.1502e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 840000 | 840000 | 2026-02-21 12:24 | 0.90179 | 0 | 0.76376 | 2.3083e-06 | 840 | 1.0074 | 0.76376 | 0.76376 | 0.76376 | 0.76375 | 0.76375
| 1000 | 1000 | 2026-02-21 12:24 | 0.89868 | 0 | 0.76674 | 1.1555e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 855000 | 855000 | 2026-02-21 12:24 | 0.90106 | 0 | 0.76369 | 2.6787e-06 | 855 | 1.0067 | 0.76369 | 0.76369 | 0.76369 | 0.76368 | 0.76368
| 1000 | 1000 | 2026-02-21 12:25 | 0.89865 | 0 | 0.76671 | 1.6017e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 870000 | 870000 | 2026-02-21 12:25 | 0.90037 | 0 | 0.76368 | 1.8187e-06 | 870 | 1.0059 | 0.76368 | 0.76368 | 0.76368 | 0.76368 | 0.76367
| 1000 | 1000 | 2026-02-21 12:26 | 0.89867 | 0 | 0.76673 | 3.0937e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 885000 | 885000 | 2026-02-21 12:26 | 0.89958 | 0 | 0.7636 | 5.5482e-07 | 885 | 1.0051 | 0.7636 | 0.7636 | 0.7636 | 0.7636 | 0.7636
| 1000 | 1000 | 2026-02-21 12:27 | 0.89871 | 0 | 0.76677 | 1.1851e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76676 | 0.76677 | 0.76677
| 900000 | 900000 | 2026-02-21 12:27 | 0.89887 | 0 | 0.76361 | 4.0179e-06 | 900 | 1.0043 | 0.76363 | 0.76362 | 0.76361 | 0.76361 | 0.7636
| 1000 | 1000 | 2026-02-21 12:27 | 0.89866 | 0 | 0.76672 | 2.1865e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76672
| 915000 | 915000 | 2026-02-21 12:28 | 0.89814 | 0 | 0.76361 | 2.7542e-06 | 915 | 1.0035 | 0.76362 | 0.76361 | 0.76361 | 0.76361 | 0.7636
| 1000 | 1000 | 2026-02-21 12:28 | 0.89868 | 0 | 0.76675 | 2.9853e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 930000 | 930000 | 2026-02-21 12:29 | 0.8973 | 0 | 0.76352 | 2.4223e-06 | 930 | 1.0027 | 0.76353 | 0.76352 | 0.76352 | 0.76352 | 0.76352
| 1000 | 1000 | 2026-02-21 12:29 | 0.89867 | 0 | 0.76673 | 7.7874e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 945000 | 945000 | 2026-02-21 12:30 | 0.89647 | 0 | 0.76346 | 3.3545e-06 | 945 | 1.0018 | 0.76347 | 0.76346 | 0.76345 | 0.76345 | 0.76345
| 1000 | 1000 | 2026-02-21 12:30 | 0.89868 | 0 | 0.76674 | 1.3691e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 960000 | 960000 | 2026-02-21 12:30 | 0.89578 | 0 | 0.76351 | 1.7482e-06 | 960 | 1.001 | 0.76351 | 0.76351 | 0.76351 | 0.76351 | 0.7635
| 1000 | 1000 | 2026-02-21 12:31 | 0.89873 | 0 | 0.76678 | 4.3557e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76679 | 0.76679
| 975000 | 975000 | 2026-02-21 12:31 | 0.895 | 0 | 0.76349 | 4.4959e-06 | 975 | 1.0001 | 0.76351 | 0.7635 | 0.76349 | 0.76349 | 0.76348
| 1000 | 1000 | 2026-02-21 12:32 | 0.89874 | 0 | 0.76679 | 1.0204e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.7668
| 990000 | 990000 | 2026-02-21 12:32 | 0.89418 | 0 | 0.76343 | 1.6892e-06 | 990 | 0.9993 | 0.76343 | 0.76343 | 0.76344 | 0.76344 | 0.76343
| 1000 | 1000 | 2026-02-21 12:32 | 0.8988 | 0 | 0.76684 | 9.871e-07 | #NAME? | #NAME? | 0.76684 | 0.76684 | 0.76684 | 0.76685 | 0.76684
| 1005000 | 1005000 | 2026-02-21 12:33 | 0.89341 | 0 | 0.76342 | 3.1723e-06 | 1005 | 0.99846 | 0.76343 | 0.76343 | 0.76342 | 0.76342 | 0.76342
| 1000 | 1000 | 2026-02-21 12:33 | 0.89873 | 0 | 0.76679 | 4.5049e-07 | #NAME? | #NAME? | 0.76679 | 0.76678 | 0.76679 | 0.76678 | 0.76679
| 1020000 | 1020000 | 2026-02-21 12:34 | 0.89257 | 0 | 0.76333 | 1.6451e-06 | 1020 | 0.99764 | 0.76334 | 0.76333 | 0.76333 | 0.76333 | 0.76333
| 1000 | 1000 | 2026-02-21 12:34 | 0.89871 | 0 | 0.76677 | 7.5765e-07 | #NAME? | #NAME? | 0.76678 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 1035000 | 1035000 | 2026-02-21 12:35 | 0.89176 | 0 | 0.76327 | 3.6256e-06 | 1035 | 0.99681 | 0.76328 | 0.76327 | 0.76327 | 0.76326 | 0.76325
| 1000 | 1000 | 2026-02-21 12:35 | 0.89869 | 0 | 0.76674 | 1.7883e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 1050000 | 1050000 | 2026-02-21 12:35 | 0.89116 | 0 | 0.76335 | 1.9096e-06 | 1050 | 0.99602 | 0.76335 | 0.76335 | 0.76335 | 0.76335 | 0.76336
| 1000 | 1000 | 2026-02-21 12:36 | 0.89877 | 0 | 0.76681 | 3.1743e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.76681 | 0.76681 | 0.76682
| 1065000 | 1065000 | 2026-02-21 12:36 | 0.89033 | 0 | 0.76325 | 2.4458e-06 | 1065 | 0.99523 | 0.76326 | 0.76325 | 0.76324 | 0.76325 | 0.76324
| 1000 | 1000 | 2026-02-21 12:37 | 0.89872 | 0 | 0.76676 | 2.76e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 1080000 | 1080000 | 2026-02-21 12:37 | 0.88969 | 0 | 0.76327 | 2.5046e-06 | 1080 | 0.9945 | 0.76327 | 0.76327 | 0.76327 | 0.76326 | 0.76326
| 1000 | 1000 | 2026-02-21 12:37 | 0.89874 | 0 | 0.76679 | 1.5924e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.7668
| 1095000 | 1095000 | 2026-02-21 12:38 | 0.88894 | 0 | 0.76319 | 2.0089e-06 | 1095 | 0.99378 | 0.76319 | 0.76319 | 0.76318 | 0.76318 | 0.76318
| 1000 | 1000 | 2026-02-21 12:38 | 0.89871 | 0 | 0.76677 | 1.6509e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 1110000 | 1110000 | 2026-02-21 12:39 | 0.88821 | 0 | 0.7631 | 1.3752e-06 | 1110 | 0.99307 | 0.7631 | 0.7631 | 0.7631 | 0.7631 | 0.7631
| 1000 | 1000 | 2026-02-21 12:39 | 0.89871 | 0 | 0.76677 | 1.4492e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76676
| 1125000 | 1125000 | 2026-02-21 12:40 | 0.88755 | 0 | 0.76303 | 2.1014e-06 | 1125 | 0.99241 | 0.76304 | 0.76304 | 0.76303 | 0.76303 | 0.76303
| 1000 | 1000 | 2026-02-21 12:40 | 0.89868 | 0 | 0.76674 | 2.4288e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 1140000 | 1140000 | 2026-02-21 12:40 | 0.887 | 0 | 0.76303 | 1.036e-06 | 1140 | 0.9918 | 0.76303 | 0.76303 | 0.76303 | 0.76303 | 0.76303
| 1000 | 1000 | 2026-02-21 12:41 | 0.89875 | 0 | 0.7668 | 3.1447e-06 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.76681 | 0.76681
| 1155000 | 1155000 | 2026-02-21 12:41 | 0.88643 | 0 | 0.763 | 1.5868e-06 | 1155 | 0.99121 | 0.763 | 0.763 | 0.763 | 0.76299 | 0.76299
| 1000 | 1000 | 2026-02-21 12:42 | 0.8987 | 0 | 0.76676 | 2.6005e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76675 | 0.76676 | 0.76677
| 1170000 | 1170000 | 2026-02-21 12:42 | 0.88596 | 0 | 0.76301 | 1.5669e-06 | 1170 | 0.99066 | 0.76301 | 0.76301 | 0.76301 | 0.763 | 0.76301
| 1000 | 1000 | 2026-02-21 12:42 | 0.89872 | 0 | 0.76677 | 2.0362e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 1185000 | 1185000 | 2026-02-21 12:43 | 0.88555 | 0 | 0.76303 | 1.6254e-06 | 1185 | 0.99019 | 0.76302 | 0.76302 | 0.76303 | 0.76303 | 0.76303
| 1000 | 1000 | 2026-02-21 12:43 | 0.89877 | 0 | 0.76682 | 7.3738e-07 | #NAME? | #NAME? | 0.76682 | 0.76682 | 0.76682 | 0.76682 | 0.76682
| 1200000 | 1200000 | 2026-02-21 12:44 | 0.88502 | 0 | 0.76291 | 2.3489e-06 | 1200 | 0.98973 | 0.76292 | 0.76292 | 0.76292 | 0.76291 | 0.76291
| 1000 | 1000 | 2026-02-21 12:44 | 0.89872 | 0 | 0.76678 | 1.0491e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76677 | 0.76678 | 0.76678
| 1215000 | 1215000 | 2026-02-21 12:45 | 0.8846 | 0 | 0.76285 | 2.7897e-06 | 1215 | 0.98935 | 0.76286 | 0.76286 | 0.76285 | 0.76285 | 0.76285
| 1000 | 1000 | 2026-02-21 12:45 | 0.89877 | 0 | 0.76681 | 1.9325e-06 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76681 | 0.76682 | 0.76682
| 1230000 | 1230000 | 2026-02-21 12:45 | 0.88429 | 0 | 0.76284 | 2.0486e-06 | 1230 | 0.98902 | 0.76284 | 0.76284 | 0.76284 | 0.76284 | 0.76283
| 1000 | 1000 | 2026-02-21 12:46 | 0.89869 | 0 | 0.76675 | 2.8399e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 1245000 | 1245000 | 2026-02-21 12:46 | 0.88398 | 0 | 0.7628 | 3.3855e-06 | 1245 | 0.98872 | 0.76281 | 0.7628 | 0.7628 | 0.76279 | 0.76279
| 1000 | 1000 | 2026-02-21 12:47 | 0.89869 | 0 | 0.76675 | 8.9137e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 1260000 | 1260000 | 2026-02-21 12:47 | 0.88365 | 0 | 0.76269 | 3.3916e-06 | 1260 | 0.98849 | 0.7627 | 0.7627 | 0.76269 | 0.76269 | 0.76268
| 1000 | 1000 | 2026-02-21 12:47 | 0.89865 | 0 | 0.76671 | 1.4255e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1275000 | 1275000 | 2026-02-21 12:48 | 0.88349 | 0 | 0.76269 | 8.386e-07 | 1275 | 0.98832 | 0.76269 | 0.76269 | 0.76269 | 0.76268 | 0.76268
| 1000 | 1000 | 2026-02-21 12:48 | 0.89866 | 0 | 0.76673 | 3.8031e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76674
| 1290000 | 1290000 | 2026-02-21 12:49 | 0.88342 | 0 | 0.7627 | 1.8484e-06 | 1290 | 0.98821 | 0.76271 | 0.7627 | 0.7627 | 0.7627 | 0.7627
| 1000 | 1000 | 2026-02-21 12:49 | 0.89875 | 0 | 0.76679 | 1.2824e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.76679
| 1305000 | 1305000 | 2026-02-21 12:50 | 0.88342 | 0 | 0.76275 | 1.7754e-06 | 1305 | 0.98816 | 0.76275 | 0.76275 | 0.76275 | 0.76274 | 0.76274
| 1000 | 1000 | 2026-02-21 12:50 | 0.89881 | 0 | 0.76685 | 1.926e-06 | #NAME? | #NAME? | 0.76685 | 0.76685 | 0.76685 | 0.76685 | 0.76685
| 1320000 | 1320000 | 2026-02-21 12:50 | 0.88325 | 0 | 0.7626 | 3.0797e-06 | 1320 | 0.98816 | 0.76261 | 0.76261 | 0.7626 | 0.76259 | 0.76259
| 1000 | 1000 | 2026-02-21 12:51 | 0.8987 | 0 | 0.76675 | 1.5385e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 1335000 | 1335000 | 2026-02-21 12:51 | 0.8832 | 0 | 0.76249 | 3.7201e-06 | 1335 | 0.98824 | 0.7625 | 0.7625 | 0.76249 | 0.76249 | 0.76248
| 1000 | 1000 | 2026-02-21 12:52 | 0.89864 | 0 | 0.76671 | 2.3527e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 1350000 | 1350000 | 2026-02-21 12:52 | 0.88328 | 0 | 0.76247 | 1.9791e-06 | 1350 | 0.98837 | 0.76248 | 0.76248 | 0.76247 | 0.76247 | 0.76247
| 1000 | 1000 | 2026-02-21 12:52 | 0.89863 | 0 | 0.7667 | 9.0142e-07 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.7667 | 0.76671
| 1365000 | 1365000 | 2026-02-21 12:53 | 0.88344 | 0 | 0.76246 | 2.2263e-06 | 1365 | 0.98857 | 0.76246 | 0.76246 | 0.76246 | 0.76246 | 0.76245
| 1000 | 1000 | 2026-02-21 12:53 | 0.89871 | 0 | 0.76677 | 4.3829e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 1380000 | 1380000 | 2026-02-21 12:54 | 0.88365 | 0 | 0.76246 | 2.249e-06 | 1380 | 0.9888 | 0.76246 | 0.76247 | 0.76246 | 0.76246 | 0.76245
| 1000 | 1000 | 2026-02-21 12:54 | 0.89871 | 0 | 0.76677 | 1.536e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 1395000 | 1395000 | 2026-02-21 12:54 | 0.88389 | 0 | 0.76243 | 4.1224e-06 | 1395 | 0.9891 | 0.76244 | 0.76243 | 0.76243 | 0.76242 | 0.76242
| 1000 | 1000 | 2026-02-21 12:55 | 0.8987 | 0 | 0.76676 | 6.6715e-07 | #NAME? | #NAME? | 0.76676 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 1410000 | 1410000 | 2026-02-21 12:55 | 0.88407 | 0 | 0.76231 | 3.3926e-06 | 1410 | 0.98946 | 0.76232 | 0.76231 | 0.76231 | 0.7623 | 0.7623
| 1000 | 1000 | 2026-02-21 12:56 | 0.89864 | 0 | 0.76671 | 9.2336e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1425000 | 1425000 | 2026-02-21 12:56 | 0.88435 | 0 | 0.76223 | 2.375e-06 | 1425 | 0.98988 | 0.76224 | 0.76224 | 0.76223 | 0.76223 | 0.76222
| 1000 | 1000 | 2026-02-21 12:57 | 0.89862 | 0 | 0.76669 | 1.2552e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 1440000 | 1440000 | 2026-02-21 12:57 | 0.88482 | 0 | 0.76227 | 4.9081e-07 | 1440 | 0.99035 | 0.76227 | 0.76227 | 0.76227 | 0.76227 | 0.76227
| 1000 | 1000 | 2026-02-21 12:57 | 0.89868 | 0 | 0.76673 | 1.7325e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 1455000 | 1455000 | 2026-02-21 12:58 | 0.8852 | 0 | 0.7622 | 3.7127e-06 | 1455 | 0.99088 | 0.76221 | 0.7622 | 0.76219 | 0.76219 | 0.76219
| 1000 | 1000 | 2026-02-21 12:58 | 0.89869 | 0 | 0.76675 | 2.7575e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 1470000 | 1470000 | 2026-02-21 12:59 | 0.88567 | 0 | 0.76217 | 2.1901e-06 | 1470 | 0.99145 | 0.76217 | 0.76217 | 0.76217 | 0.76216 | 0.76216
| 1000 | 1000 | 2026-02-21 12:59 | 0.89872 | 0 | 0.76678 | 1.8364e-06 | #NAME? | #NAME? | 0.76679 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 1485000 | 1485000 | 2026-02-21 13:00 | 0.8862 | 0 | 0.76214 | 1.4412e-06 | 1485 | 0.99207 | 0.76214 | 0.76214 | 0.76213 | 0.76214 | 0.76213
| 1000 | 1000 | 2026-02-21 13:00 | 0.89868 | 0 | 0.76674 | 1.1546e-06 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 1500000 | 1500000 | 2026-02-21 13:00 | 0.88671 | 0 | 0.76207 | 1.9112e-06 | 1500 | 0.99272 | 0.76207 | 0.76207 | 0.76206 | 0.76207 | 0.76206
| 1000 | 1000 | 2026-02-21 13:01 | 0.89866 | 0 | 0.76672 | 2.1538e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 1515000 | 1515000 | 2026-02-21 13:01 | 0.88727 | 0 | 0.76201 | 1.3421e-06 | 1515 | 0.99344 | 0.76202 | 0.76201 | 0.76201 | 0.76201 | 0.76202
| 1000 | 1000 | 2026-02-21 13:02 | 0.89869 | 0 | 0.76675 | 1.5369e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76675 | 0.76676
| 1530000 | 1530000 | 2026-02-21 13:02 | 0.88788 | 0 | 0.76196 | 1.8089e-06 | 1530 | 0.99417 | 0.76196 | 0.76196 | 0.76196 | 0.76196 | 0.76195
| 1000 | 1000 | 2026-02-21 13:02 | 0.89867 | 0 | 0.76673 | 3.3013e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76674
| 1545000 | 1545000 | 2026-02-21 13:03 | 0.88855 | 0 | 0.76195 | 6.8807e-07 | 1545 | 0.99495 | 0.76194 | 0.76195 | 0.76195 | 0.76195 | 0.76195
| 1000 | 1000 | 2026-02-21 13:03 | 0.89874 | 0 | 0.76679 | 1.9944e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.76679 | 0.76679
| 1560000 | 1560000 | 2026-02-21 13:04 | 0.88924 | 0 | 0.76191 | 1.1641e-06 | 1560 | 0.99577 | 0.76191 | 0.76192 | 0.76191 | 0.76191 | 0.76191
| 1000 | 1000 | 2026-02-21 13:04 | 0.89867 | 0 | 0.76673 | 1.1636e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76673 | 0.76673 | 0.76674
| 1575000 | 1575000 | 2026-02-21 13:05 | 0.88989 | 0 | 0.76183 | 1.974e-06 | 1575 | 0.9966 | 0.76183 | 0.76183 | 0.76182 | 0.76183 | 0.76182
| 1000 | 1000 | 2026-02-21 13:05 | 0.89865 | 0 | 0.76671 | 5.1112e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76672 | 0.76673
| 1590000 | 1590000 | 2026-02-21 13:05 | 0.89061 | 0 | 0.76179 | 2.5929e-06 | 1590 | 0.99747 | 0.7618 | 0.76179 | 0.76179 | 0.76179 | 0.76178
| 1000 | 1000 | 2026-02-21 13:06 | 0.89864 | 0 | 0.76671 | 2.0332e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 1605000 | 1605000 | 2026-02-21 13:06 | 0.89135 | 0 | 0.76174 | 1.9716e-06 | 1605 | 0.99835 | 0.76173 | 0.76173 | 0.76173 | 0.76174 | 0.76174
| 1000 | 1000 | 2026-02-21 13:07 | 0.89867 | 0 | 0.76673 | 3.9327e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76673 | 0.76674
| 1620000 | 1620000 | 2026-02-21 13:07 | 0.89216 | 0 | 0.76175 | 2.3478e-06 | 1620 | 0.99925 | 0.76176 | 0.76175 | 0.76175 | 0.76174 | 0.76175
| 1000 | 1000 | 2026-02-21 13:07 | 0.89865 | 0 | 0.76672 | 1.5215e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 1635000 | 1635000 | 2026-02-21 13:08 | 0.89289 | 0 | 0.76167 | 7.0267e-07 | 1635 | 1.0002 | 0.76167 | 0.76168 | 0.76167 | 0.76167 | 0.76167
| 1000 | 1000 | 2026-02-21 13:08 | 0.89865 | 0 | 0.76671 | 1.604e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76671 | 0.76672 | 0.76672
| 1650000 | 1650000 | 2026-02-21 13:09 | 0.89374 | 0 | 0.7617 | 1.3356e-06 | 1650 | 1.0011 | 0.76169 | 0.7617 | 0.7617 | 0.7617 | 0.7617
| 1000 | 1000 | 2026-02-21 13:09 | 0.89875 | 0 | 0.7668 | 4.5559e-07 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.7668
| 1665000 | 1665000 | 2026-02-21 13:10 | 0.89443 | 0 | 0.76159 | 4.1727e-06 | 1665 | 1.002 | 0.7616 | 0.76159 | 0.76159 | 0.76158 | 0.76158
| 1000 | 1000 | 2026-02-21 13:10 | 0.89867 | 0 | 0.76673 | 8.7031e-07 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76673 | 0.76673
| 1680000 | 1680000 | 2026-02-21 13:10 | 0.89521 | 0 | 0.76154 | 2.2485e-06 | 1680 | 1.0029 | 0.76155 | 0.76155 | 0.76154 | 0.76154 | 0.76154
| 1000 | 1000 | 2026-02-21 13:11 | 0.8987 | 0 | 0.76676 | 2.7239e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 1695000 | 1695000 | 2026-02-21 13:11 | 0.8959 | 0 | 0.76143 | 2.6928e-06 | 1695 | 1.0038 | 0.76144 | 0.76144 | 0.76143 | 0.76143 | 0.76143
| 1000 | 1000 | 2026-02-21 13:12 | 0.89865 | 0 | 0.76671 | 3.4928e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76672 | 0.76672 | 0.76672
| 1710000 | 1710000 | 2026-02-21 13:12 | 0.89669 | 0 | 0.7614 | 2.0544e-06 | 1710 | 1.0048 | 0.76141 | 0.76141 | 0.7614 | 0.7614 | 0.7614
| 1000 | 1000 | 2026-02-21 13:12 | 0.89868 | 0 | 0.76674 | 2.5869e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 1725000 | 1725000 | 2026-02-21 13:13 | 0.89747 | 0 | 0.76139 | 3.6473e-06 | 1725 | 1.0057 | 0.76141 | 0.7614 | 0.76139 | 0.76139 | 0.76139
| 1000 | 1000 | 2026-02-21 13:13 | 0.89868 | 0 | 0.76674 | 2.8761e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 1740000 | 1740000 | 2026-02-21 13:14 | 0.8982 | 0 | 0.76134 | 1.2474e-06 | 1740 | 1.0066 | 0.76135 | 0.76134 | 0.76134 | 0.76134 | 0.76134
| 1000 | 1000 | 2026-02-21 13:14 | 0.89867 | 0 | 0.76673 | 3.0962e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76674
| 1755000 | 1755000 | 2026-02-21 13:15 | 0.89885 | 0 | 0.76124 | 1.8497e-06 | 1755 | 1.0074 | 0.76124 | 0.76124 | 0.76124 | 0.76123 | 0.76123
| 1000 | 1000 | 2026-02-21 13:15 | 0.89863 | 0 | 0.7667 | 2.8891e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76671 | 0.7667
| 1770000 | 1770000 | 2026-02-21 13:15 | 0.89952 | 0 | 0.76116 | 1.3726e-06 | 1770 | 1.0083 | 0.76117 | 0.76117 | 0.76116 | 0.76116 | 0.76117
| 1000 | 1000 | 2026-02-21 13:16 | 0.89863 | 0 | 0.7667 | 1.5326e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.76669
| 1785000 | 1785000 | 2026-02-21 13:16 | 0.90017 | 0 | 0.7611 | 3.2915e-06 | 1785 | 1.0091 | 0.76111 | 0.7611 | 0.7611 | 0.7611 | 0.76109
| 1000 | 1000 | 2026-02-21 13:17 | 0.89863 | 0 | 0.76669 | 2.2794e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 1800000 | 1800000 | 2026-02-21 13:17 | 0.90085 | 0 | 0.76109 | 2.0165e-06 | 1800 | 1.0099 | 0.76109 | 0.76109 | 0.76109 | 0.76109 | 0.76108
| 1000 | 1000 | 2026-02-21 13:17 | 0.89867 | 0 | 0.76673 | 1.6595e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 1815000 | 1815000 | 2026-02-21 13:18 | 0.90143 | 0 | 0.76102 | 3.3271e-06 | 1815 | 1.0106 | 0.76103 | 0.76103 | 0.76102 | 0.76101 | 0.76101
| 1000 | 1000 | 2026-02-21 13:18 | 0.89868 | 0 | 0.76673 | 3.0954e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 1830000 | 1830000 | 2026-02-21 13:19 | 0.9021 | 0 | 0.76105 | 1.3224e-06 | 1830 | 1.0113 | 0.76105 | 0.76105 | 0.76105 | 0.76105 | 0.76104
| 1000 | 1000 | 2026-02-21 13:19 | 0.89868 | 0 | 0.76674 | 2.0009e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 1845000 | 1845000 | 2026-02-21 13:20 | 0.90251 | 0 | 0.7609 | 3.1993e-06 | 1845 | 1.012 | 0.76091 | 0.7609 | 0.7609 | 0.7609 | 0.7609
| 1000 | 1000 | 2026-02-21 13:20 | 0.89867 | 0 | 0.76673 | 2.6527e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 1860000 | 1860000 | 2026-02-21 13:20 | 0.90308 | 0 | 0.76091 | 2.0864e-06 | 1860 | 1.0126 | 0.76091 | 0.76091 | 0.76091 | 0.76091 | 0.7609
| 1000 | 1000 | 2026-02-21 13:21 | 0.89868 | 0 | 0.76674 | 1.0922e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 1875000 | 1875000 | 2026-02-21 13:21 | 0.90349 | 0 | 0.76083 | 2.195e-06 | 1875 | 1.0132 | 0.76084 | 0.76083 | 0.76083 | 0.76083 | 0.76083
| 1000 | 1000 | 2026-02-21 13:22 | 0.89867 | 0 | 0.76673 | 1.012e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 1890000 | 1890000 | 2026-02-21 13:22 | 0.9039 | 0 | 0.76079 | 3.5239e-06 | 1890 | 1.0137 | 0.7608 | 0.76079 | 0.76079 | 0.76078 | 0.76078
| 1000 | 1000 | 2026-02-21 13:22 | 0.89867 | 0 | 0.76673 | 1.373e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 1905000 | 1905000 | 2026-02-21 13:23 | 0.90417 | 0 | 0.76066 | 1.3386e-06 | 1905 | 1.0141 | 0.76066 | 0.76067 | 0.76066 | 0.76066 | 0.76066
| 1000 | 1000 | 2026-02-21 13:23 | 0.89865 | 0 | 0.76671 | 3.5906e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 1920000 | 1920000 | 2026-02-21 13:24 | 0.90453 | 0 | 0.76066 | 1.2218e-06 | 1920 | 1.0146 | 0.76066 | 0.76066 | 0.76066 | 0.76066 | 0.76066
| 1000 | 1000 | 2026-02-21 13:24 | 0.89868 | 0 | 0.76674 | 8.7795e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 1935000 | 1935000 | 2026-02-21 13:25 | 0.90477 | 0 | 0.76059 | 1.117e-06 | 1935 | 1.0149 | 0.76059 | 0.76059 | 0.76059 | 0.76059 | 0.76058
| 1000 | 1000 | 2026-02-21 13:25 | 0.89866 | 0 | 0.76672 | 1.7275e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1950000 | 1950000 | 2026-02-21 13:25 | 0.90499 | 0 | 0.76055 | 9.5733e-07 | 1950 | 1.0152 | 0.76055 | 0.76055 | 0.76055 | 0.76055 | 0.76055
| 1000 | 1000 | 2026-02-21 13:26 | 0.89866 | 0 | 0.76673 | 3.6772e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 1965000 | 1965000 | 2026-02-21 13:26 | 0.90521 | 0 | 0.76056 | 1.9689e-06 | 1965 | 1.0154 | 0.76056 | 0.76057 | 0.76056 | 0.76056 | 0.76057
| 1000 | 1000 | 2026-02-21 13:27 | 0.89869 | 0 | 0.76675 | 1.1084e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76675 | 0.76676
| 1980000 | 1980000 | 2026-02-21 13:27 | 0.90524 | 0 | 0.76046 | 2.7678e-06 | 1980 | 1.0156 | 0.76047 | 0.76046 | 0.76046 | 0.76046 | 0.76046
| 1000 | 1000 | 2026-02-21 13:27 | 0.89866 | 0 | 0.76672 | 2.215e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 1995000 | 1995000 | 2026-02-21 13:28 | 0.90531 | 0 | 0.76043 | 9.9672e-07 | 1995 | 1.0157 | 0.76043 | 0.76043 | 0.76044 | 0.76044 | 0.76043
| 1000 | 1000 | 2026-02-21 13:28 | 0.89866 | 0 | 0.76672 | 7.211e-07 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76672 | 0.76673
| 2010000 | 2010000 | 2026-02-21 13:29 | 0.9053 | 0 | 0.76039 | 2.7648e-06 | 2010 | 1.0158 | 0.7604 | 0.7604 | 0.76039 | 0.76038 | 0.76038
| 1000 | 1000 | 2026-02-21 13:29 | 0.89868 | 0 | 0.76674 | 1.3941e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76673 | 0.76674 | 0.76673
| 2025000 | 2025000 | 2026-02-21 13:29 | 0.90513 | 0 | 0.76026 | 2.7791e-06 | 2025 | 1.0158 | 0.76026 | 0.76026 | 0.76026 | 0.76025 | 0.76025
| 1000 | 1000 | 2026-02-21 13:30 | 0.89868 | 0 | 0.76675 | 3.2908e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76676
| 2040000 | 2040000 | 2026-02-21 13:30 | 0.9051 | 0 | 0.7603 | 3.1457e-06 | 2040 | 1.0157 | 0.7603 | 0.7603 | 0.7603 | 0.76029 | 0.76029
| 1000 | 1000 | 2026-02-21 13:31 | 0.8987 | 0 | 0.76676 | 3.8587e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76675 | 0.76677 | 0.76677
| 2055000 | 2055000 | 2026-02-21 13:31 | 0.90483 | 0 | 0.76017 | 2.71e-06 | 2055 | 1.0155 | 0.76017 | 0.76017 | 0.76017 | 0.76016 | 0.76016
| 1000 | 1000 | 2026-02-21 13:32 | 0.89866 | 0 | 0.76672 | 2.1435e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 2070000 | 2070000 | 2026-02-21 13:32 | 0.90456 | 0 | 0.76009 | 2.6613e-06 | 2070 | 1.0153 | 0.76009 | 0.76009 | 0.76009 | 0.76008 | 0.76008
| 1000 | 1000 | 2026-02-21 13:32 | 0.89867 | 0 | 0.76672 | 4.7722e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76674
| 2085000 | 2085000 | 2026-02-21 13:33 | 0.90424 | 0 | 0.76001 | 2.931e-06 | 2085 | 1.0151 | 0.76002 | 0.76001 | 0.76001 | 0.76001 | 0.76
| 1000 | 1000 | 2026-02-21 13:33 | 0.89867 | 0 | 0.76673 | 2.0715e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76673 | 0.76673
| 2100000 | 2100000 | 2026-02-21 13:34 | 0.90396 | 0 | 0.76001 | 9.2182e-07 | 2100 | 1.0148 | 0.76001 | 0.76002 | 0.76001 | 0.76001 | 0.76001
| 1000 | 1000 | 2026-02-21 13:34 | 0.89869 | 0 | 0.76674 | 4.1263e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76676
| 2115000 | 2115000 | 2026-02-21 13:34 | 0.90356 | 0 | 0.75996 | 1.5038e-06 | 2115 | 1.0144 | 0.75996 | 0.75996 | 0.75996 | 0.75996 | 0.75996
| 1000 | 1000 | 2026-02-21 13:35 | 0.89875 | 0 | 0.7668 | 4.559e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.76679 | 0.7668 | 0.76681
| 2130000 | 2130000 | 2026-02-21 13:35 | 0.90317 | 0 | 0.75996 | 3.0455e-06 | 2130 | 1.014 | 0.75996 | 0.75996 | 0.75996 | 0.75995 | 0.75995
| 1000 | 1000 | 2026-02-21 13:36 | 0.89873 | 0 | 0.76678 | 2.2668e-07 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 2145000 | 2145000 | 2026-02-21 13:36 | 0.90259 | 0 | 0.75982 | 1.9479e-06 | 2145 | 1.0135 | 0.75983 | 0.75983 | 0.75982 | 0.75982 | 0.75982
| 1000 | 1000 | 2026-02-21 13:37 | 0.89869 | 0 | 0.76674 | 2.5325e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 2160000 | 2160000 | 2026-02-21 13:37 | 0.90205 | 0 | 0.75978 | 3.5059e-06 | 2160 | 1.0129 | 0.75978 | 0.75978 | 0.75977 | 0.75977 | 0.75976
| 1000 | 1000 | 2026-02-21 13:37 | 0.89868 | 0 | 0.76673 | 2.2935e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 2175000 | 2175000 | 2026-02-21 13:38 | 0.9014 | 0 | 0.75966 | 2.8217e-06 | 2175 | 1.0124 | 0.75967 | 0.75967 | 0.75966 | 0.75966 | 0.75965
| 1000 | 1000 | 2026-02-21 13:38 | 0.89865 | 0 | 0.76671 | 1.941e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2190000 | 2190000 | 2026-02-21 13:39 | 0.90078 | 0 | 0.75961 | 1.7484e-06 | 2190 | 1.0117 | 0.75961 | 0.75961 | 0.75961 | 0.75961 | 0.7596
| 1000 | 1000 | 2026-02-21 13:39 | 0.89866 | 0 | 0.76672 | 2.3644e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76672
| 2205000 | 2205000 | 2026-02-21 13:39 | 0.9002 | 0 | 0.75963 | 1.0126e-06 | 2205 | 1.0111 | 0.75963 | 0.75963 | 0.75963 | 0.75963 | 0.75963
| 1000 | 1000 | 2026-02-21 13:40 | 0.89871 | 0 | 0.76677 | 8.7107e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 2220000 | 2220000 | 2026-02-21 13:40 | 0.89944 | 0 | 0.75952 | 1.1587e-06 | 2220 | 1.0104 | 0.75952 | 0.75952 | 0.75952 | 0.75952 | 0.75952
| 1000 | 1000 | 2026-02-21 13:41 | 0.8987 | 0 | 0.76676 | 1.4253e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 2235000 | 2235000 | 2026-02-21 13:41 | 0.89871 | 0 | 0.75946 | 1.7991e-06 | 2235 | 1.0096 | 0.75946 | 0.75946 | 0.75947 | 0.75947 | 0.75947
| 1000 | 1000 | 2026-02-21 13:42 | 0.8987 | 0 | 0.76676 | 3.8795e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 2250000 | 2250000 | 2026-02-21 13:42 | 0.89786 | 0 | 0.75933 | 3.3148e-06 | 2250 | 1.0088 | 0.75934 | 0.75934 | 0.75933 | 0.75933 | 0.75932
| 1000 | 1000 | 2026-02-21 13:42 | 0.89865 | 0 | 0.76671 | 8.998e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 2265000 | 2265000 | 2026-02-21 13:43 | 0.89701 | 0 | 0.75921 | 1.767e-06 | 2265 | 1.008 | 0.75922 | 0.75922 | 0.75921 | 0.75921 | 0.75921
| 1000 | 1000 | 2026-02-21 13:43 | 0.89867 | 0 | 0.76674 | 4.0356e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 2280000 | 2280000 | 2026-02-21 13:44 | 0.89627 | 0 | 0.75922 | 3.9772e-06 | 2280 | 1.0072 | 0.75923 | 0.75923 | 0.75922 | 0.75921 | 0.75921
| 1000 | 1000 | 2026-02-21 13:44 | 0.89873 | 0 | 0.76679 | 1.4742e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.7668
| 2295000 | 2295000 | 2026-02-21 13:44 | 0.89538 | 0 | 0.75909 | 7.2627e-07 | 2295 | 1.0064 | 0.7591 | 0.75909 | 0.75909 | 0.75909 | 0.75909
| 1000 | 1000 | 2026-02-21 13:45 | 0.89869 | 0 | 0.76675 | 1.0011e-06 | #NAME? | #NAME? | 0.76675 | 0.76674 | 0.76675 | 0.76674 | 0.76675
| 2310000 | 2310000 | 2026-02-21 13:45 | 0.89451 | 0 | 0.75901 | 1.3078e-06 | 2310 | 1.0055 | 0.75902 | 0.75902 | 0.75902 | 0.75901 | 0.75901
| 1000 | 1000 | 2026-02-21 13:46 | 0.89866 | 0 | 0.76672 | 4.1093e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76673 | 0.76673
| 2325000 | 2325000 | 2026-02-21 13:46 | 0.8936 | 0 | 0.75891 | 3.874e-06 | 2325 | 1.0046 | 0.75891 | 0.75891 | 0.75891 | 0.7589 | 0.75889
| 1000 | 1000 | 2026-02-21 13:47 | 0.89865 | 0 | 0.76671 | 2.0456e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76671
| 2340000 | 2340000 | 2026-02-21 13:47 | 0.89278 | 0 | 0.75888 | 2.3931e-06 | 2340 | 1.0037 | 0.75888 | 0.75888 | 0.75888 | 0.75887 | 0.75887
| 1000 | 1000 | 2026-02-21 13:47 | 0.89869 | 0 | 0.76674 | 5.1897e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76675 | 0.76675 | 0.76675
| 2355000 | 2355000 | 2026-02-21 13:48 | 0.89199 | 0 | 0.75888 | 2.3664e-06 | 2355 | 1.0028 | 0.75889 | 0.75888 | 0.75888 | 0.75888 | 0.75887
| 1000 | 1000 | 2026-02-21 13:48 | 0.89868 | 0 | 0.76674 | 2.2394e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 2370000 | 2370000 | 2026-02-21 13:49 | 0.89109 | 0 | 0.75878 | 2.116e-06 | 2370 | 1.002 | 0.75878 | 0.75878 | 0.75878 | 0.75877 | 0.75877
| 1000 | 1000 | 2026-02-21 13:49 | 0.89866 | 0 | 0.76673 | 4.0172e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76674
| 2385000 | 2385000 | 2026-02-21 13:49 | 0.89016 | 0 | 0.75866 | 2.0716e-06 | 2385 | 1.0011 | 0.75866 | 0.75866 | 0.75866 | 0.75866 | 0.75865
| 1000 | 1000 | 2026-02-21 13:50 | 0.89865 | 0 | 0.76672 | 2.1073e-06 | #NAME? | #NAME? | 0.76672 | 0.76671 | 0.76672 | 0.76672 | 0.76673
| 2400000 | 2400000 | 2026-02-21 13:50 | 0.88934 | 0 | 0.75862 | 2.4659e-06 | 2400 | 1.0002 | 0.75863 | 0.75863 | 0.75862 | 0.75862 | 0.75862
| 1000 | 1000 | 2026-02-21 13:51 | 0.89862 | 0 | 0.7667 | 2.0729e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 2415000 | 2415000 | 2026-02-21 13:51 | 0.88843 | 0 | 0.75849 | 2.993e-06 | 2415 | 0.99933 | 0.7585 | 0.7585 | 0.7585 | 0.75849 | 0.75848
| 1000 | 1000 | 2026-02-21 13:52 | 0.89863 | 0 | 0.76669 | 1.7117e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 2430000 | 2430000 | 2026-02-21 13:52 | 0.88757 | 0 | 0.7584 | 1.701e-06 | 2430 | 0.9985 | 0.7584 | 0.7584 | 0.75839 | 0.7584 | 0.7584
| 1000 | 1000 | 2026-02-21 13:52 | 0.8986 | 0 | 0.76668 | 2.9086e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76668 | 0.76667 | 0.76669
| 2445000 | 2445000 | 2026-02-21 13:53 | 0.88682 | 0 | 0.75837 | 1.541e-06 | 2445 | 0.9977 | 0.75837 | 0.75836 | 0.75837 | 0.75837 | 0.75836
| 1000 | 1000 | 2026-02-21 13:53 | 0.89861 | 0 | 0.76669 | 2.5252e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 2460000 | 2460000 | 2026-02-21 13:54 | 0.88604 | 0 | 0.75831 | 3.0201e-06 | 2460 | 0.99689 | 0.75831 | 0.75831 | 0.75831 | 0.7583 | 0.7583
| 1000 | 1000 | 2026-02-21 13:54 | 0.89865 | 0 | 0.76671 | 3.8174e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2475000 | 2475000 | 2026-02-21 13:54 | 0.8854 | 0 | 0.75833 | 9.2455e-07 | 2475 | 0.99615 | 0.75834 | 0.75833 | 0.75833 | 0.75833 | 0.75833
| 1000 | 1000 | 2026-02-21 13:55 | 0.89864 | 0 | 0.7667 | 3.0514e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 2490000 | 2490000 | 2026-02-21 13:55 | 0.88464 | 0 | 0.75823 | 2.8661e-06 | 2490 | 0.99542 | 0.75824 | 0.75823 | 0.75823 | 0.75823 | 0.75822
| 1000 | 1000 | 2026-02-21 13:56 | 0.89865 | 0 | 0.76672 | 1.2996e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 2505000 | 2505000 | 2026-02-21 13:56 | 0.88385 | 0 | 0.75808 | 3.7381e-06 | 2505 | 0.99473 | 0.75809 | 0.75809 | 0.75808 | 0.75808 | 0.75807
| 1000 | 1000 | 2026-02-21 13:57 | 0.89866 | 0 | 0.76672 | 9.377e-07 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76672 | 0.76672
| 2520000 | 2520000 | 2026-02-21 13:57 | 0.88325 | 0 | 0.75805 | 2.5713e-06 | 2520 | 0.99409 | 0.75806 | 0.75806 | 0.75805 | 0.75805 | 0.75805
| 1000 | 1000 | 2026-02-21 13:57 | 0.89863 | 0 | 0.76669 | 1.7873e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 2535000 | 2535000 | 2026-02-21 13:58 | 0.88266 | 0 | 0.758 | 1.239e-06 | 2535 | 0.99349 | 0.758 | 0.758 | 0.758 | 0.758 | 0.758
| 1000 | 1000 | 2026-02-21 13:58 | 0.8987 | 0 | 0.76675 | 1.7369e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 2550000 | 2550000 | 2026-02-21 13:59 | 0.88211 | 0 | 0.75795 | 2.5885e-06 | 2550 | 0.99294 | 0.75796 | 0.75796 | 0.75795 | 0.75795 | 0.75794
| 1000 | 1000 | 2026-02-21 13:59 | 0.89866 | 0 | 0.76672 | 1.7866e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 2565000 | 2565000 | 2026-02-21 13:59 | 0.88153 | 0 | 0.75783 | 2.7048e-06 | 2565 | 0.99245 | 0.75784 | 0.75784 | 0.75783 | 0.75783 | 0.75783
| 1000 | 1000 | 2026-02-21 14:00 | 0.89863 | 0 | 0.7667 | 2.9594e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76669 | 0.7667 | 0.7667
| 2580000 | 2580000 | 2026-02-21 14:00 | 0.88112 | 0 | 0.75783 | 1.9666e-06 | 2580 | 0.992 | 0.75783 | 0.75783 | 0.75782 | 0.75782 | 0.75783
| 1000 | 1000 | 2026-02-21 14:01 | 0.89864 | 0 | 0.76671 | 2.2737e-06 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 2595000 | 2595000 | 2026-02-21 14:01 | 0.88068 | 0 | 0.75774 | 2.6818e-06 | 2595 | 0.9916 | 0.75774 | 0.75774 | 0.75773 | 0.75774 | 0.75773
| 1000 | 1000 | 2026-02-21 14:01 | 0.89866 | 0 | 0.76672 | 2.2502e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 2610000 | 2610000 | 2026-02-21 14:02 | 0.88039 | 0 | 0.75775 | 1.7741e-06 | 2610 | 0.99128 | 0.75775 | 0.75775 | 0.75775 | 0.75775 | 0.75774
| 1000 | 1000 | 2026-02-21 14:02 | 0.89867 | 0 | 0.76673 | 1.4036e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 2625000 | 2625000 | 2026-02-21 14:03 | 0.88 | 0 | 0.75762 | 3.3405e-06 | 2625 | 0.991 | 0.75763 | 0.75762 | 0.75762 | 0.75761 | 0.75761
| 1000 | 1000 | 2026-02-21 14:03 | 0.89864 | 0 | 0.7667 | 2.0684e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 2640000 | 2640000 | 2026-02-21 14:04 | 0.8798 | 0 | 0.75761 | 2.3575e-06 | 2640 | 0.99079 | 0.75761 | 0.75761 | 0.75761 | 0.7576 | 0.7576
| 1000 | 1000 | 2026-02-21 14:04 | 0.89871 | 0 | 0.76677 | 2.0235e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 2655000 | 2655000 | 2026-02-21 14:04 | 0.87956 | 0 | 0.75752 | 3.945e-06 | 2655 | 0.99064 | 0.75753 | 0.75752 | 0.75751 | 0.75751 | 0.75751
| 1000 | 1000 | 2026-02-21 14:05 | 0.89865 | 0 | 0.76671 | 2.0094e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 2670000 | 2670000 | 2026-02-21 14:05 | 0.87944 | 0 | 0.75748 | 3.9592e-06 | 2670 | 0.99056 | 0.75749 | 0.75749 | 0.75748 | 0.75747 | 0.75747
| 1000 | 1000 | 2026-02-21 14:06 | 0.89865 | 0 | 0.76672 | 9.6085e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 2685000 | 2685000 | 2026-02-21 14:06 | 0.87941 | 0 | 0.75746 | 6.2521e-07 | 2685 | 0.99054 | 0.75746 | 0.75746 | 0.75746 | 0.75746 | 0.75746
| 1000 | 1000 | 2026-02-21 14:06 | 0.89864 | 0 | 0.76671 | 1.9012e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 2700000 | 2700000 | 2026-02-21 14:07 | 0.87942 | 0 | 0.75744 | 1.0242e-06 | 2700 | 0.99059 | 0.75744 | 0.75743 | 0.75744 | 0.75744 | 0.75744
| 1000 | 1000 | 2026-02-21 14:07 | 0.89867 | 0 | 0.76674 | 1.8025e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 2715000 | 2715000 | 2026-02-21 14:08 | 0.87945 | 0 | 0.75739 | 3.2817e-06 | 2715 | 0.99068 | 0.75739 | 0.75739 | 0.75739 | 0.75739 | 0.75738
| 1000 | 1000 | 2026-02-21 14:08 | 0.89865 | 0 | 0.76671 | 2.393e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 2730000 | 2730000 | 2026-02-21 14:09 | 0.87956 | 0 | 0.75736 | 2.7933e-06 | 2730 | 0.99083 | 0.75737 | 0.75736 | 0.75736 | 0.75736 | 0.75735
| 1000 | 1000 | 2026-02-21 14:09 | 0.89872 | 0 | 0.76676 | 4.0093e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 2745000 | 2745000 | 2026-02-21 14:09 | 0.87982 | 0 | 0.75741 | 2.0646e-06 | 2745 | 0.99108 | 0.75741 | 0.75741 | 0.75741 | 0.75741 | 0.7574
| 1000 | 1000 | 2026-02-21 14:10 | 0.89872 | 0 | 0.76677 | 2.1619e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76678
| 2760000 | 2760000 | 2026-02-21 14:10 | 0.87996 | 0 | 0.7573 | 1.4326e-06 | 2760 | 0.99138 | 0.75729 | 0.7573 | 0.7573 | 0.7573 | 0.75729
| 1000 | 1000 | 2026-02-21 14:11 | 0.89871 | 0 | 0.76676 | 8.8583e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76676 | 0.76676
| 2775000 | 2775000 | 2026-02-21 14:11 | 0.88016 | 0 | 0.75721 | 3.5662e-06 | 2775 | 0.99173 | 0.75721 | 0.75721 | 0.7572 | 0.7572 | 0.7572
| 1000 | 1000 | 2026-02-21 14:12 | 0.89865 | 0 | 0.76672 | 2.8762e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76673
| 2790000 | 2790000 | 2026-02-21 14:12 | 0.88049 | 0 | 0.75717 | 2.2463e-06 | 2790 | 0.99215 | 0.75717 | 0.75717 | 0.75717 | 0.75716 | 0.75716
| 1000 | 1000 | 2026-02-21 14:12 | 0.89862 | 0 | 0.76669 | 2.9467e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 2805000 | 2805000 | 2026-02-21 14:13 | 0.88082 | 0 | 0.75708 | 1.5433e-06 | 2805 | 0.99262 | 0.75709 | 0.75709 | 0.75708 | 0.75708 | 0.75708
| 1000 | 1000 | 2026-02-21 14:13 | 0.89862 | 0 | 0.76668 | 2.9032e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 2820000 | 2820000 | 2026-02-21 14:14 | 0.88126 | 0 | 0.75706 | 2.9121e-06 | 2820 | 0.99315 | 0.75707 | 0.75706 | 0.75706 | 0.75706 | 0.75706
| 1000 | 1000 | 2026-02-21 14:14 | 0.89862 | 0 | 0.76669 | 2.8156e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 2835000 | 2835000 | 2026-02-21 14:14 | 0.88175 | 0 | 0.75704 | 3.8907e-06 | 2835 | 0.99372 | 0.75705 | 0.75705 | 0.75705 | 0.75704 | 0.75703
| 1000 | 1000 | 2026-02-21 14:15 | 0.89861 | 0 | 0.76668 | 3.5524e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76669
| 2850000 | 2850000 | 2026-02-21 14:15 | 0.88224 | 0 | 0.75698 | 2.8456e-06 | 2850 | 0.99436 | 0.75699 | 0.75699 | 0.75698 | 0.75698 | 0.75697
| 1000 | 1000 | 2026-02-21 14:16 | 0.89869 | 0 | 0.76674 | 4.2692e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 2865000 | 2865000 | 2026-02-21 14:16 | 0.88284 | 0 | 0.75697 | 2.0604e-06 | 2865 | 0.99504 | 0.75698 | 0.75698 | 0.75697 | 0.75697 | 0.75697
| 1000 | 1000 | 2026-02-21 14:17 | 0.89865 | 0 | 0.76671 | 2.4437e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 2880000 | 2880000 | 2026-02-21 14:17 | 0.88349 | 0 | 0.75698 | 2.1367e-06 | 2880 | 0.99576 | 0.75699 | 0.75698 | 0.75698 | 0.75698 | 0.75698
| 1000 | 1000 | 2026-02-21 14:17 | 0.89868 | 0 | 0.76673 | 5.5701e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 2895000 | 2895000 | 2026-02-21 14:18 | 0.88405 | 0 | 0.75688 | 1.7926e-06 | 2895 | 0.99653 | 0.75688 | 0.75688 | 0.75688 | 0.75687 | 0.75687
| 1000 | 1000 | 2026-02-21 14:18 | 0.89865 | 0 | 0.76671 | 2.5381e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 2910000 | 2910000 | 2026-02-21 14:19 | 0.88468 | 0 | 0.75682 | 3.5032e-06 | 2910 | 0.99733 | 0.75683 | 0.75682 | 0.75682 | 0.75681 | 0.75681
| 1000 | 1000 | 2026-02-21 14:19 | 0.89868 | 0 | 0.76673 | 5.8369e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 2925000 | 2925000 | 2026-02-21 14:19 | 0.8855 | 0 | 0.75687 | 1.2071e-06 | 2925 | 0.99818 | 0.75687 | 0.75687 | 0.75687 | 0.75687 | 0.75686
| 1000 | 1000 | 2026-02-21 14:20 | 0.89872 | 0 | 0.76678 | 1.706e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 2940000 | 2940000 | 2026-02-21 14:20 | 0.88622 | 0 | 0.75684 | 3.7822e-06 | 2940 | 0.99904 | 0.75685 | 0.75685 | 0.75684 | 0.75683 | 0.75683
| 1000 | 1000 | 2026-02-21 14:21 | 0.89868 | 0 | 0.76674 | 9.4019e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 2955000 | 2955000 | 2026-02-21 14:21 | 0.88699 | 0 | 0.75681 | 3.7978e-06 | 2955 | 0.99993 | 0.75682 | 0.75682 | 0.75681 | 0.75681 | 0.7568
| 1000 | 1000 | 2026-02-21 14:22 | 0.89871 | 0 | 0.76676 | 2.0475e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 2970000 | 2970000 | 2026-02-21 14:22 | 0.88769 | 0 | 0.75672 | 3.1943e-06 | 2970 | 1.0008 | 0.75673 | 0.75672 | 0.75672 | 0.75672 | 0.75671
| 1000 | 1000 | 2026-02-21 14:22 | 0.89866 | 0 | 0.76673 | 2.028e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 2985000 | 2985000 | 2026-02-21 14:23 | 0.88842 | 0 | 0.75664 | 4.3705e-06 | 2985 | 1.0018 | 0.75665 | 0.75665 | 0.75665 | 0.75663 | 0.75663
| 1000 | 1000 | 2026-02-21 14:23 | 0.89875 | 0 | 0.7668 | 4.4589e-06 | #NAME? | #NAME? | 0.76678 | 0.76679 | 0.7668 | 0.7668 | 0.76681
| 3000000 | 3000000 | 2026-02-21 14:24 | 0.88931 | 0 | 0.75668 | 2.7941e-06 | 3000 | 1.0027 | 0.75669 | 0.75668 | 0.75668 | 0.75668 | 0.75668
| 1000 | 1000 | 2026-02-21 14:24 | 0.89879 | 0 | 0.76683 | 2.747e-06 | #NAME? | #NAME? | 0.76682 | 0.76683 | 0.76683 | 0.76683 | 0.76684
| 3015000 | 3015000 | 2026-02-21 14:24 | 0.89016 | 0 | 0.75668 | 2.5175e-06 | 3015 | 1.0037 | 0.75669 | 0.75669 | 0.75668 | 0.75668 | 0.75668
| 1000 | 1000 | 2026-02-21 14:25 | 0.89875 | 0 | 0.7668 | 5.6267e-07 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.7668
| 3030000 | 3030000 | 2026-02-21 14:25 | 0.8908 | 0 | 0.75652 | 2.9544e-06 | 3030 | 1.0046 | 0.75653 | 0.75652 | 0.75652 | 0.75651 | 0.75651
| 1000 | 1000 | 2026-02-21 14:26 | 0.89867 | 0 | 0.76673 | 6.6167e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 3045000 | 3045000 | 2026-02-21 14:26 | 0.89151 | 0 | 0.75639 | 2.3969e-06 | 3045 | 1.0056 | 0.7564 | 0.75639 | 0.75639 | 0.75639 | 0.75638
| 1000 | 1000 | 2026-02-21 14:27 | 0.89864 | 0 | 0.76671 | 8.3824e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3060000 | 3060000 | 2026-02-21 14:27 | 0.89227 | 0 | 0.75633 | 3.3307e-06 | 3060 | 1.0065 | 0.75634 | 0.75633 | 0.75633 | 0.75632 | 0.75632
| 1000 | 1000 | 2026-02-21 14:27 | 0.89862 | 0 | 0.76669 | 1.6073e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 3075000 | 3075000 | 2026-02-21 14:28 | 0.89303 | 0 | 0.75626 | 1.0629e-06 | 3075 | 1.0075 | 0.75626 | 0.75626 | 0.75626 | 0.75626 | 0.75626
| 1000 | 1000 | 2026-02-21 14:28 | 0.89865 | 0 | 0.76671 | 2.4218e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3090000 | 3090000 | 2026-02-21 14:29 | 0.89377 | 0 | 0.75619 | 2.1285e-06 | 3090 | 1.0084 | 0.75619 | 0.75619 | 0.75619 | 0.75619 | 0.75618
| 1000 | 1000 | 2026-02-21 14:29 | 0.89863 | 0 | 0.76669 | 3.2247e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 3105000 | 3105000 | 2026-02-21 14:30 | 0.89457 | 0 | 0.75618 | 1.5161e-06 | 3105 | 1.0093 | 0.75618 | 0.75618 | 0.75618 | 0.75618 | 0.75618
| 1000 | 1000 | 2026-02-21 14:30 | 0.89868 | 0 | 0.76674 | 1.5363e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76673 | 0.76674 | 0.76674
| 3120000 | 3120000 | 2026-02-21 14:30 | 0.89525 | 0 | 0.7561 | 3.4233e-06 | 3120 | 1.0102 | 0.75611 | 0.7561 | 0.7561 | 0.75609 | 0.75609
| 1000 | 1000 | 2026-02-21 14:31 | 0.89867 | 0 | 0.76673 | 2.1008e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 3135000 | 3135000 | 2026-02-21 14:31 | 0.89589 | 0 | 0.756 | 2.6314e-06 | 3135 | 1.0111 | 0.75601 | 0.75601 | 0.756 | 0.756 | 0.75599
| 1000 | 1000 | 2026-02-21 14:32 | 0.89865 | 0 | 0.76671 | 3.119e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3150000 | 3150000 | 2026-02-21 14:32 | 0.8965 | 0 | 0.7559 | 3.1621e-06 | 3150 | 1.0119 | 0.75591 | 0.7559 | 0.75589 | 0.75589 | 0.75589
| 1000 | 1000 | 2026-02-21 14:32 | 0.89867 | 0 | 0.76673 | 2.6218e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 3165000 | 3165000 | 2026-02-21 14:33 | 0.89717 | 0 | 0.75587 | 2.9595e-06 | 3165 | 1.0127 | 0.75587 | 0.75587 | 0.75587 | 0.75586 | 0.75586
| 1000 | 1000 | 2026-02-21 14:33 | 0.89862 | 0 | 0.7667 | 1.8827e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 3180000 | 3180000 | 2026-02-21 14:34 | 0.89772 | 0 | 0.75578 | 1.6148e-06 | 3180 | 1.0134 | 0.75577 | 0.75578 | 0.75578 | 0.75577 | 0.75577
| 1000 | 1000 | 2026-02-21 14:34 | 0.89866 | 0 | 0.76672 | 2.0129e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 3195000 | 3195000 | 2026-02-21 14:34 | 0.89825 | 0 | 0.75569 | 2.4593e-06 | 3195 | 1.0141 | 0.75569 | 0.75569 | 0.75569 | 0.75568 | 0.75568
| 1000 | 1000 | 2026-02-21 14:35 | 0.89863 | 0 | 0.7667 | 3.0385e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 3210000 | 3210000 | 2026-02-21 14:35 | 0.89875 | 0 | 0.75561 | 1.9521e-06 | 3210 | 1.0148 | 0.75561 | 0.75561 | 0.75561 | 0.7556 | 0.7556
| 1000 | 1000 | 2026-02-21 14:36 | 0.89861 | 0 | 0.76668 | 2.8136e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76668
| 3225000 | 3225000 | 2026-02-21 14:36 | 0.89927 | 0 | 0.75558 | 1.3992e-06 | 3225 | 1.0154 | 0.75559 | 0.75559 | 0.75558 | 0.75558 | 0.75558
| 1000 | 1000 | 2026-02-21 14:37 | 0.89867 | 0 | 0.76673 | 3.5152e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76674
| 3240000 | 3240000 | 2026-02-21 14:37 | 0.89971 | 0 | 0.75553 | 2.8405e-06 | 3240 | 1.016 | 0.75554 | 0.75554 | 0.75553 | 0.75553 | 0.75552
| 1000 | 1000 | 2026-02-21 14:37 | 0.89866 | 0 | 0.76672 | 5.2776e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76673 | 0.76673
| 3255000 | 3255000 | 2026-02-21 14:38 | 0.89997 | 0 | 0.75538 | 3.9565e-06 | 3255 | 1.0165 | 0.75539 | 0.75539 | 0.75538 | 0.75537 | 0.75537
| 1000 | 1000 | 2026-02-21 14:38 | 0.89865 | 0 | 0.76672 | 3.0108e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 3270000 | 3270000 | 2026-02-21 14:39 | 0.90033 | 0 | 0.75534 | 1.0868e-06 | 3270 | 1.017 | 0.75534 | 0.75534 | 0.75534 | 0.75534 | 0.75534
| 1000 | 1000 | 2026-02-21 14:39 | 0.89864 | 0 | 0.76671 | 1.797e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3285000 | 3285000 | 2026-02-21 14:39 | 0.90055 | 0 | 0.75523 | 1.6699e-06 | 3285 | 1.0174 | 0.75523 | 0.75523 | 0.75523 | 0.75522 | 0.75523
| 1000 | 1000 | 2026-02-21 14:40 | 0.89864 | 0 | 0.76671 | 3.0152e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 3300000 | 3300000 | 2026-02-21 14:40 | 0.90078 | 0 | 0.75519 | 1.9021e-06 | 3300 | 1.0177 | 0.75519 | 0.75519 | 0.75519 | 0.75519 | 0.75519
| 1000 | 1000 | 2026-02-21 14:41 | 0.89868 | 0 | 0.76674 | 3.3382e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 3315000 | 3315000 | 2026-02-21 14:41 | 0.90088 | 0 | 0.75506 | 4.1513e-06 | 3315 | 1.0179 | 0.75507 | 0.75507 | 0.75506 | 0.75506 | 0.75505
| 1000 | 1000 | 2026-02-21 14:42 | 0.89864 | 0 | 0.76671 | 2.3116e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 3330000 | 3330000 | 2026-02-21 14:42 | 0.90097 | 0 | 0.75499 | 2.3871e-06 | 3330 | 1.0182 | 0.75499 | 0.75499 | 0.75499 | 0.75498 | 0.75498
| 1000 | 1000 | 2026-02-21 14:42 | 0.89869 | 0 | 0.76676 | 1.3866e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76675 | 0.76675 | 0.76676
| 3345000 | 3345000 | 2026-02-21 14:43 | 0.90098 | 0 | 0.75489 | 1.6566e-06 | 3345 | 1.0183 | 0.75489 | 0.75489 | 0.75489 | 0.75489 | 0.75488
| 1000 | 1000 | 2026-02-21 14:43 | 0.89865 | 0 | 0.76672 | 7.2644e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 3360000 | 3360000 | 2026-02-21 14:44 | 0.90097 | 0 | 0.75482 | 3.7745e-06 | 3360 | 1.0184 | 0.75483 | 0.75482 | 0.75482 | 0.75481 | 0.75481
| 1000 | 1000 | 2026-02-21 14:44 | 0.8987 | 0 | 0.76675 | 3.6605e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 3375000 | 3375000 | 2026-02-21 14:44 | 0.90079 | 0 | 0.75466 | 3.4655e-06 | 3375 | 1.0184 | 0.75467 | 0.75466 | 0.75466 | 0.75466 | 0.75465
| 1000 | 1000 | 2026-02-21 14:45 | 0.89865 | 0 | 0.76671 | 1.4727e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3390000 | 3390000 | 2026-02-21 14:45 | 0.90067 | 0 | 0.7546 | 3.4811e-06 | 3390 | 1.0183 | 0.75461 | 0.7546 | 0.75461 | 0.75459 | 0.75459
| 1000 | 1000 | 2026-02-21 14:46 | 0.89866 | 0 | 0.76671 | 2.2281e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3405000 | 3405000 | 2026-02-21 14:46 | 0.90045 | 0 | 0.75451 | 9.411e-07 | 3405 | 1.0182 | 0.75451 | 0.75451 | 0.75451 | 0.75451 | 0.75451
| 1000 | 1000 | 2026-02-21 14:47 | 0.89869 | 0 | 0.76674 | 9.6676e-07 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 3420000 | 3420000 | 2026-02-21 14:47 | 0.90016 | 0 | 0.7544 | 2.0303e-06 | 3420 | 1.018 | 0.75441 | 0.7544 | 0.7544 | 0.7544 | 0.75439
| 1000 | 1000 | 2026-02-21 14:47 | 0.89864 | 0 | 0.76671 | 2.2196e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 3435000 | 3435000 | 2026-02-21 14:48 | 0.8998 | 0 | 0.75429 | 2.0711e-06 | 3435 | 1.0178 | 0.7543 | 0.75429 | 0.75429 | 0.75429 | 0.75428
| 1000 | 1000 | 2026-02-21 14:48 | 0.89865 | 0 | 0.76672 | 1.72e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 3450000 | 3450000 | 2026-02-21 14:49 | 0.89943 | 0 | 0.75421 | 2.2083e-06 | 3450 | 1.0175 | 0.75422 | 0.75421 | 0.75421 | 0.75421 | 0.75421
| 1000 | 1000 | 2026-02-21 14:49 | 0.89866 | 0 | 0.76672 | 1.4971e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 3465000 | 3465000 | 2026-02-21 14:50 | 0.89898 | 0 | 0.7541 | 3.8973e-06 | 3465 | 1.0171 | 0.75412 | 0.75411 | 0.7541 | 0.7541 | 0.75409
| 1000 | 1000 | 2026-02-21 14:50 | 0.89863 | 0 | 0.7667 | 2.1443e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 3480000 | 3480000 | 2026-02-21 14:50 | 0.8986 | 0 | 0.7541 | 5.4015e-07 | 3480 | 1.0167 | 0.7541 | 0.7541 | 0.7541 | 0.7541 | 0.7541
| 1000 | 1000 | 2026-02-21 14:51 | 0.89871 | 0 | 0.76677 | 2.2439e-06 | #NAME? | #NAME? | 0.76677 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 3495000 | 3495000 | 2026-02-21 14:51 | 0.89799 | 0 | 0.75395 | 1.5428e-06 | 3495 | 1.0162 | 0.75395 | 0.75395 | 0.75395 | 0.75395 | 0.75394
| 1000 | 1000 | 2026-02-21 14:52 | 0.8987 | 0 | 0.76675 | 1.6067e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 3510000 | 3510000 | 2026-02-21 14:52 | 0.89734 | 0 | 0.7538 | 3.3827e-06 | 3510 | 1.0156 | 0.75381 | 0.75381 | 0.7538 | 0.75379 | 0.75379
| 1000 | 1000 | 2026-02-21 14:52 | 0.89863 | 0 | 0.76669 | 2.0909e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 3525000 | 3525000 | 2026-02-21 14:53 | 0.89664 | 0 | 0.75365 | 1.0868e-06 | 3525 | 1.0151 | 0.75366 | 0.75365 | 0.75365 | 0.75365 | 0.75365
| 1000 | 1000 | 2026-02-21 14:53 | 0.89865 | 0 | 0.76671 | 2.2749e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3540000 | 3540000 | 2026-02-21 14:54 | 0.89601 | 0 | 0.75359 | 2.217e-06 | 3540 | 1.0144 | 0.7536 | 0.75359 | 0.75359 | 0.7536 | 0.75359
| 1000 | 1000 | 2026-02-21 14:54 | 0.89868 | 0 | 0.76674 | 3.5695e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 3555000 | 3555000 | 2026-02-21 14:55 | 0.89527 | 0 | 0.75349 | 4.1671e-06 | 3555 | 1.0137 | 0.7535 | 0.75349 | 0.75348 | 0.75348 | 0.75348
| 1000 | 1000 | 2026-02-21 14:55 | 0.89872 | 0 | 0.76677 | 1.6171e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76677
| 3570000 | 3570000 | 2026-02-21 14:55 | 0.89458 | 0 | 0.75344 | 9.544e-07 | 3570 | 1.013 | 0.75344 | 0.75344 | 0.75345 | 0.75345 | 0.75345
| 1000 | 1000 | 2026-02-21 14:56 | 0.89877 | 0 | 0.76682 | 2.8932e-06 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76681 | 0.76682 | 0.76682
| 3585000 | 3585000 | 2026-02-21 14:56 | 0.89372 | 0 | 0.75329 | 3.1098e-06 | 3585 | 1.0122 | 0.7533 | 0.75329 | 0.75329 | 0.75328 | 0.75328
| 1000 | 1000 | 2026-02-21 14:57 | 0.8987 | 0 | 0.76675 | 1.3298e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 3600000 | 3600000 | 2026-02-21 14:57 | 0.89293 | 0 | 0.75322 | 2.6909e-06 | 3600 | 1.0114 | 0.75322 | 0.75322 | 0.75322 | 0.75321 | 0.75321
| 1000 | 1000 | 2026-02-21 14:58 | 0.89866 | 0 | 0.76672 | 3.455e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 3615000 | 3615000 | 2026-02-21 14:58 | 0.89203 | 0 | 0.75307 | 4.1348e-06 | 3615 | 1.0106 | 0.75308 | 0.75307 | 0.75307 | 0.75307 | 0.75306
| 1000 | 1000 | 2026-02-21 14:58 | 0.89863 | 0 | 0.7667 | 2.2182e-06 | #NAME? | #NAME? | 0.7667 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 3630000 | 3630000 | 2026-02-21 14:59 | 0.89114 | 0 | 0.75295 | 2.4143e-06 | 3630 | 1.0098 | 0.75296 | 0.75296 | 0.75295 | 0.75295 | 0.75294
| 1000 | 1000 | 2026-02-21 14:59 | 0.89868 | 0 | 0.76673 | 5.3344e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 3645000 | 3645000 | 2026-02-21 15:00 | 0.89037 | 0 | 0.75295 | 2.8797e-06 | 3645 | 1.0089 | 0.75296 | 0.75296 | 0.75295 | 0.75295 | 0.75295
| 1000 | 1000 | 2026-02-21 15:00 | 0.8987 | 0 | 0.76675 | 2.5352e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 3660000 | 3660000 | 2026-02-21 15:00 | 0.88953 | 0 | 0.7529 | 2.908e-06 | 3660 | 1.008 | 0.75291 | 0.75291 | 0.75291 | 0.7529 | 0.75289
| 1000 | 1000 | 2026-02-21 15:01 | 0.89871 | 0 | 0.76676 | 2.8728e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 3675000 | 3675000 | 2026-02-21 15:01 | 0.88845 | 0 | 0.75266 | 3.8601e-06 | 3675 | 1.0071 | 0.75267 | 0.75266 | 0.75266 | 0.75265 | 0.75265
| 1000 | 1000 | 2026-02-21 15:02 | 0.89865 | 0 | 0.76671 | 1.4596e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 3690000 | 3690000 | 2026-02-21 15:02 | 0.88748 | 0 | 0.75252 | 2.8478e-06 | 3690 | 1.0062 | 0.75253 | 0.75252 | 0.75252 | 0.75251 | 0.75251
| 1000 | 1000 | 2026-02-21 15:03 | 0.89867 | 0 | 0.76673 | 2.0746e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 3705000 | 3705000 | 2026-02-21 15:03 | 0.88659 | 0 | 0.75245 | 2.0481e-06 | 3705 | 1.0053 | 0.75245 | 0.75245 | 0.75245 | 0.75245 | 0.75244
| 1000 | 1000 | 2026-02-21 15:03 | 0.89869 | 0 | 0.76675 | 1.6802e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76675 | 0.76675
| 3720000 | 3720000 | 2026-02-21 15:04 | 0.88572 | 0 | 0.75239 | 2.0874e-06 | 3720 | 1.0044 | 0.7524 | 0.75239 | 0.75239 | 0.75239 | 0.75239
| 1000 | 1000 | 2026-02-21 15:04 | 0.89868 | 0 | 0.76674 | 2.5141e-06 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 3735000 | 3735000 | 2026-02-21 15:05 | 0.88481 | 0 | 0.75228 | 2.0673e-06 | 3735 | 1.0035 | 0.75228 | 0.75228 | 0.75228 | 0.75229 | 0.75229
| 1000 | 1000 | 2026-02-21 15:05 | 0.89869 | 0 | 0.76674 | 5.9885e-07 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76674 | 0.76674
| 3750000 | 3750000 | 2026-02-21 15:06 | 0.88393 | 0 | 0.75221 | 4.6093e-06 | 3750 | 1.0026 | 0.75223 | 0.75222 | 0.75221 | 0.75221 | 0.7522
| 1000 | 1000 | 2026-02-21 15:06 | 0.89865 | 0 | 0.76671 | 2.6649e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 3765000 | 3765000 | 2026-02-21 15:06 | 0.88293 | 0 | 0.75202 | 3.905e-06 | 3765 | 1.0017 | 0.75203 | 0.75202 | 0.75202 | 0.75201 | 0.75201
| 1000 | 1000 | 2026-02-21 15:07 | 0.89863 | 0 | 0.76669 | 3.0385e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 3780000 | 3780000 | 2026-02-21 15:07 | 0.88213 | 0 | 0.75197 | 1.1027e-06 | 3780 | 1.0009 | 0.75197 | 0.75197 | 0.75197 | 0.75197 | 0.75197
| 1000 | 1000 | 2026-02-21 15:08 | 0.89865 | 0 | 0.76671 | 1.4445e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3795000 | 3795000 | 2026-02-21 15:08 | 0.88129 | 0 | 0.75187 | 2.1704e-06 | 3795 | 1 | 0.75187 | 0.75188 | 0.75187 | 0.75187 | 0.75186
| 1000 | 1000 | 2026-02-21 15:08 | 0.89865 | 0 | 0.76671 | 7.7532e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 3810000 | 3810000 | 2026-02-21 15:09 | 0.88042 | 0 | 0.75173 | 1.3977e-06 | 3810 | 0.99924 | 0.75172 | 0.75172 | 0.75172 | 0.75173 | 0.75173
| 1000 | 1000 | 2026-02-21 15:09 | 0.89871 | 0 | 0.76676 | 1.928e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 3825000 | 3825000 | 2026-02-21 15:10 | 0.87966 | 0 | 0.75166 | 1.8068e-06 | 3825 | 0.99847 | 0.75167 | 0.75167 | 0.75166 | 0.75166 | 0.75166
| 1000 | 1000 | 2026-02-21 15:10 | 0.89864 | 0 | 0.76671 | 2.1403e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 3840000 | 3840000 | 2026-02-21 15:11 | 0.87896 | 0 | 0.7516 | 4.6203e-07 | 3840 | 0.99776 | 0.7516 | 0.7516 | 0.7516 | 0.7516 | 0.7516
| 1000 | 1000 | 2026-02-21 15:11 | 0.89868 | 0 | 0.76674 | 3.015e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 3855000 | 3855000 | 2026-02-21 15:11 | 0.87826 | 0 | 0.75151 | 1.9997e-06 | 3855 | 0.99707 | 0.75152 | 0.75151 | 0.75151 | 0.75151 | 0.75151
| 1000 | 1000 | 2026-02-21 15:12 | 0.8987 | 0 | 0.76676 | 5.5268e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76675 | 0.76675
| 3870000 | 3870000 | 2026-02-21 15:12 | 0.87751 | 0 | 0.75136 | 1.3157e-06 | 3870 | 0.99643 | 0.75136 | 0.75136 | 0.75136 | 0.75135 | 0.75136
| 1000 | 1000 | 2026-02-21 15:13 | 0.89879 | 0 | 0.76682 | 3.6094e-06 | #NAME? | #NAME? | 0.76681 | 0.76682 | 0.76682 | 0.76683 | 0.76683
| 3885000 | 3885000 | 2026-02-21 15:13 | 0.87691 | 0 | 0.75129 | 1.0509e-06 | 3885 | 0.99585 | 0.75129 | 0.75129 | 0.75129 | 0.75129 | 0.75129
| 1000 | 1000 | 2026-02-21 15:13 | 0.8987 | 0 | 0.76676 | 3.8546e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 3900000 | 3900000 | 2026-02-21 15:14 | 0.87625 | 0 | 0.75112 | 3.7642e-06 | 3900 | 0.99532 | 0.75113 | 0.75113 | 0.75112 | 0.75112 | 0.75111
| 1000 | 1000 | 2026-02-21 15:14 | 0.89867 | 0 | 0.76674 | 1.5529e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76673 | 0.76674
| 3915000 | 3915000 | 2026-02-21 15:15 | 0.87567 | 0 | 0.75099 | 2.5093e-06 | 3915 | 0.99483 | 0.751 | 0.751 | 0.751 | 0.75099 | 0.75099
| 1000 | 1000 | 2026-02-21 15:15 | 0.89867 | 0 | 0.76673 | 1.774e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 3930000 | 3930000 | 2026-02-21 15:16 | 0.87521 | 0 | 0.75091 | 1.5544e-06 | 3930 | 0.99442 | 0.75091 | 0.75091 | 0.7509 | 0.75091 | 0.75091
| 1000 | 1000 | 2026-02-21 15:16 | 0.89866 | 0 | 0.76673 | 1.6096e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 3945000 | 3945000 | 2026-02-21 15:16 | 0.87466 | 0 | 0.75071 | 2.6624e-06 | 3945 | 0.99405 | 0.75072 | 0.75072 | 0.75071 | 0.75071 | 0.75071
| 1000 | 1000 | 2026-02-21 15:17 | 0.89863 | 0 | 0.7667 | 1.2482e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 3960000 | 3960000 | 2026-02-21 15:17 | 0.87432 | 0 | 0.75065 | 1.994e-06 | 3960 | 0.99375 | 0.75065 | 0.75065 | 0.75065 | 0.75065 | 0.75064
| 1000 | 1000 | 2026-02-21 15:18 | 0.89863 | 0 | 0.7667 | 3.5969e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 3975000 | 3975000 | 2026-02-21 15:18 | 0.87402 | 0 | 0.75056 | 2.14e-06 | 3975 | 0.99353 | 0.75057 | 0.75056 | 0.75056 | 0.75056 | 0.75056
| 1000 | 1000 | 2026-02-21 15:18 | 0.89865 | 0 | 0.76672 | 1.7768e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 3990000 | 3990000 | 2026-02-21 15:19 | 0.87376 | 0 | 0.75047 | 1.541e-06 | 3990 | 0.99335 | 0.75048 | 0.75047 | 0.75047 | 0.75047 | 0.75047
| 1000 | 1000 | 2026-02-21 15:19 | 0.89866 | 0 | 0.76672 | 1.0029e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 4005000 | 4005000 | 2026-02-21 15:20 | 0.87348 | 0 | 0.75031 | 2.212e-06 | 4005 | 0.99324 | 0.75032 | 0.75031 | 0.75031 | 0.7503 | 0.75031
| 1000 | 1000 | 2026-02-21 15:20 | 0.89863 | 0 | 0.7667 | 3.4233e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 4020000 | 4020000 | 2026-02-21 15:21 | 0.87338 | 0 | 0.75024 | 2.0716e-06 | 4020 | 0.99321 | 0.75025 | 0.75025 | 0.75024 | 0.75024 | 0.75024
| 1000 | 1000 | 2026-02-21 15:21 | 0.89868 | 0 | 0.76673 | 5.8328e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 4035000 | 4035000 | 2026-02-21 15:21 | 0.87332 | 0 | 0.75016 | 3.1494e-06 | 4035 | 0.99325 | 0.75017 | 0.75016 | 0.75016 | 0.75016 | 0.75015
| 1000 | 1000 | 2026-02-21 15:22 | 0.89873 | 0 | 0.76678 | 3.9389e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.76678
| 4050000 | 4050000 | 2026-02-21 15:22 | 0.87326 | 0 | 0.75004 | 2.1962e-06 | 4050 | 0.99335 | 0.75005 | 0.75004 | 0.75003 | 0.75003 | 0.75004
| 1000 | 1000 | 2026-02-21 15:23 | 0.89869 | 0 | 0.76675 | 1.2642e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 4065000 | 4065000 | 2026-02-21 15:23 | 0.87322 | 0 | 0.74987 | 2.0811e-06 | 4065 | 0.99352 | 0.74988 | 0.74987 | 0.74987 | 0.74987 | 0.74987
| 1000 | 1000 | 2026-02-21 15:23 | 0.89865 | 0 | 0.76671 | 2.2448e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76672 | 0.76671 | 0.76671
| 4080000 | 4080000 | 2026-02-21 15:24 | 0.87331 | 0 | 0.74976 | 1.4365e-06 | 4080 | 0.99376 | 0.74977 | 0.74977 | 0.74976 | 0.74976 | 0.74976
| 1000 | 1000 | 2026-02-21 15:24 | 0.89868 | 0 | 0.76674 | 2.4977e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 4095000 | 4095000 | 2026-02-21 15:25 | 0.87344 | 0 | 0.74966 | 2.5919e-06 | 4095 | 0.99406 | 0.74967 | 0.74967 | 0.74966 | 0.74965 | 0.74966
| 1000 | 1000 | 2026-02-21 15:25 | 0.89868 | 0 | 0.76674 | 1.1812e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 4110000 | 4110000 | 2026-02-21 15:26 | 0.87364 | 0 | 0.74955 | 3.8341e-06 | 4110 | 0.99443 | 0.74956 | 0.74955 | 0.74955 | 0.74954 | 0.74953
| 1000 | 1000 | 2026-02-21 15:26 | 0.89867 | 0 | 0.76673 | 1.4042e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 4125000 | 4125000 | 2026-02-21 15:26 | 0.87387 | 0 | 0.74942 | 3.5202e-06 | 4125 | 0.99486 | 0.74943 | 0.74942 | 0.74942 | 0.74941 | 0.74941
| 1000 | 1000 | 2026-02-21 15:27 | 0.8987 | 0 | 0.76675 | 3.5347e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76676 | 0.76676
| 4140000 | 4140000 | 2026-02-21 15:27 | 0.8742 | 0 | 0.74933 | 1.3496e-06 | 4140 | 0.99536 | 0.74933 | 0.74933 | 0.74933 | 0.74932 | 0.74932
| 1000 | 1000 | 2026-02-21 15:28 | 0.89866 | 0 | 0.76673 | 8.9761e-07 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76672 | 0.76673 | 0.76673
| 4155000 | 4155000 | 2026-02-21 15:28 | 0.87448 | 0 | 0.74915 | 1.7929e-06 | 4155 | 0.99592 | 0.74915 | 0.74915 | 0.74915 | 0.74915 | 0.74914
| 1000 | 1000 | 2026-02-21 15:29 | 0.8987 | 0 | 0.76676 | 4.5472e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 4170000 | 4170000 | 2026-02-21 15:29 | 0.87492 | 0 | 0.74907 | 4.2548e-06 | 4170 | 0.99653 | 0.74908 | 0.74907 | 0.74907 | 0.74906 | 0.74906
| 1000 | 1000 | 2026-02-21 15:29 | 0.89866 | 0 | 0.76672 | 1.1819e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 4185000 | 4185000 | 2026-02-21 15:30 | 0.87536 | 0 | 0.74894 | 2.567e-06 | 4185 | 0.99719 | 0.74895 | 0.74894 | 0.74894 | 0.74894 | 0.74894
| 1000 | 1000 | 2026-02-21 15:30 | 0.89872 | 0 | 0.76677 | 2.7185e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76678
| 4200000 | 4200000 | 2026-02-21 15:31 | 0.87605 | 0 | 0.749 | 6.389e-07 | 4200 | 0.99791 | 0.749 | 0.749 | 0.749 | 0.749 | 0.749
| 1000 | 1000 | 2026-02-21 15:31 | 0.89873 | 0 | 0.76678 | 1.0337e-06 | #NAME? | #NAME? | 0.76679 | 0.76678 | 0.76678 | 0.76679 | 0.76679
| 4215000 | 4215000 | 2026-02-21 15:31 | 0.87641 | 0 | 0.74873 | 3.9425e-06 | 4215 | 0.99868 | 0.74874 | 0.74874 | 0.74873 | 0.74872 | 0.74872
| 1000 | 1000 | 2026-02-21 15:32 | 0.89866 | 0 | 0.76672 | 1.8356e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 4230000 | 4230000 | 2026-02-21 15:32 | 0.877 | 0 | 0.74863 | 1.657e-06 | 4230 | 0.99948 | 0.74864 | 0.74864 | 0.74863 | 0.74863 | 0.74863
| 1000 | 1000 | 2026-02-21 15:33 | 0.89865 | 0 | 0.76671 | 1.605e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76671
| 4245000 | 4245000 | 2026-02-21 15:33 | 0.87755 | 0 | 0.74847 | 2.6212e-06 | 4245 | 1.0003 | 0.74847 | 0.74847 | 0.74847 | 0.74846 | 0.74846
| 1000 | 1000 | 2026-02-21 15:34 | 0.89863 | 0 | 0.7667 | 2.8845e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 4260000 | 4260000 | 2026-02-21 15:34 | 0.87828 | 0 | 0.74843 | 2.8265e-06 | 4260 | 1.0012 | 0.74843 | 0.74843 | 0.74843 | 0.74842 | 0.74842
| 1000 | 1000 | 2026-02-21 15:34 | 0.89871 | 0 | 0.76676 | 1.6393e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 4275000 | 4275000 | 2026-02-21 15:35 | 0.87886 | 0 | 0.74824 | 3.0261e-06 | 4275 | 1.0021 | 0.74825 | 0.74825 | 0.74824 | 0.74824 | 0.74823
| 1000 | 1000 | 2026-02-21 15:35 | 0.89876 | 0 | 0.76681 | 3.1205e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.76681 | 0.76681
| 4290000 | 4290000 | 2026-02-21 15:36 | 0.8797 | 0 | 0.74826 | 4.6868e-06 | 4290 | 1.0031 | 0.74827 | 0.74827 | 0.74826 | 0.74825 | 0.74825
| 1000 | 1000 | 2026-02-21 15:36 | 0.89872 | 0 | 0.76678 | 1.7997e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76678
| 4305000 | 4305000 | 2026-02-21 15:37 | 0.88035 | 0 | 0.74809 | 1.559e-06 | 4305 | 1.004 | 0.7481 | 0.74809 | 0.74809 | 0.74809 | 0.74809
| 1000 | 1000 | 2026-02-21 15:37 | 0.89867 | 0 | 0.76674 | 2.5662e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 4320000 | 4320000 | 2026-02-21 15:37 | 0.88114 | 0 | 0.74803 | 3.266e-06 | 4320 | 1.005 | 0.74804 | 0.74804 | 0.74803 | 0.74803 | 0.74803
| 1000 | 1000 | 2026-02-21 15:38 | 0.89869 | 0 | 0.76676 | 2.0071e-06 | #NAME? | #NAME? | 0.76676 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 4335000 | 4335000 | 2026-02-21 15:38 | 0.88173 | 0 | 0.7478 | 3.2432e-06 | 4335 | 1.006 | 0.74781 | 0.7478 | 0.7478 | 0.74779 | 0.74779
| 1000 | 1000 | 2026-02-21 15:39 | 0.89867 | 0 | 0.76673 | 5.3618e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76674 | 0.76675
| 4350000 | 4350000 | 2026-02-21 15:39 | 0.88254 | 0 | 0.74775 | 4.8483e-06 | 4350 | 1.007 | 0.74776 | 0.74775 | 0.74775 | 0.74774 | 0.74773
| 1000 | 1000 | 2026-02-21 15:40 | 0.8988 | 0 | 0.76684 | 3.574e-06 | #NAME? | #NAME? | 0.76683 | 0.76684 | 0.76684 | 0.76684 | 0.76685
| 4365000 | 4365000 | 2026-02-21 15:40 | 0.88341 | 0 | 0.74774 | 3.318e-06 | 4365 | 1.008 | 0.74775 | 0.74774 | 0.74773 | 0.74773 | 0.74773
| 1000 | 1000 | 2026-02-21 15:40 | 0.89875 | 0 | 0.7668 | 2.0214e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.76681
| 4380000 | 4380000 | 2026-02-21 15:41 | 0.88405 | 0 | 0.74754 | 1.2486e-06 | 4380 | 1.009 | 0.74754 | 0.74754 | 0.74754 | 0.74753 | 0.74754
| 1000 | 1000 | 2026-02-21 15:41 | 0.8987 | 0 | 0.76676 | 1.3962e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 4395000 | 4395000 | 2026-02-21 15:42 | 0.8847 | 0 | 0.74736 | 2.2008e-06 | 4395 | 1.01 | 0.74736 | 0.74736 | 0.74736 | 0.74735 | 0.74735
| 1000 | 1000 | 2026-02-21 15:42 | 0.89868 | 0 | 0.76674 | 2.9242e-06 | #NAME? | #NAME? | 0.76674 | 0.76673 | 0.76674 | 0.76675 | 0.76675
| 4410000 | 4410000 | 2026-02-21 15:42 | 0.88545 | 0 | 0.74727 | 3.572e-06 | 4410 | 1.011 | 0.74728 | 0.74728 | 0.74727 | 0.74727 | 0.74727
| 1000 | 1000 | 2026-02-21 15:43 | 0.89868 | 0 | 0.76675 | 2.3882e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76676
| 4425000 | 4425000 | 2026-02-21 15:43 | 0.88606 | 0 | 0.74707 | 2.3868e-06 | 4425 | 1.0119 | 0.74707 | 0.74707 | 0.74707 | 0.74706 | 0.74706
| 1000 | 1000 | 2026-02-21 15:44 | 0.89865 | 0 | 0.76672 | 1.551e-06 | #NAME? | #NAME? | 0.76672 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 4440000 | 4440000 | 2026-02-21 15:44 | 0.88678 | 0 | 0.74699 | 2.0044e-06 | 4440 | 1.0129 | 0.74699 | 0.74699 | 0.74699 | 0.74699 | 0.74698
| 1000 | 1000 | 2026-02-21 15:45 | 0.89871 | 0 | 0.76676 | 1.2605e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 4455000 | 4455000 | 2026-02-21 15:45 | 0.8874 | 0 | 0.74684 | 1.9064e-06 | 4455 | 1.0138 | 0.74684 | 0.74684 | 0.74684 | 0.74683 | 0.74683
| 1000 | 1000 | 2026-02-21 15:45 | 0.89866 | 0 | 0.76671 | 2.6924e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 4470000 | 4470000 | 2026-02-21 15:46 | 0.88809 | 0 | 0.74678 | 1.9136e-06 | 4470 | 1.0146 | 0.74678 | 0.74678 | 0.74679 | 0.74678 | 0.74678
| 1000 | 1000 | 2026-02-21 15:46 | 0.89876 | 0 | 0.76681 | 3.3705e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.76681 | 0.76681
| 4485000 | 4485000 | 2026-02-21 15:47 | 0.88863 | 0 | 0.74662 | 3.8851e-06 | 4485 | 1.0155 | 0.74663 | 0.74663 | 0.74663 | 0.74662 | 0.74661
| 1000 | 1000 | 2026-02-21 15:47 | 0.89866 | 0 | 0.76673 | 1.9642e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 4500000 | 4500000 | 2026-02-21 15:48 | 0.88914 | 0 | 0.74647 | 3.7006e-06 | 4500 | 1.0163 | 0.74648 | 0.74647 | 0.74647 | 0.74646 | 0.74646
| 1000 | 1000 | 2026-02-21 15:48 | 0.89866 | 0 | 0.76673 | 1.4747e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 4515000 | 4515000 | 2026-02-21 15:48 | 0.88965 | 0 | 0.74634 | 2.4303e-06 | 4515 | 1.017 | 0.74635 | 0.74634 | 0.74634 | 0.74634 | 0.74634
| 1000 | 1000 | 2026-02-21 15:49 | 0.89867 | 0 | 0.76674 | 3.4203e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 4530000 | 4530000 | 2026-02-21 15:49 | 0.89011 | 0 | 0.74621 | 1.736e-06 | 4530 | 1.0177 | 0.74622 | 0.74621 | 0.74621 | 0.74621 | 0.74621
| 1000 | 1000 | 2026-02-21 15:50 | 0.8987 | 0 | 0.76675 | 1.4893e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76676 | 0.76675
| 4545000 | 4545000 | 2026-02-21 15:50 | 0.89054 | 0 | 0.7461 | 2.0362e-06 | 4545 | 1.0184 | 0.74611 | 0.7461 | 0.7461 | 0.7461 | 0.7461
| 1000 | 1000 | 2026-02-21 15:50 | 0.89867 | 0 | 0.76673 | 1.7616e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 4560000 | 4560000 | 2026-02-21 15:51 | 0.89095 | 0 | 0.74602 | 1.2977e-06 | 4560 | 1.0189 | 0.74601 | 0.74602 | 0.74602 | 0.74602 | 0.74602
| 1000 | 1000 | 2026-02-21 15:51 | 0.89869 | 0 | 0.76675 | 2.2343e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76674 | 0.76675 | 0.76675
| 4575000 | 4575000 | 2026-02-21 15:52 | 0.89129 | 0 | 0.74592 | 2.1321e-06 | 4575 | 1.0195 | 0.74593 | 0.74592 | 0.74592 | 0.74592 | 0.74591
| 1000 | 1000 | 2026-02-21 15:52 | 0.89864 | 0 | 0.76671 | 2.7319e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.7667 | 0.76671 | 0.76672
| 4590000 | 4590000 | 2026-02-21 15:52 | 0.89145 | 0 | 0.74572 | 2.3157e-06 | 4590 | 1.0199 | 0.74573 | 0.74573 | 0.74572 | 0.74572 | 0.74571
| 1000 | 1000 | 2026-02-21 15:53 | 0.89872 | 0 | 0.76678 | 5.3585e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76678 | 0.76678 | 0.76679
| 4605000 | 4605000 | 2026-02-21 15:53 | 0.89178 | 0 | 0.74571 | 1.9229e-06 | 4605 | 1.0203 | 0.74571 | 0.74571 | 0.74571 | 0.7457 | 0.74571
| 1000 | 1000 | 2026-02-21 15:54 | 0.8987 | 0 | 0.76675 | 1.1895e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76676
| 4620000 | 4620000 | 2026-02-21 15:54 | 0.8918 | 0 | 0.74548 | 3.7142e-06 | 4620 | 1.0206 | 0.74549 | 0.74548 | 0.74548 | 0.74547 | 0.74547
| 1000 | 1000 | 2026-02-21 15:55 | 0.89865 | 0 | 0.76671 | 1.9262e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4635000 | 4635000 | 2026-02-21 15:55 | 0.89185 | 0 | 0.74534 | 2.1467e-06 | 4635 | 1.0209 | 0.74534 | 0.74534 | 0.74534 | 0.74533 | 0.74533
| 1000 | 1000 | 2026-02-21 15:55 | 0.89866 | 0 | 0.76672 | 2.018e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 4650000 | 4650000 | 2026-02-21 15:56 | 0.89184 | 0 | 0.74519 | 8.9016e-07 | 4650 | 1.0211 | 0.74519 | 0.74518 | 0.74519 | 0.74519 | 0.74519
| 1000 | 1000 | 2026-02-21 15:56 | 0.89863 | 0 | 0.7667 | 1.8266e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 4665000 | 4665000 | 2026-02-21 15:57 | 0.89188 | 0 | 0.74514 | 9.2241e-07 | 4665 | 1.0212 | 0.74514 | 0.74514 | 0.74514 | 0.74515 | 0.74515
| 1000 | 1000 | 2026-02-21 15:57 | 0.89874 | 0 | 0.7668 | 1.4684e-06 | #NAME? | #NAME? | 0.7668 | 0.7668 | 0.7668 | 0.7668 | 0.7668
| 4680000 | 4680000 | 2026-02-21 15:57 | 0.89185 | 0 | 0.74509 | 3.3563e-06 | 4680 | 1.0212 | 0.7451 | 0.74509 | 0.74509 | 0.74508 | 0.74508
| 1000 | 1000 | 2026-02-21 15:58 | 0.89871 | 0 | 0.76676 | 9.1623e-07 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76676 | 0.76676 | 0.76676
| 4695000 | 4695000 | 2026-02-21 15:58 | 0.89158 | 0 | 0.74489 | 1.2881e-06 | 4695 | 1.0212 | 0.74489 | 0.74488 | 0.74489 | 0.74489 | 0.74488
| 1000 | 1000 | 2026-02-21 15:59 | 0.89868 | 0 | 0.76674 | 1.9871e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76674
| 4710000 | 4710000 | 2026-02-21 15:59 | 0.89137 | 0 | 0.74479 | 3.2201e-06 | 4710 | 1.0211 | 0.7448 | 0.74479 | 0.74479 | 0.74478 | 0.74478
| 1000 | 1000 | 2026-02-21 16:00 | 0.89867 | 0 | 0.76673 | 1.5316e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 4725000 | 4725000 | 2026-02-21 16:00 | 0.891 | 0 | 0.7446 | 1.8317e-06 | 4725 | 1.0209 | 0.74461 | 0.74461 | 0.7446 | 0.7446 | 0.7446
| 1000 | 1000 | 2026-02-21 16:00 | 0.89866 | 0 | 0.76672 | 2.1104e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 4740000 | 4740000 | 2026-02-21 16:01 | 0.89063 | 0 | 0.74448 | 1.6108e-06 | 4740 | 1.0207 | 0.74449 | 0.74449 | 0.74448 | 0.74448 | 0.74448
| 1000 | 1000 | 2026-02-21 16:01 | 0.89863 | 0 | 0.7667 | 3.6544e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.76671 | 0.76671
| 4755000 | 4755000 | 2026-02-21 16:02 | 0.89018 | 0 | 0.74435 | 3.0257e-06 | 4755 | 1.0203 | 0.74436 | 0.74435 | 0.74435 | 0.74435 | 0.74434
| 1000 | 1000 | 2026-02-21 16:02 | 0.89865 | 0 | 0.76671 | 1.1739e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 4770000 | 4770000 | 2026-02-21 16:03 | 0.88961 | 0 | 0.74416 | 3.363e-06 | 4770 | 1.0199 | 0.74417 | 0.74417 | 0.74416 | 0.74416 | 0.74415
| 1000 | 1000 | 2026-02-21 16:03 | 0.89863 | 0 | 0.7667 | 1.605e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 4785000 | 4785000 | 2026-02-21 16:03 | 0.88906 | 0 | 0.74404 | 2.7384e-06 | 4785 | 1.0195 | 0.74405 | 0.74404 | 0.74404 | 0.74404 | 0.74403
| 1000 | 1000 | 2026-02-21 16:04 | 0.89862 | 0 | 0.76669 | 2.5355e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 4800000 | 4800000 | 2026-02-21 16:04 | 0.88847 | 0 | 0.74393 | 1.0165e-06 | 4800 | 1.019 | 0.74393 | 0.74393 | 0.74393 | 0.74393 | 0.74392
| 1000 | 1000 | 2026-02-21 16:05 | 0.89862 | 0 | 0.76669 | 2.2947e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 4815000 | 4815000 | 2026-02-21 16:05 | 0.88785 | 0 | 0.74384 | 2.0006e-06 | 4815 | 1.0184 | 0.74384 | 0.74384 | 0.74384 | 0.74384 | 0.74385
| 1000 | 1000 | 2026-02-21 16:05 | 0.89866 | 0 | 0.76672 | 2.5739e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76673
| 4830000 | 4830000 | 2026-02-21 16:06 | 0.88711 | 0 | 0.7437 | 2.5403e-06 | 4830 | 1.0177 | 0.7437 | 0.7437 | 0.74369 | 0.74369 | 0.74369
| 1000 | 1000 | 2026-02-21 16:06 | 0.89863 | 0 | 0.7667 | 1.1262e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.7667 | 0.7667
| 4845000 | 4845000 | 2026-02-21 16:07 | 0.88646 | 0 | 0.74366 | 3.4927e-06 | 4845 | 1.017 | 0.74367 | 0.74367 | 0.74367 | 0.74366 | 0.74365
| 1000 | 1000 | 2026-02-21 16:07 | 0.89863 | 0 | 0.7667 | 6.6379e-07 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 4860000 | 4860000 | 2026-02-21 16:08 | 0.88553 | 0 | 0.74343 | 2.5781e-06 | 4860 | 1.0163 | 0.74344 | 0.74344 | 0.74343 | 0.74343 | 0.74343
| 1000 | 1000 | 2026-02-21 16:08 | 0.89868 | 0 | 0.76674 | 1.955e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 4875000 | 4875000 | 2026-02-21 16:08 | 0.88467 | 0 | 0.7433 | 1.8882e-06 | 4875 | 1.0155 | 0.7433 | 0.7433 | 0.7433 | 0.74329 | 0.74329
| 1000 | 1000 | 2026-02-21 16:09 | 0.89862 | 0 | 0.7667 | 2.2098e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 4890000 | 4890000 | 2026-02-21 16:09 | 0.88378 | 0 | 0.74317 | 1.4539e-06 | 4890 | 1.0146 | 0.74317 | 0.74317 | 0.74317 | 0.74317 | 0.74316
| 1000 | 1000 | 2026-02-21 16:10 | 0.89867 | 0 | 0.76673 | 3.5119e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 4905000 | 4905000 | 2026-02-21 16:10 | 0.88295 | 0 | 0.74312 | 1.8112e-06 | 4905 | 1.0137 | 0.74312 | 0.74312 | 0.74312 | 0.74312 | 0.74311
| 1000 | 1000 | 2026-02-21 16:10 | 0.89869 | 0 | 0.76675 | 3.0208e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76676
| 4920000 | 4920000 | 2026-02-21 16:11 | 0.882 | 0 | 0.74299 | 3.3623e-06 | 4920 | 1.0128 | 0.743 | 0.74299 | 0.74299 | 0.74298 | 0.74298
| 1000 | 1000 | 2026-02-21 16:11 | 0.89871 | 0 | 0.76677 | 1.9383e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76676 | 0.76677 | 0.76678
| 4935000 | 4935000 | 2026-02-21 16:12 | 0.88106 | 0 | 0.74289 | 2.3115e-06 | 4935 | 1.0119 | 0.74289 | 0.74289 | 0.74289 | 0.74288 | 0.74288
| 1000 | 1000 | 2026-02-21 16:12 | 0.89866 | 0 | 0.76673 | 1.7374e-06 | #NAME? | #NAME? | 0.76673 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 4950000 | 4950000 | 2026-02-21 16:13 | 0.88009 | 0 | 0.74278 | 1.3583e-06 | 4950 | 1.0109 | 0.74278 | 0.74278 | 0.74278 | 0.74278 | 0.74277
| 1000 | 1000 | 2026-02-21 16:13 | 0.89867 | 0 | 0.76673 | 1.2859e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76674
| 4965000 | 4965000 | 2026-02-21 16:13 | 0.87904 | 0 | 0.74262 | 3.7491e-06 | 4965 | 1.0099 | 0.74263 | 0.74262 | 0.74262 | 0.74261 | 0.74261
| 1000 | 1000 | 2026-02-21 16:14 | 0.89865 | 0 | 0.76672 | 1.254e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76672 | 0.76672
| 4980000 | 4980000 | 2026-02-21 16:14 | 0.87799 | 0 | 0.74246 | 1.8906e-06 | 4980 | 1.0089 | 0.74247 | 0.74246 | 0.74246 | 0.74246 | 0.74246
| 1000 | 1000 | 2026-02-21 16:15 | 0.89865 | 0 | 0.76671 | 1.1339e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 4995000 | 4995000 | 2026-02-21 16:15 | 0.87694 | 0 | 0.74231 | 3.7224e-06 | 4995 | 1.0079 | 0.74232 | 0.74232 | 0.74231 | 0.74231 | 0.7423
| 1000 | 1000 | 2026-02-21 16:15 | 0.89869 | 0 | 0.76675 | 3.8184e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76676 | 0.76676
| 5010000 | 5010000 | 2026-02-21 16:16 | 0.87594 | 0 | 0.74221 | 3.1032e-06 | 5010 | 1.0069 | 0.74222 | 0.74221 | 0.7422 | 0.7422 | 0.7422
| 1000 | 1000 | 2026-02-21 16:16 | 0.89867 | 0 | 0.76673 | 1.6752e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 5025000 | 5025000 | 2026-02-21 16:17 | 0.87492 | 0 | 0.74206 | 2.9208e-06 | 5025 | 1.0059 | 0.74207 | 0.74206 | 0.74206 | 0.74206 | 0.74205
| 1000 | 1000 | 2026-02-21 16:17 | 0.89865 | 0 | 0.76672 | 3.0722e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 5040000 | 5040000 | 2026-02-21 16:18 | 0.87397 | 0 | 0.74197 | 1.3738e-06 | 5040 | 1.005 | 0.74197 | 0.74197 | 0.74197 | 0.74197 | 0.74197
| 1000 | 1000 | 2026-02-21 16:18 | 0.89873 | 0 | 0.76678 | 4.9509e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76679 | 0.7668
| 5055000 | 5055000 | 2026-02-21 16:18 | 0.87298 | 0 | 0.74184 | 3.5111e-06 | 5055 | 1.004 | 0.74186 | 0.74185 | 0.74184 | 0.74184 | 0.74184
| 1000 | 1000 | 2026-02-21 16:19 | 0.89867 | 0 | 0.76673 | 1.6194e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76673 | 0.76673
| 5070000 | 5070000 | 2026-02-21 16:19 | 0.87205 | 0 | 0.74174 | 1.7602e-06 | 5070 | 1.0031 | 0.74175 | 0.74174 | 0.74174 | 0.74174 | 0.74174
| 1000 | 1000 | 2026-02-21 16:20 | 0.89875 | 0 | 0.7668 | 4.3317e-06 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.7668 | 0.7668 | 0.76681
| 5085000 | 5085000 | 2026-02-21 16:20 | 0.87119 | 0 | 0.74167 | 4.4012e-06 | 5085 | 1.0022 | 0.74168 | 0.74168 | 0.74167 | 0.74167 | 0.74166
| 1000 | 1000 | 2026-02-21 16:20 | 0.89871 | 0 | 0.76677 | 1.2212e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76677
| 5100000 | 5100000 | 2026-02-21 16:21 | 0.87025 | 0 | 0.74152 | 2.5036e-06 | 5100 | 1.0013 | 0.74152 | 0.74152 | 0.74151 | 0.74151 | 0.74151
| 1000 | 1000 | 2026-02-21 16:21 | 0.8987 | 0 | 0.76676 | 2.1748e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76677
| 5115000 | 5115000 | 2026-02-21 16:22 | 0.86941 | 0 | 0.7414 | 2.5548e-06 | 5115 | 1.0005 | 0.74141 | 0.7414 | 0.7414 | 0.7414 | 0.74139
| 1000 | 1000 | 2026-02-21 16:22 | 0.89868 | 0 | 0.76675 | 7.6295e-07 | #NAME? | #NAME? | 0.76675 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 5130000 | 5130000 | 2026-02-21 16:23 | 0.86859 | 0 | 0.74127 | 8.8251e-07 | 5130 | 0.99974 | 0.74127 | 0.74127 | 0.74127 | 0.74127 | 0.74127
| 1000 | 1000 | 2026-02-21 16:23 | 0.89869 | 0 | 0.76675 | 5.1724e-07 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 5145000 | 5145000 | 2026-02-21 16:23 | 0.86784 | 0 | 0.74116 | 1.2674e-06 | 5145 | 0.99902 | 0.74116 | 0.74116 | 0.74116 | 0.74115 | 0.74115
| 1000 | 1000 | 2026-02-21 16:24 | 0.89872 | 0 | 0.76677 | 9.429e-07 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76678 | 0.76677
| 5160000 | 5160000 | 2026-02-21 16:24 | 0.86712 | 0 | 0.74103 | 2.5228e-06 | 5160 | 0.99836 | 0.74103 | 0.74103 | 0.74103 | 0.74102 | 0.74102
| 1000 | 1000 | 2026-02-21 16:25 | 0.89869 | 0 | 0.76675 | 2.1991e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 5175000 | 5175000 | 2026-02-21 16:25 | 0.86636 | 0 | 0.74082 | 2.7124e-06 | 5175 | 0.99777 | 0.74082 | 0.74083 | 0.74082 | 0.74082 | 0.74081
| 1000 | 1000 | 2026-02-21 16:25 | 0.89864 | 0 | 0.76671 | 2.05e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.7667 | 0.76671
| 5190000 | 5190000 | 2026-02-21 16:26 | 0.86583 | 0 | 0.74075 | 1.2102e-06 | 5190 | 0.99725 | 0.74075 | 0.74075 | 0.74074 | 0.74074 | 0.74074
| 1000 | 1000 | 2026-02-21 16:26 | 0.89863 | 0 | 0.7667 | 1.025e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 5205000 | 5205000 | 2026-02-21 16:27 | 0.86529 | 0 | 0.74061 | 2.2727e-06 | 5205 | 0.9968 | 0.74062 | 0.74062 | 0.74062 | 0.74061 | 0.74061
| 1000 | 1000 | 2026-02-21 16:27 | 0.89878 | 0 | 0.76682 | 5.3652e-06 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76682 | 0.76683 | 0.76684
| 5220000 | 5220000 | 2026-02-21 16:28 | 0.86493 | 0 | 0.74058 | 2.2082e-06 | 5220 | 0.99643 | 0.74059 | 0.74059 | 0.74058 | 0.74058 | 0.74058
| 1000 | 1000 | 2026-02-21 16:28 | 0.89871 | 0 | 0.76676 | 4.7179e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76676 | 0.76677
| 5235000 | 5235000 | 2026-02-21 16:28 | 0.86447 | 0 | 0.74041 | 1.5334e-06 | 5235 | 0.99615 | 0.74041 | 0.74041 | 0.74041 | 0.74041 | 0.7404
| 1000 | 1000 | 2026-02-21 16:29 | 0.89867 | 0 | 0.76674 | 1.6675e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76673 | 0.76674
| 5250000 | 5250000 | 2026-02-21 16:29 | 0.86405 | 0 | 0.7402 | 2.4075e-06 | 5250 | 0.99594 | 0.74021 | 0.7402 | 0.7402 | 0.74019 | 0.74019
| 1000 | 1000 | 2026-02-21 16:30 | 0.89863 | 0 | 0.7667 | 3.0174e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 5265000 | 5265000 | 2026-02-21 16:30 | 0.8638 | 0 | 0.74007 | 2.4604e-06 | 5265 | 0.99583 | 0.74008 | 0.74008 | 0.74008 | 0.74007 | 0.74007
| 1000 | 1000 | 2026-02-21 16:31 | 0.89861 | 0 | 0.76668 | 9.7625e-07 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76668 | 0.76668
| 5280000 | 5280000 | 2026-02-21 16:31 | 0.86363 | 0 | 0.73995 | 3.9753e-06 | 5280 | 0.99579 | 0.73996 | 0.73997 | 0.73996 | 0.73995 | 0.73994
| 1000 | 1000 | 2026-02-21 16:31 | 0.8986 | 0 | 0.76668 | 3.457e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 5295000 | 5295000 | 2026-02-21 16:32 | 0.86354 | 0 | 0.73983 | 1.85e-06 | 5295 | 0.99585 | 0.73983 | 0.73984 | 0.73983 | 0.73983 | 0.73983
| 1000 | 1000 | 2026-02-21 16:32 | 0.89867 | 0 | 0.76673 | 3.116e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 5310000 | 5310000 | 2026-02-21 16:33 | 0.86355 | 0 | 0.73973 | 3.2182e-06 | 5310 | 0.996 | 0.73974 | 0.73973 | 0.73973 | 0.73973 | 0.73972
| 1000 | 1000 | 2026-02-21 16:33 | 0.89873 | 0 | 0.76678 | 1.8582e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76678 | 0.76679 | 0.76679
| 5325000 | 5325000 | 2026-02-21 16:33 | 0.86369 | 0 | 0.73966 | 4.4417e-06 | 5325 | 0.99625 | 0.73967 | 0.73967 | 0.73966 | 0.73966 | 0.73965
| 1000 | 1000 | 2026-02-21 16:34 | 0.89867 | 0 | 0.76674 | 1.2448e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 5340000 | 5340000 | 2026-02-21 16:34 | 0.86374 | 0 | 0.73946 | 4.587e-06 | 5340 | 0.99656 | 0.73948 | 0.73946 | 0.73946 | 0.73946 | 0.73945
| 1000 | 1000 | 2026-02-21 16:35 | 0.89864 | 0 | 0.7667 | 1.4361e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 5355000 | 5355000 | 2026-02-21 16:35 | 0.86397 | 0 | 0.73935 | 4.0769e-07 | 5355 | 0.99698 | 0.73935 | 0.73935 | 0.73935 | 0.73935 | 0.73935
| 1000 | 1000 | 2026-02-21 16:36 | 0.89865 | 0 | 0.7667 | 3.7902e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.76671 | 0.76672
| 5370000 | 5370000 | 2026-02-21 16:36 | 0.86423 | 0 | 0.73919 | 2.5965e-06 | 5370 | 0.9975 | 0.7392 | 0.7392 | 0.7392 | 0.73919 | 0.73919
| 1000 | 1000 | 2026-02-21 16:36 | 0.89869 | 0 | 0.76674 | 5.3514e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76675 | 0.76676
| 5385000 | 5385000 | 2026-02-21 16:37 | 0.8647 | 0 | 0.73915 | 4.6286e-07 | 5385 | 0.99811 | 0.73915 | 0.73914 | 0.73915 | 0.73915 | 0.73914
| 1000 | 1000 | 2026-02-21 16:37 | 0.89879 | 0 | 0.76683 | 2.4414e-06 | #NAME? | #NAME? | 0.76682 | 0.76683 | 0.76683 | 0.76684 | 0.76684
| 5400000 | 5400000 | 2026-02-21 16:38 | 0.86514 | 0 | 0.73903 | 4.858e-06 | 5400 | 0.99877 | 0.73904 | 0.73903 | 0.73902 | 0.73902 | 0.73901
| 1000 | 1000 | 2026-02-21 16:38 | 0.8987 | 0 | 0.76675 | 4.2135e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76676
| 5415000 | 5415000 | 2026-02-21 16:38 | 0.86556 | 0 | 0.73882 | 2.8986e-06 | 5415 | 0.99955 | 0.73883 | 0.73882 | 0.73882 | 0.73881 | 0.73881
| 1000 | 1000 | 2026-02-21 16:39 | 0.89867 | 0 | 0.76673 | 1.6965e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76674
| 5430000 | 5430000 | 2026-02-21 16:39 | 0.86619 | 0 | 0.73873 | 3.6737e-06 | 5430 | 1.0004 | 0.73874 | 0.73874 | 0.73873 | 0.73873 | 0.73872
| 1000 | 1000 | 2026-02-21 16:40 | 0.89865 | 0 | 0.76671 | 1.774e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 5445000 | 5445000 | 2026-02-21 16:40 | 0.86677 | 0 | 0.73855 | 1.6004e-06 | 5445 | 1.0013 | 0.73856 | 0.73855 | 0.73855 | 0.73855 | 0.73855
| 1000 | 1000 | 2026-02-21 16:41 | 0.89864 | 0 | 0.76671 | 3.1975e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76672
| 5460000 | 5460000 | 2026-02-21 16:41 | 0.86758 | 0 | 0.73851 | 3.0329e-07 | 5460 | 1.0023 | 0.73851 | 0.73851 | 0.73851 | 0.73851 | 0.73851
| 1000 | 1000 | 2026-02-21 16:41 | 0.89873 | 0 | 0.76679 | 7.7313e-07 | #NAME? | #NAME? | 0.76679 | 0.76679 | 0.76679 | 0.76679 | 0.76679
| 5475000 | 5475000 | 2026-02-21 16:42 | 0.86831 | 0 | 0.73838 | 4.5606e-06 | 5475 | 1.0033 | 0.73838 | 0.73839 | 0.73838 | 0.73837 | 0.73836
| 1000 | 1000 | 2026-02-21 16:42 | 0.89871 | 0 | 0.76677 | 2.2978e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76677
| 5490000 | 5490000 | 2026-02-21 16:43 | 0.86919 | 0 | 0.73832 | 2.2557e-06 | 5490 | 1.0044 | 0.73833 | 0.73832 | 0.73832 | 0.73832 | 0.73831
| 1000 | 1000 | 2026-02-21 16:43 | 0.89872 | 0 | 0.76677 | 1.3419e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 5505000 | 5505000 | 2026-02-21 16:43 | 0.86995 | 0 | 0.73812 | 1.9147e-06 | 5505 | 1.0056 | 0.73812 | 0.73812 | 0.73812 | 0.73812 | 0.73811
| 1000 | 1000 | 2026-02-21 16:44 | 0.89869 | 0 | 0.76675 | 1.1064e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 5520000 | 5520000 | 2026-02-21 16:44 | 0.87082 | 0 | 0.73799 | 2.5369e-06 | 5520 | 1.0068 | 0.738 | 0.738 | 0.73799 | 0.73799 | 0.73798
| 1000 | 1000 | 2026-02-21 16:45 | 0.89866 | 0 | 0.76672 | 7.9315e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76672
| 5535000 | 5535000 | 2026-02-21 16:45 | 0.87165 | 0 | 0.73781 | 2.7877e-06 | 5535 | 1.008 | 0.73782 | 0.73781 | 0.73781 | 0.7378 | 0.7378
| 1000 | 1000 | 2026-02-21 16:46 | 0.89868 | 0 | 0.76674 | 1.1232e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76674
| 5550000 | 5550000 | 2026-02-21 16:46 | 0.87263 | 0 | 0.73773 | 2.1656e-06 | 5550 | 1.0092 | 0.73773 | 0.73773 | 0.73773 | 0.73773 | 0.73772
| 1000 | 1000 | 2026-02-21 16:46 | 0.89864 | 0 | 0.7667 | 1.6472e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 5565000 | 5565000 | 2026-02-21 16:47 | 0.87356 | 0 | 0.7376 | 2.4992e-06 | 5565 | 1.0105 | 0.7376 | 0.7376 | 0.7376 | 0.73759 | 0.73759
| 1000 | 1000 | 2026-02-21 16:47 | 0.89866 | 0 | 0.76672 | 6.2119e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 5580000 | 5580000 | 2026-02-21 16:48 | 0.87447 | 0 | 0.73745 | 3.0617e-06 | 5580 | 1.0117 | 0.73746 | 0.73745 | 0.73745 | 0.73744 | 0.73744
| 1000 | 1000 | 2026-02-21 16:48 | 0.89864 | 0 | 0.7667 | 2.6643e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 5595000 | 5595000 | 2026-02-21 16:48 | 0.8754 | 0 | 0.73733 | 2.2833e-06 | 5595 | 1.0129 | 0.73733 | 0.73733 | 0.73732 | 0.73732 | 0.73732
| 1000 | 1000 | 2026-02-21 16:49 | 0.89864 | 0 | 0.7667 | 1.1768e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 5610000 | 5610000 | 2026-02-21 16:49 | 0.8763 | 0 | 0.73719 | 2.4613e-06 | 5610 | 1.0142 | 0.73719 | 0.73719 | 0.73719 | 0.73719 | 0.73718
| 1000 | 1000 | 2026-02-21 16:50 | 0.89863 | 0 | 0.76669 | 2.7632e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 5625000 | 5625000 | 2026-02-21 16:50 | 0.87718 | 0 | 0.73706 | 1.8751e-06 | 5625 | 1.0154 | 0.73707 | 0.73706 | 0.73706 | 0.73706 | 0.73706
| 1000 | 1000 | 2026-02-21 16:51 | 0.89872 | 0 | 0.76677 | 3.4438e-06 | #NAME? | #NAME? | 0.76676 | 0.76677 | 0.76677 | 0.76677 | 0.76678
| 5640000 | 5640000 | 2026-02-21 16:51 | 0.87806 | 0 | 0.73697 | 2.9127e-06 | 5640 | 1.0165 | 0.73698 | 0.73697 | 0.73697 | 0.73696 | 0.73696
| 1000 | 1000 | 2026-02-21 16:51 | 0.89866 | 0 | 0.76672 | 6.9105e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 5655000 | 5655000 | 2026-02-21 16:52 | 0.87882 | 0 | 0.7368 | 2.654e-06 | 5655 | 1.0176 | 0.73681 | 0.7368 | 0.7368 | 0.7368 | 0.73679
| 1000 | 1000 | 2026-02-21 16:52 | 0.89866 | 0 | 0.76672 | 2.3938e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 5670000 | 5670000 | 2026-02-21 16:53 | 0.87957 | 0 | 0.73668 | 1.7317e-06 | 5670 | 1.0187 | 0.73668 | 0.73668 | 0.73668 | 0.73667 | 0.73667
| 1000 | 1000 | 2026-02-21 16:53 | 0.89866 | 0 | 0.76672 | 2.6987e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 5685000 | 5685000 | 2026-02-21 16:53 | 0.88019 | 0 | 0.73653 | 5.2942e-06 | 5685 | 1.0196 | 0.73654 | 0.73653 | 0.73652 | 0.73651 | 0.73651
| 1000 | 1000 | 2026-02-21 16:54 | 0.89865 | 0 | 0.76671 | 7.925e-07 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 5700000 | 5700000 | 2026-02-21 16:54 | 0.88087 | 0 | 0.73646 | 1.812e-06 | 5700 | 1.0205 | 0.73645 | 0.73645 | 0.73646 | 0.73646 | 0.73646
| 1000 | 1000 | 2026-02-21 16:55 | 0.89874 | 0 | 0.76679 | 2.7994e-06 | #NAME? | #NAME? | 0.76678 | 0.76678 | 0.76679 | 0.7668 | 0.76679
| 5715000 | 5715000 | 2026-02-21 16:55 | 0.88133 | 0 | 0.73631 | 1.6045e-06 | 5715 | 1.0212 | 0.73631 | 0.73631 | 0.73631 | 0.73631 | 0.7363
| 1000 | 1000 | 2026-02-21 16:56 | 0.89865 | 0 | 0.76672 | 1.3577e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76672 | 0.76672
| 5730000 | 5730000 | 2026-02-21 16:56 | 0.88167 | 0 | 0.73614 | 7.5557e-07 | 5730 | 1.0219 | 0.73614 | 0.73614 | 0.73614 | 0.73614 | 0.73614
| 1000 | 1000 | 2026-02-21 16:56 | 0.89865 | 0 | 0.76672 | 1.8945e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 5745000 | 5745000 | 2026-02-21 16:57 | 0.88192 | 0 | 0.73599 | 3.7745e-06 | 5745 | 1.0223 | 0.73601 | 0.736 | 0.73599 | 0.73599 | 0.73598
| 1000 | 1000 | 2026-02-21 16:57 | 0.89861 | 0 | 0.76668 | 2.362e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 5760000 | 5760000 | 2026-02-21 16:58 | 0.88213 | 0 | 0.73593 | 3.3909e-06 | 5760 | 1.0227 | 0.73594 | 0.73594 | 0.73594 | 0.73593 | 0.73592
| 1000 | 1000 | 2026-02-21 16:58 | 0.89862 | 0 | 0.76669 | 2.6631e-06 | #NAME? | #NAME? | 0.76669 | 0.76668 | 0.76669 | 0.76669 | 0.7667
| 5775000 | 5775000 | 2026-02-21 16:58 | 0.88207 | 0 | 0.73574 | 3.4518e-06 | 5775 | 1.0229 | 0.73575 | 0.73575 | 0.73574 | 0.73574 | 0.73573
| 1000 | 1000 | 2026-02-21 16:59 | 0.89865 | 0 | 0.76671 | 4.0701e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76672 | 0.76672
| 5790000 | 5790000 | 2026-02-21 16:59 | 0.88192 | 0 | 0.73562 | 2.0279e-06 | 5790 | 1.0229 | 0.73563 | 0.73562 | 0.73562 | 0.73562 | 0.73562
| 1000 | 1000 | 2026-02-21 17:00 | 0.89862 | 0 | 0.76669 | 2.4432e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 5805000 | 5805000 | 2026-02-21 17:00 | 0.88158 | 0 | 0.73549 | 1.4774e-06 | 5805 | 1.0227 | 0.73549 | 0.73549 | 0.73549 | 0.73549 | 0.73548
| 1000 | 1000 | 2026-02-21 17:01 | 0.89863 | 0 | 0.7667 | 1.5443e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 5820000 | 5820000 | 2026-02-21 17:01 | 0.88108 | 0 | 0.73537 | 2.1399e-06 | 5820 | 1.0222 | 0.73538 | 0.73537 | 0.73537 | 0.73537 | 0.73536
| 1000 | 1000 | 2026-02-21 17:01 | 0.89861 | 0 | 0.76667 | 2.1003e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76667 | 0.76668 | 0.76668
| 5835000 | 5835000 | 2026-02-21 17:02 | 0.88039 | 0 | 0.73524 | 4.1654e-07 | 5835 | 1.0216 | 0.73524 | 0.73524 | 0.73524 | 0.73524 | 0.73524
| 1000 | 1000 | 2026-02-21 17:02 | 0.89872 | 0 | 0.76678 | 1.783e-06 | #NAME? | #NAME? | 0.76677 | 0.76677 | 0.76678 | 0.76678 | 0.76678
| 5850000 | 5850000 | 2026-02-21 17:03 | 0.8796 | 0 | 0.73521 | 1.4603e-06 | 5850 | 1.0208 | 0.73521 | 0.73521 | 0.73521 | 0.7352 | 0.7352
| 1000 | 1000 | 2026-02-21 17:03 | 0.89867 | 0 | 0.76674 | 2.5789e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76674 | 0.76675
| 5865000 | 5865000 | 2026-02-21 17:04 | 0.87848 | 0 | 0.73507 | 2.3448e-06 | 5865 | 1.0196 | 0.73507 | 0.73507 | 0.73508 | 0.73507 | 0.73506
| 1000 | 1000 | 2026-02-21 17:04 | 0.89868 | 0 | 0.76674 | 2.4389e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 5880000 | 5880000 | 2026-02-21 17:04 | 0.87709 | 0 | 0.73492 | 2.3837e-06 | 5880 | 1.0182 | 0.73493 | 0.73492 | 0.73492 | 0.73491 | 0.73491
| 1000 | 1000 | 2026-02-21 17:05 | 0.89864 | 0 | 0.76671 | 4.5e-06 | #NAME? | #NAME? | 0.76669 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 5895000 | 5895000 | 2026-02-21 17:05 | 0.87548 | 0 | 0.73476 | 2.3082e-06 | 5895 | 1.0166 | 0.73476 | 0.73476 | 0.73475 | 0.73476 | 0.73475
| 1000 | 1000 | 2026-02-21 17:06 | 0.8986 | 0 | 0.76668 | 2.834e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 5910000 | 5910000 | 2026-02-21 17:06 | 0.87367 | 0 | 0.73466 | 1.8561e-06 | 5910 | 1.0146 | 0.73466 | 0.73466 | 0.73466 | 0.73465 | 0.73465
| 1000 | 1000 | 2026-02-21 17:06 | 0.89863 | 0 | 0.76671 | 3.0079e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 5925000 | 5925000 | 2026-02-21 17:07 | 0.87154 | 0 | 0.73449 | 1.1699e-06 | 5925 | 1.0124 | 0.7345 | 0.7345 | 0.73449 | 0.73449 | 0.73449
| 1000 | 1000 | 2026-02-21 17:07 | 0.89863 | 0 | 0.7667 | 7.6574e-07 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 5940000 | 5940000 | 2026-02-21 17:08 | 0.86923 | 0 | 0.73441 | 2.3033e-06 | 5940 | 1.0098 | 0.73441 | 0.73441 | 0.73441 | 0.73441 | 0.7344
| 1000 | 1000 | 2026-02-21 17:08 | 0.89868 | 0 | 0.76674 | 2.7514e-06 | #NAME? | #NAME? | 0.76673 | 0.76674 | 0.76674 | 0.76674 | 0.76674
| 5955000 | 5955000 | 2026-02-21 17:09 | 0.86651 | 0 | 0.73423 | 2.9149e-06 | 5955 | 1.0069 | 0.73423 | 0.73423 | 0.73423 | 0.73422 | 0.73422
| 1000 | 1000 | 2026-02-21 17:09 | 0.89862 | 0 | 0.76669 | 1.9764e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 5970000 | 5970000 | 2026-02-21 17:09 | 0.8636 | 0 | 0.73413 | 2.428e-06 | 5970 | 1.0037 | 0.73413 | 0.73413 | 0.73413 | 0.73413 | 0.73412
| 1000 | 1000 | 2026-02-21 17:10 | 0.89863 | 0 | 0.7667 | 3.9225e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.76671
| 5985000 | 5985000 | 2026-02-21 17:10 | 0.86031 | 0 | 0.73396 | 2.7436e-06 | 5985 | 1.0001 | 0.73397 | 0.73396 | 0.73396 | 0.73395 | 0.73395
| 1000 | 1000 | 2026-02-21 17:11 | 0.89862 | 0 | 0.76669 | 1.4618e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 6000000 | 6000000 | 2026-02-21 17:11 | 0.85681 | 0 | 0.73388 | 1.6121e-06 | 6000 | 0.99611 | 0.73388 | 0.73388 | 0.73388 | 0.73388 | 0.73388
| 1000 | 1000 | 2026-02-21 17:11 | 0.89865 | 0 | 0.76672 | 1.8654e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 6015000 | 6015000 | 2026-02-21 17:12 | 0.85297 | 0 | 0.73377 | 1.8694e-06 | 6015 | 0.99177 | 0.73377 | 0.73378 | 0.73378 | 0.73377 | 0.73377
| 1000 | 1000 | 2026-02-21 17:12 | 0.89871 | 0 | 0.76676 | 4.8856e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76676 | 0.76677 | 0.76678
| 6030000 | 6030000 | 2026-02-21 17:13 | 0.8489 | 0 | 0.73375 | 6.7691e-07 | 6030 | 0.98708 | 0.73375 | 0.73375 | 0.73374 | 0.73374 | 0.73375
| 1000 | 1000 | 2026-02-21 17:13 | 0.89872 | 0 | 0.76678 | 6.4608e-07 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76678
| 6045000 | 6045000 | 2026-02-21 17:14 | 0.84426 | 0 | 0.73353 | 2.3587e-06 | 6045 | 0.98197 | 0.73353 | 0.73353 | 0.73353 | 0.73352 | 0.73352
| 1000 | 1000 | 2026-02-21 17:14 | 0.89867 | 0 | 0.76673 | 2.9267e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76674 | 0.76673
| 6060000 | 6060000 | 2026-02-21 17:14 | 0.83934 | 0 | 0.73335 | 1.4442e-06 | 6060 | 0.9765 | 0.73335 | 0.73335 | 0.73335 | 0.73335 | 0.73335
| 1000 | 1000 | 2026-02-21 17:15 | 0.89864 | 0 | 0.76671 | 2.1505e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 6075000 | 6075000 | 2026-02-21 17:15 | 0.83404 | 0 | 0.73314 | 3.4084e-06 | 6075 | 0.9706 | 0.73315 | 0.73315 | 0.73314 | 0.73314 | 0.73313
| 1000 | 1000 | 2026-02-21 17:16 | 0.89864 | 0 | 0.7667 | 6.1626e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.7667 | 0.7667 | 0.76672
| 6090000 | 6090000 | 2026-02-21 17:16 | 0.82856 | 0 | 0.73308 | 3.3122e-06 | 6090 | 0.9643 | 0.73309 | 0.73309 | 0.73308 | 0.73308 | 0.73308
| 1000 | 1000 | 2026-02-21 17:17 | 0.89866 | 0 | 0.76672 | 1.0599e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 6105000 | 6105000 | 2026-02-21 17:17 | 0.82257 | 0 | 0.73288 | 3.1773e-06 | 6105 | 0.95759 | 0.73289 | 0.73289 | 0.73288 | 0.73288 | 0.73288
| 1000 | 1000 | 2026-02-21 17:17 | 0.89868 | 0 | 0.76674 | 1.1561e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76674 | 0.76674 | 0.76675
| 6120000 | 6120000 | 2026-02-21 17:18 | 0.81639 | 0 | 0.73283 | 2.8299e-06 | 6120 | 0.95047 | 0.73284 | 0.73284 | 0.73283 | 0.73283 | 0.73283
| 1000 | 1000 | 2026-02-21 17:18 | 0.89864 | 0 | 0.76671 | 2.602e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 6135000 | 6135000 | 2026-02-21 17:19 | 0.80969 | 0 | 0.73264 | 3.4279e-06 | 6135 | 0.94291 | 0.73265 | 0.73265 | 0.73264 | 0.73263 | 0.73263
| 1000 | 1000 | 2026-02-21 17:19 | 0.89865 | 0 | 0.76671 | 1.4306e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 6150000 | 6150000 | 2026-02-21 17:19 | 0.80273 | 0 | 0.73253 | 2.591e-06 | 6150 | 0.93494 | 0.73254 | 0.73254 | 0.73253 | 0.73253 | 0.73252
| 1000 | 1000 | 2026-02-21 17:20 | 0.89871 | 0 | 0.76677 | 2.6229e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 6165000 | 6165000 | 2026-02-21 17:20 | 0.79535 | 0 | 0.73238 | 3.3165e-06 | 6165 | 0.92655 | 0.73239 | 0.73238 | 0.73238 | 0.73237 | 0.73237
| 1000 | 1000 | 2026-02-21 17:21 | 0.89865 | 0 | 0.76671 | 1.3792e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76671 | 0.76672
| 6180000 | 6180000 | 2026-02-21 17:21 | 0.78767 | 0 | 0.73228 | 2.3193e-06 | 6180 | 0.91772 | 0.73228 | 0.73228 | 0.73228 | 0.73228 | 0.73227
| 1000 | 1000 | 2026-02-21 17:21 | 0.89866 | 0 | 0.76672 | 3.0927e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 6195000 | 6195000 | 2026-02-21 17:22 | 0.77959 | 0 | 0.73214 | 3.9283e-06 | 6195 | 0.90848 | 0.73215 | 0.73215 | 0.73214 | 0.73213 | 0.73213
| 1000 | 1000 | 2026-02-21 17:22 | 0.89864 | 0 | 0.7667 | 1.9579e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 6210000 | 6210000 | 2026-02-21 17:23 | 0.77117 | 0 | 0.73201 | 2.9625e-06 | 6210 | 0.89882 | 0.73202 | 0.73202 | 0.73202 | 0.73201 | 0.73201
| 1000 | 1000 | 2026-02-21 17:23 | 0.89863 | 0 | 0.7667 | 2.4619e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 6225000 | 6225000 | 2026-02-21 17:24 | 0.76235 | 0 | 0.73187 | 2.9175e-06 | 6225 | 0.88872 | 0.73188 | 0.73187 | 0.73187 | 0.73186 | 0.73186
| 1000 | 1000 | 2026-02-21 17:24 | 0.89861 | 0 | 0.76668 | 2.4507e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 6240000 | 6240000 | 2026-02-21 17:24 | 0.75323 | 0 | 0.73177 | 3.2681e-06 | 6240 | 0.8782 | 0.73178 | 0.73177 | 0.73178 | 0.73177 | 0.73176
| 1000 | 1000 | 2026-02-21 17:25 | 0.89863 | 0 | 0.76669 | 2.3504e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.76669 | 0.7667
| 6255000 | 6255000 | 2026-02-21 17:25 | 0.74367 | 0 | 0.73158 | 2.15e-06 | 6255 | 0.86728 | 0.73159 | 0.73158 | 0.73158 | 0.73158 | 0.73158
| 1000 | 1000 | 2026-02-21 17:26 | 0.8987 | 0 | 0.76675 | 3.0338e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76675 | 0.76675 | 0.76675
| 6270000 | 6270000 | 2026-02-21 17:26 | 0.73391 | 0 | 0.73153 | 2.0084e-06 | 6270 | 0.85596 | 0.73154 | 0.73153 | 0.73153 | 0.73153 | 0.73152
| 1000 | 1000 | 2026-02-21 17:27 | 0.89866 | 0 | 0.76673 | 8.6836e-07 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 6285000 | 6285000 | 2026-02-21 17:27 | 0.72364 | 0 | 0.73132 | 1.838e-06 | 6285 | 0.84422 | 0.73133 | 0.73133 | 0.73133 | 0.73132 | 0.73132
| 1000 | 1000 | 2026-02-21 17:27 | 0.89861 | 0 | 0.76668 | 3.319e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76668 | 0.76669
| 6300000 | 6300000 | 2026-02-21 17:28 | 0.71309 | 0 | 0.73117 | 2.6446e-06 | 6300 | 0.8321 | 0.73118 | 0.73117 | 0.73117 | 0.73117 | 0.73117
| 1000 | 1000 | 2026-02-21 17:28 | 0.89868 | 0 | 0.76674 | 1.7132e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76674
| 6315000 | 6315000 | 2026-02-21 17:29 | 0.70225 | 0 | 0.73104 | 2.918e-06 | 6315 | 0.81958 | 0.73105 | 0.73105 | 0.73104 | 0.73104 | 0.73104
| 1000 | 1000 | 2026-02-21 17:29 | 0.89862 | 0 | 0.76669 | 1.622e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 6330000 | 6330000 | 2026-02-21 17:29 | 0.69109 | 0 | 0.73092 | 2.9601e-06 | 6330 | 0.80668 | 0.73093 | 0.73093 | 0.73092 | 0.73092 | 0.73091
| 1000 | 1000 | 2026-02-21 17:30 | 0.89864 | 0 | 0.76671 | 2.3699e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 6345000 | 6345000 | 2026-02-21 17:30 | 0.67961 | 0 | 0.73081 | 1.9567e-06 | 6345 | 0.79342 | 0.73082 | 0.73081 | 0.73081 | 0.73081 | 0.73081
| 1000 | 1000 | 2026-02-21 17:31 | 0.89864 | 0 | 0.7667 | 1.4824e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 6360000 | 6360000 | 2026-02-21 17:31 | 0.66785 | 0 | 0.7307 | 3.1006e-06 | 6360 | 0.7798 | 0.7307 | 0.73069 | 0.73069 | 0.7307 | 0.73071
| 1000 | 1000 | 2026-02-21 17:32 | 0.89863 | 0 | 0.7667 | 2.6546e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.7667 | 0.7667
| 6375000 | 6375000 | 2026-02-21 17:32 | 0.65572 | 0 | 0.73052 | 2.8165e-06 | 6375 | 0.76582 | 0.73052 | 0.73052 | 0.73052 | 0.73051 | 0.73051
| 1000 | 1000 | 2026-02-21 17:32 | 0.89862 | 0 | 0.76669 | 3.059e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.7667
| 6390000 | 6390000 | 2026-02-21 17:33 | 0.64336 | 0 | 0.73039 | 1.727e-06 | 6390 | 0.75153 | 0.73039 | 0.73039 | 0.73038 | 0.73039 | 0.73039
| 1000 | 1000 | 2026-02-21 17:33 | 0.89863 | 0 | 0.7667 | 1.5743e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.7667
| 6405000 | 6405000 | 2026-02-21 17:34 | 0.63075 | 0 | 0.73029 | 2.7266e-06 | 6405 | 0.7369 | 0.7303 | 0.73029 | 0.73029 | 0.73029 | 0.73028
| 1000 | 1000 | 2026-02-21 17:34 | 0.89864 | 0 | 0.7667 | 3.6595e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 6420000 | 6420000 | 2026-02-21 17:34 | 0.61781 | 0 | 0.73011 | 1.7197e-06 | 6420 | 0.72197 | 0.73011 | 0.7301 | 0.73011 | 0.7301 | 0.7301
| 1000 | 1000 | 2026-02-21 17:35 | 0.89861 | 0 | 0.76668 | 2.3569e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76668 | 0.76669
| 6435000 | 6435000 | 2026-02-21 17:35 | 0.60469 | 0 | 0.72999 | 3.7301e-06 | 6435 | 0.70674 | 0.73 | 0.73 | 0.72999 | 0.72999 | 0.72998
| 1000 | 1000 | 2026-02-21 17:36 | 0.89864 | 0 | 0.76671 | 3.7384e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.76671 | 0.76672 | 0.76672
| 6450000 | 6450000 | 2026-02-21 17:36 | 0.59136 | 0 | 0.72991 | 1.459e-06 | 6450 | 0.69123 | 0.72991 | 0.72991 | 0.72991 | 0.72991 | 0.72991
| 1000 | 1000 | 2026-02-21 17:37 | 0.89865 | 0 | 0.76672 | 1.9007e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 6465000 | 6465000 | 2026-02-21 17:37 | 0.57779 | 0 | 0.72981 | 2.1933e-06 | 6465 | 0.67547 | 0.72981 | 0.72981 | 0.72981 | 0.7298 | 0.7298
| 1000 | 1000 | 2026-02-21 17:37 | 0.89867 | 0 | 0.76673 | 1.6926e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 6480000 | 6480000 | 2026-02-21 17:38 | 0.564 | 0 | 0.72968 | 8.4567e-07 | 6480 | 0.65947 | 0.72968 | 0.72969 | 0.72969 | 0.72968 | 0.72968
| 1000 | 1000 | 2026-02-21 17:38 | 0.89869 | 0 | 0.76675 | 2.0433e-06 | #NAME? | #NAME? | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 6495000 | 6495000 | 2026-02-21 17:39 | 0.55 | 0 | 0.72953 | 2.8316e-06 | 6495 | 0.64323 | 0.72954 | 0.72953 | 0.72953 | 0.72952 | 0.72952
| 1000 | 1000 | 2026-02-21 17:39 | 0.89867 | 0 | 0.76673 | 2.1184e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 6510000 | 6510000 | 2026-02-21 17:39 | 0.5358 | 0 | 0.72934 | 2.5255e-06 | 6510 | 0.62678 | 0.72935 | 0.72935 | 0.72934 | 0.72934 | 0.72934
| 1000 | 1000 | 2026-02-21 17:40 | 0.89866 | 0 | 0.76672 | 2.9778e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76673 | 0.76673
| 6525000 | 6525000 | 2026-02-21 17:40 | 0.52149 | 0 | 0.72923 | 3.4171e-06 | 6525 | 0.61013 | 0.72924 | 0.72923 | 0.72923 | 0.72922 | 0.72922
| 1000 | 1000 | 2026-02-21 17:41 | 0.89864 | 0 | 0.76671 | 1.658e-06 | #NAME? | #NAME? | 0.76671 | 0.7667 | 0.76671 | 0.76671 | 0.76671
| 6540000 | 6540000 | 2026-02-21 17:41 | 0.50702 | 0 | 0.72908 | 3.2395e-06 | 6540 | 0.59332 | 0.72909 | 0.72909 | 0.72909 | 0.72908 | 0.72908
| 1000 | 1000 | 2026-02-21 17:42 | 0.89862 | 0 | 0.76669 | 2.5334e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 6555000 | 6555000 | 2026-02-21 17:42 | 0.4924 | 0 | 0.72891 | 3.3081e-06 | 6555 | 0.57635 | 0.72892 | 0.72891 | 0.7289 | 0.7289 | 0.7289
| 1000 | 1000 | 2026-02-21 17:42 | 0.89863 | 0 | 0.76669 | 7.8061e-07 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 6570000 | 6570000 | 2026-02-21 17:43 | 0.47773 | 0 | 0.72881 | 1.6611e-06 | 6570 | 0.55925 | 0.72881 | 0.72881 | 0.72881 | 0.72881 | 0.7288
| 1000 | 1000 | 2026-02-21 17:43 | 0.89862 | 0 | 0.76669 | 1.7484e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 6585000 | 6585000 | 2026-02-21 17:44 | 0.46297 | 0 | 0.72875 | 2.2864e-06 | 6585 | 0.54203 | 0.72874 | 0.72874 | 0.72875 | 0.72875 | 0.72875
| 1000 | 1000 | 2026-02-21 17:44 | 0.89874 | 0 | 0.7668 | 3.3485e-06 | #NAME? | #NAME? | 0.76679 | 0.7668 | 0.7668 | 0.7668 | 0.76681
| 6600000 | 6600000 | 2026-02-21 17:44 | 0.44812 | 0 | 0.72864 | 7.7808e-07 | 6600 | 0.52472 | 0.72864 | 0.72864 | 0.72864 | 0.72864 | 0.72864
| 1000 | 1000 | 2026-02-21 17:45 | 0.8987 | 0 | 0.76676 | 9.9073e-07 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76676 | 0.76676 | 0.76676
| 6615000 | 6615000 | 2026-02-21 17:45 | 0.43313 | 0 | 0.72841 | 3.7804e-06 | 6615 | 0.50732 | 0.72842 | 0.72842 | 0.72841 | 0.72841 | 0.7284
| 1000 | 1000 | 2026-02-21 17:46 | 0.89866 | 0 | 0.76672 | 2.3229e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 6630000 | 6630000 | 2026-02-21 17:46 | 0.41814 | 0 | 0.72824 | 1.3683e-06 | 6630 | 0.48988 | 0.72824 | 0.72825 | 0.72824 | 0.72824 | 0.72824
| 1000 | 1000 | 2026-02-21 17:47 | 0.89864 | 0 | 0.76671 | 1.6619e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76672 | 0.76671 | 0.76672
| 6645000 | 6645000 | 2026-02-21 17:47 | 0.40313 | 0 | 0.72808 | 3.0785e-06 | 6645 | 0.47241 | 0.72808 | 0.72808 | 0.72809 | 0.72808 | 0.72807
| 1000 | 1000 | 2026-02-21 17:47 | 0.89861 | 0 | 0.76669 | 1.5648e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 6660000 | 6660000 | 2026-02-21 17:48 | 0.38813 | 0 | 0.72796 | 5.1849e-07 | 6660 | 0.45491 | 0.72796 | 0.72796 | 0.72796 | 0.72796 | 0.72796
| 1000 | 1000 | 2026-02-21 17:48 | 0.89866 | 0 | 0.76673 | 1.3454e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 6675000 | 6675000 | 2026-02-21 17:49 | 0.37319 | 0 | 0.72792 | 3.169e-06 | 6675 | 0.43741 | 0.72793 | 0.72793 | 0.72792 | 0.72792 | 0.72791
| 1000 | 1000 | 2026-02-21 17:49 | 0.89872 | 0 | 0.76677 | 6.1753e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76677 | 0.76678 | 0.76678
| 6690000 | 6690000 | 2026-02-21 17:49 | 0.3582 | 0 | 0.72773 | 3.0042e-06 | 6690 | 0.41996 | 0.72774 | 0.72773 | 0.72773 | 0.72773 | 0.72772
| 1000 | 1000 | 2026-02-21 17:50 | 0.89866 | 0 | 0.76673 | 2.5238e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76673 | 0.76673 | 0.76673
| 6705000 | 6705000 | 2026-02-21 17:50 | 0.34327 | 0 | 0.72757 | 1.4839e-06 | 6705 | 0.40254 | 0.72757 | 0.72757 | 0.72757 | 0.72757 | 0.72757
| 1000 | 1000 | 2026-02-21 17:51 | 0.89863 | 0 | 0.7667 | 2.7399e-06 | #NAME? | #NAME? | 0.76669 | 0.76671 | 0.7667 | 0.7667 | 0.76671
| 6720000 | 6720000 | 2026-02-21 17:51 | 0.32846 | 0 | 0.72753 | 1.7215e-06 | 6720 | 0.3852 | 0.72753 | 0.72753 | 0.72753 | 0.72753 | 0.72752
| 1000 | 1000 | 2026-02-21 17:52 | 0.89867 | 0 | 0.76674 | 1.5408e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76674
| 6735000 | 6735000 | 2026-02-21 17:52 | 0.31365 | 0 | 0.72732 | 3.6657e-06 | 6735 | 0.36793 | 0.72732 | 0.72732 | 0.72732 | 0.72731 | 0.72731
| 1000 | 1000 | 2026-02-21 17:52 | 0.89864 | 0 | 0.76671 | 2.2997e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76672
| 6750000 | 6750000 | 2026-02-21 17:53 | 0.29896 | 0 | 0.72714 | 1.6191e-06 | 6750 | 0.35077 | 0.72714 | 0.72714 | 0.72715 | 0.72714 | 0.72714
| 1000 | 1000 | 2026-02-21 17:53 | 0.89867 | 0 | 0.76673 | 1.9902e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 6765000 | 6765000 | 2026-02-21 17:54 | 0.2844 | 0 | 0.72704 | 3.6357e-06 | 6765 | 0.33375 | 0.72705 | 0.72704 | 0.72704 | 0.72703 | 0.72703
| 1000 | 1000 | 2026-02-21 17:54 | 0.89866 | 0 | 0.76672 | 1.341e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 6780000 | 6780000 | 2026-02-21 17:54 | 0.26995 | 0 | 0.72687 | 4.3989e-06 | 6780 | 0.31686 | 0.72688 | 0.72687 | 0.72687 | 0.72686 | 0.72686
| 1000 | 1000 | 2026-02-21 17:55 | 0.89863 | 0 | 0.76669 | 3.7004e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.76671
| 6795000 | 6795000 | 2026-02-21 17:55 | 0.25565 | 0 | 0.72671 | 1.5848e-06 | 6795 | 0.30014 | 0.72672 | 0.72672 | 0.72671 | 0.72671 | 0.72671
| 1000 | 1000 | 2026-02-21 17:56 | 0.89862 | 0 | 0.76669 | 3.2498e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.7667
| 6810000 | 6810000 | 2026-02-21 17:56 | 0.24153 | 0 | 0.72661 | 1.6114e-06 | 6810 | 0.28361 | 0.72661 | 0.72661 | 0.72661 | 0.7266 | 0.7266
| 1000 | 1000 | 2026-02-21 17:57 | 0.89872 | 0 | 0.76678 | 2.5522e-06 | #NAME? | #NAME? | 0.76677 | 0.76678 | 0.76678 | 0.76678 | 0.76679
| 6825000 | 6825000 | 2026-02-21 17:57 | 0.22759 | 0 | 0.72652 | 3.8747e-06 | 6825 | 0.26727 | 0.72653 | 0.72652 | 0.72651 | 0.72651 | 0.7265
| 1000 | 1000 | 2026-02-21 17:57 | 0.89869 | 0 | 0.76674 | 2.8327e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76675 | 0.76675
| 6840000 | 6840000 | 2026-02-21 17:58 | 0.21383 | 0 | 0.7264 | 2.3494e-06 | 6840 | 0.25115 | 0.7264 | 0.7264 | 0.7264 | 0.72639 | 0.72639
| 1000 | 1000 | 2026-02-21 17:58 | 0.89865 | 0 | 0.76671 | 1.9722e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 6855000 | 6855000 | 2026-02-21 17:59 | 0.20025 | 0 | 0.72622 | 1.9993e-06 | 6855 | 0.23527 | 0.72622 | 0.72622 | 0.72622 | 0.72621 | 0.72621
| 1000 | 1000 | 2026-02-21 17:59 | 0.89865 | 0 | 0.76672 | 8.9857e-07 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76672 | 0.76672
| 6870000 | 6870000 | 2026-02-21 17:59 | 0.18691 | 0 | 0.72608 | 1.4508e-06 | 6870 | 0.21963 | 0.72608 | 0.72608 | 0.72608 | 0.72608 | 0.72607
| 1000 | 1000 | 2026-02-21 18:00 | 0.89866 | 0 | 0.76672 | 3.8825e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76673 | 0.76673
| 6885000 | 6885000 | 2026-02-21 18:00 | 0.17378 | 0 | 0.72591 | 2.3218e-06 | 6885 | 0.20426 | 0.72591 | 0.72591 | 0.72591 | 0.7259 | 0.7259
| 1000 | 1000 | 2026-02-21 18:01 | 0.89868 | 0 | 0.76674 | 3.1929e-06 | #NAME? | #NAME? | 0.76673 | 0.76673 | 0.76674 | 0.76674 | 0.76675
| 6900000 | 6900000 | 2026-02-21 18:01 | 0.16093 | 0 | 0.72582 | 4.3613e-07 | 6900 | 0.18917 | 0.72582 | 0.72582 | 0.72582 | 0.72582 | 0.72582
| 1000 | 1000 | 2026-02-21 18:01 | 0.89867 | 0 | 0.76674 | 1.0438e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76673 | 0.76674 | 0.76673
| 6915000 | 6915000 | 2026-02-21 18:02 | 0.1483 | 0 | 0.72563 | 2.4174e-06 | 6915 | 0.17437 | 0.72564 | 0.72564 | 0.72563 | 0.72563 | 0.72562
| 1000 | 1000 | 2026-02-21 18:02 | 0.89863 | 0 | 0.76669 | 3.9069e-06 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.76669 | 0.7667 | 0.76671
| 6930000 | 6930000 | 2026-02-21 18:03 | 0.13595 | 0 | 0.72549 | 2.2179e-06 | 6930 | 0.15988 | 0.7255 | 0.7255 | 0.72549 | 0.72549 | 0.72549
| 1000 | 1000 | 2026-02-21 18:03 | 0.89862 | 0 | 0.76669 | 1.3985e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.76669 | 0.76669
| 6945000 | 6945000 | 2026-02-21 18:04 | 0.12389 | 0 | 0.72538 | 2.3703e-06 | 6945 | 0.14572 | 0.72539 | 0.72538 | 0.72537 | 0.72538 | 0.72537
| 1000 | 1000 | 2026-02-21 18:04 | 0.89866 | 0 | 0.76672 | 3.0009e-06 | #NAME? | #NAME? | 0.76671 | 0.76672 | 0.76672 | 0.76672 | 0.76673
| 6960000 | 6960000 | 2026-02-21 18:04 | 0.11212 | 0 | 0.72532 | 2.008e-06 | 6960 | 0.13189 | 0.72532 | 0.72532 | 0.72532 | 0.72532 | 0.72531
| 1000 | 1000 | 2026-02-21 18:05 | 0.89868 | 0 | 0.76675 | 1.0176e-06 | #NAME? | #NAME? | 0.76675 | 0.76675 | 0.76675 | 0.76675 | 0.76675
| 6975000 | 6975000 | 2026-02-21 18:05 | 0.10063 | 0 | 0.72512 | 2.8134e-06 | 6975 | 0.1184 | 0.72513 | 0.72512 | 0.72512 | 0.72512 | 0.72511
| 1000 | 1000 | 2026-02-21 18:06 | 0.89871 | 0 | 0.76677 | 2.6238e-06 | #NAME? | #NAME? | 0.76676 | 0.76676 | 0.76677 | 0.76677 | 0.76678
| 6990000 | 6990000 | 2026-02-21 18:06 | 0.089457 | 0 | 0.72501 | 2.5812e-06 | 6990 | 0.10527 | 0.72502 | 0.72502 | 0.72501 | 0.72502 | 0.72501
| 1000 | 1000 | 2026-02-21 18:06 | 0.89868 | 0 | 0.76674 | 5.8102e-06 | #NAME? | #NAME? | 0.76672 | 0.76673 | 0.76674 | 0.76675 | 0.76676
| 7005000 | 7005000 | 2026-02-21 18:07 | 0.07859 | 0 | 0.72486 | 9.7905e-07 | 7005 | 0.092505 | 0.72486 | 0.72486 | 0.72486 | 0.72486 | 0.72486
| 1000 | 1000 | 2026-02-21 18:07 | 0.89865 | 0 | 0.76673 | 4.3016e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76672 | 0.76673 | 0.76675
| 7020000 | 7020000 | 2026-02-21 18:08 | Error | Error | 0.72471 | 2.5179e-06 | 7020 | #VALUE! | 0.72472 | 0.72471 | 0.72471 | 0.7247 | 0.7247
| 1000 | 1000 | 2026-02-21 18:08 | 0.8986 | 0 | 0.76667 | 3.1727e-06 | #NAME? | #NAME? | 0.76666 | 0.76666 | 0.76667 | 0.76668 | 0.76668
| 7035000 | 7035000 | 2026-02-21 18:09 | Error | Error | 0.7245 | 2.4911e-06 | 7035 | #VALUE! | 0.7245 | 0.7245 | 0.7245 | 0.72451 | 0.72451
| 1000 | 1000 | 2026-02-21 18:10 | 0.89856 | 0 | 0.76664 | 2.248e-06 | #NAME? | #NAME? | 0.76663 | 0.76664 | 0.76664 | 0.76664 | 0.76665
| 7050000 | 7050000 | 2026-02-21 18:10 | Error | Error | 0.72439 | 2.0723e-06 | 7050 | #VALUE! | 0.72438 | 0.72439 | 0.72439 | 0.72439 | 0.72439
| 1000 | 1000 | 2026-02-21 18:11 | 0.89853 | 0 | 0.76661 | 2.7678e-06 | #NAME? | #NAME? | 0.7666 | 0.76661 | 0.76661 | 0.76661 | 0.76662
| 7065000 | 7065000 | 2026-02-21 18:12 | Error | Error | 0.72425 | 1.3712e-06 | 7065 | #VALUE! | 0.72425 | 0.72425 | 0.72425 | 0.72425 | 0.72424
| 1000 | 1000 | 2026-02-21 18:12 | 0.89856 | 0 | 0.76664 | 3.06e-06 | #NAME? | #NAME? | 0.76664 | 0.76663 | 0.76664 | 0.76665 | 0.76665
| 7080000 | 7080000 | 2026-02-21 18:13 | Error | Error | 0.72414 | 1.7093e-06 | 7080 | #VALUE! | 0.72414 | 0.72414 | 0.72414 | 0.72414 | 0.72413
| 1000 | 1000 | 2026-02-21 18:13 | 0.89858 | 0 | 0.76666 | 2.0216e-06 | #NAME? | #NAME? | 0.76665 | 0.76666 | 0.76666 | 0.76666 | 0.76666
| 7095000 | 7095000 | 2026-02-21 18:14 | Error | Error | 0.72399 | 2.7199e-06 | 7095 | #VALUE! | 0.724 | 0.724 | 0.72399 | 0.72399 | 0.72398
| 1000 | 1000 | 2026-02-21 18:14 | 0.8986 | 0 | 0.76667 | 1.9734e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76667 | 0.76668 | 0.76668
| 7110000 | 7110000 | 2026-02-21 18:15 | Error | Error | 0.72388 | 3.0448e-06 | 7110 | #VALUE! | 0.72389 | 0.72389 | 0.72388 | 0.72388 | 0.72388
| 1000 | 1000 | 2026-02-21 18:15 | 0.89857 | 0 | 0.76665 | 9.7803e-07 | #NAME? | #NAME? | 0.76665 | 0.76666 | 0.76665 | 0.76665 | 0.76666
| 7125000 | 7125000 | 2026-02-21 18:16 | Error | Error | 0.72373 | 1.5102e-06 | 7125 | #VALUE! | 0.72373 | 0.72373 | 0.72373 | 0.72372 | 0.72372
| 1000 | 1000 | 2026-02-21 18:17 | 0.89859 | 0 | 0.76666 | 5.5253e-06 | #NAME? | #NAME? | 0.76665 | 0.76665 | 0.76666 | 0.76667 | 0.76668
| 7140000 | 7140000 | 2026-02-21 18:17 | Error | Error | 0.72362 | 2.1224e-06 | 7140 | #VALUE! | 0.72362 | 0.72362 | 0.72362 | 0.72362 | 0.72361
| 1000 | 1000 | 2026-02-21 18:18 | 0.89858 | 0 | 0.76665 | 2.4843e-06 | #NAME? | #NAME? | 0.76665 | 0.76664 | 0.76664 | 0.76665 | 0.76666
| 7155000 | 7155000 | 2026-02-21 18:18 | Error | Error | 0.72349 | 1.6531e-06 | 7155 | #VALUE! | 0.72349 | 0.72349 | 0.72349 | 0.72349 | 0.72349
| 1000 | 1000 | 2026-02-21 18:19 | 0.8986 | 0 | 0.76667 | 4.0074e-06 | #NAME? | #NAME? | 0.76666 | 0.76667 | 0.76668 | 0.76668 | 0.76668
| 7170000 | 7170000 | 2026-02-21 18:20 | Error | Error | 0.72332 | 3.4634e-06 | 7170 | #VALUE! | 0.72333 | 0.72333 | 0.72332 | 0.72332 | 0.72331
| 1000 | 1000 | 2026-02-21 18:20 | 0.89857 | 0 | 0.76665 | 1.6744e-06 | #NAME? | #NAME? | 0.76664 | 0.76664 | 0.76665 | 0.76665 | 0.76665
| 7185000 | 7185000 | 2026-02-21 18:21 | Error | Error | 0.72317 | 2.775e-06 | 7185 | #VALUE! | 0.72318 | 0.72317 | 0.72316 | 0.72316 | 0.72316
| 1000 | 1000 | 2026-02-21 18:21 | 0.89859 | 0 | 0.76666 | 2.291e-06 | #NAME? | #NAME? | 0.76666 | 0.76666 | 0.76666 | 0.76667 | 0.76667
| 7200000 | 7200000 | 2026-02-21 18:22 | Error | Error | 0.72312 | 1.4225e-06 | 7200 | #VALUE! | 0.72312 | 0.72312 | 0.72312 | 0.72312 | 0.72312
| 1000 | 1000 | 2026-02-21 18:22 | 0.89861 | 0 | 0.76669 | 1.739e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76669 | 0.76669 | 0.76669
| 7215000 | 7215000 | 2026-02-21 18:23 | Error | Error | 0.723 | 3.5264e-06 | 7215 | #VALUE! | 0.72301 | 0.723 | 0.723 | 0.72299 | 0.72299
| 1000 | 1000 | 2026-02-21 18:24 | 0.89871 | 0 | 0.76676 | 2.6661e-06 | #NAME? | #NAME? | 0.76675 | 0.76676 | 0.76676 | 0.76677 | 0.76677
| 7230000 | 7230000 | 2026-02-21 18:24 | Error | Error | 0.72284 | 2.4016e-06 | 7230 | #VALUE! | 0.72285 | 0.72285 | 0.72284 | 0.72284 | 0.72284
| 1000 | 1000 | 2026-02-21 18:25 | 0.89862 | 0 | 0.76669 | 9.3113e-07 | #NAME? | #NAME? | 0.76669 | 0.76669 | 0.7667 | 0.76669 | 0.76669
| 7245000 | 7245000 | 2026-02-21 18:26 | Error | Error | 0.72264 | 2.3814e-06 | 7245 | #VALUE! | 0.72265 | 0.72265 | 0.72264 | 0.72264 | 0.72264
| 1000 | 1000 | 2026-02-21 18:26 | 0.89857 | 0 | 0.76664 | 2.8221e-06 | #NAME? | #NAME? | 0.76663 | 0.76664 | 0.76664 | 0.76664 | 0.76665
| 7260000 | 7260000 | 2026-02-21 18:27 | Error | Error | 0.72255 | 3.4992e-06 | 7260 | #VALUE! | 0.72256 | 0.72255 | 0.72255 | 0.72254 | 0.72254
| 1000 | 1000 | 2026-02-21 18:27 | 0.89861 | 0 | 0.76668 | 3.1503e-06 | #NAME? | #NAME? | 0.76667 | 0.76667 | 0.76668 | 0.76669 | 0.76669
| 7275000 | 7275000 | 2026-02-21 18:28 | Error | Error | 0.72242 | 2.0333e-06 | 7275 | #VALUE! | 0.72243 | 0.72242 | 0.72242 | 0.72242 | 0.72242
| 1000 | 1000 | 2026-02-21 18:28 | 0.89859 | 0 | 0.76666 | 1.9248e-06 | #NAME? | #NAME? | 0.76666 | 0.76667 | 0.76666 | 0.76666 | 0.76667
| 7290000 | 7290000 | 2026-02-21 18:29 | Error | Error | 0.7223 | 1.6569e-06 | 7290 | #VALUE! | 0.72229 | 0.72229 | 0.7223 | 0.7223 | 0.7223
| 1000 | 1000 | 2026-02-21 18:29 | 0.89862 | 0 | 0.76668 | 4.2808e-06 | #NAME? | #NAME? | 0.76667 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 7305000 | 7305000 | 2026-02-21 18:30 | Error | Error | 0.72214 | 2.6329e-06 | 7305 | #VALUE! | 0.72215 | 0.72215 | 0.72214 | 0.72213 | 0.72213
| 1000 | 1000 | 2026-02-21 18:31 | 0.89856 | 0 | 0.76663 | 3.6399e-06 | #NAME? | #NAME? | 0.76662 | 0.76663 | 0.76663 | 0.76664 | 0.76664
| 7320000 | 7320000 | 2026-02-21 18:31 | Error | Error | 0.72211 | 1.1704e-06 | 7320 | #VALUE! | 0.72211 | 0.72211 | 0.72211 | 0.72211 | 0.72211
| 1000 | 1000 | 2026-02-21 18:32 | 0.89864 | 0 | 0.76671 | 1.1048e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 7335000 | 7335000 | 2026-02-21 18:32 | 0.073806 | 0 | 0.72197 | 1.6179e-06 | 7335 | 0.087221 | 0.72197 | 0.72197 | 0.72197 | 0.72197 | 0.72198
| 1000 | 1000 | 2026-02-21 18:33 | 0.89876 | 0 | 0.76682 | 2.3736e-06 | #NAME? | #NAME? | 0.76681 | 0.76681 | 0.76682 | 0.76682 | 0.76682
| 7350000 | 7350000 | 2026-02-21 18:33 | 0.076832 | 0 | 0.72182 | 3.904e-06 | 7350 | 0.090815 | 0.72183 | 0.72183 | 0.72182 | 0.72182 | 0.72181
| 1000 | 1000 | 2026-02-21 18:33 | 0.89866 | 0 | 0.76672 | 3.063e-06 | #NAME? | #NAME? | 0.76672 | 0.76672 | 0.76673 | 0.76673 | 0.76673
| 7365000 | 7365000 | 2026-02-21 18:34 | 0.079548 | 0 | 0.72161 | 2.9446e-06 | 7365 | 0.094053 | 0.72162 | 0.72161 | 0.72161 | 0.72161 | 0.7216
| 1000 | 1000 | 2026-02-21 18:34 | 0.89861 | 0 | 0.76668 | 2.3973e-06 | #NAME? | #NAME? | 0.76668 | 0.76668 | 0.76668 | 0.76669 | 0.76669
| 7380000 | 7380000 | 2026-02-21 18:35 | 0.081985 | 0 | 0.7215 | 2.2836e-06 | 7380 | 0.096948 | 0.72151 | 0.72151 | 0.7215 | 0.7215 | 0.7215
| 1000 | 1000 | 2026-02-21 18:35 | 0.89863 | 0 | 0.76669 | 3.4665e-06 | #NAME? | #NAME? | 0.76668 | 0.76669 | 0.76669 | 0.7667 | 0.7667
| 7395000 | 7395000 | 2026-02-21 18:36 | 0.084131 | 0 | 0.7214 | 3.0806e-06 | 7395 | 0.099501 | 0.72141 | 0.7214 | 0.7214 | 0.72139 | 0.72139
| 1000 | 1000 | 2026-02-21 18:36 | 0.89863 | 0 | 0.7667 | 2.1005e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.7667 | 0.76671
| 7410000 | 7410000 | 2026-02-21 18:36 | 0.08599 | 0 | 0.72123 | 2.1594e-06 | 7410 | 0.10172 | 0.72124 | 0.72124 | 0.72124 | 0.72123 | 0.72123
| 1000 | 1000 | 2026-02-21 18:37 | 0.89864 | 0 | 0.7667 | 4.4385e-06 | #NAME? | #NAME? | 0.76669 | 0.7667 | 0.7667 | 0.76671 | 0.76672
| 7425000 | 7425000 | 2026-02-21 18:37 | 0.087587 | 0 | 0.72117 | 2.0965e-06 | 7425 | 0.10362 | 0.72118 | 0.72118 | 0.72118 | 0.72117 | 0.72117
| 1000 | 1000 | 2026-02-21 18:38 | 0.89865 | 0 | 0.76671 | 2.1138e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76672 | 0.76672
| 7440000 | 7440000 | 2026-02-21 18:38 | 0.08892 | 0 | 0.72105 | 3.0696e-06 | 7440 | 0.10521 | 0.72106 | 0.72106 | 0.72105 | 0.72105 | 0.72104
| 1000 | 1000 | 2026-02-21 18:38 | 0.89864 | 0 | 0.76671 | 1.5922e-06 | #NAME? | #NAME? | 0.76671 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 7455000 | 7455000 | 2026-02-21 18:39 | 0.089999 | 0 | 0.72095 | 1.9348e-06 | 7455 | 0.10651 | 0.72095 | 0.72094 | 0.72094 | 0.72094 | 0.72094
| 1000 | 1000 | 2026-02-21 18:39 | 0.89867 | 0 | 0.76674 | 2.4007e-06 | #NAME? | #NAME? | 0.76674 | 0.76674 | 0.76674 | 0.76674 | 0.76675
| 7470000 | 7470000 | 2026-02-21 18:40 | 0.090824 | 0 | 0.72075 | 1.3048e-06 | 7470 | 0.10751 | 0.72076 | 0.72076 | 0.72075 | 0.72075 | 0.72075
| 1000 | 1000 | 2026-02-21 18:40 | 0.89864 | 0 | 0.76671 | 2.7981e-06 | #NAME? | #NAME? | 0.7667 | 0.7667 | 0.7667 | 0.76671 | 0.76671
| 7485000 | 7485000 | 2026-02-21 18:41 | 0.091419 | 0 | 0.72061 | 2.9205e-06 | 7485 | 0.10824 | 0.72061 | 0.72061 | 0.72061 | 0.7206 | 0.7206
| 1000 | 1000 | 2026-02-21 18:41 | 0.89864 | 0 | 0.76671 | 2.0949e-06 | #NAME? | #NAME? | 0.7667 | 0.76671 | 0.76671 | 0.76671 | 0.76671
| 7500000 | 7500000 | 2026-02-21 18:41 | 0.091801 | 0 | 0.72054 | 1.5608e-06 | 7500 | 0.1087 | 0.72053 | 0.72054 | 0.72054 | 0.72054 | 0.72054
STOP | 1000 | 1000 | 2026-02-21 18:42 | 0.89869 | 0 | 0.76675 | 2.4324e-06 | 1 | #DIV/0! | 0.76674 | 0.76675 | 0.76675 | 0.76675 | 0.76676
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
[8 empty] | #NAME? | #NAME?
Note on D2: labina: Real sine amplitude, i.e. A*sin(w*t). It is NOT rms or peak-to-peak value.
Note on D7: labina: The algorithm must generate output quantity 'rms'. That is e.g. TWM-WFFT or TWM-WRMS
Note on D8: labina: Set to 1 to measure simultaneously both Digitizer amplitude and DC multimeter voltage. Set to 0 or keep empty to measure first AC and DC after.
Note on C12: Martin Šíra: The record length should be such to get coherent sampling. If the number of periods is the same as the signal frequency, coherent sampling should be always met, however the data processing can take a lot of time.

## PTB reference data
10 | 1
70 | 1.0005
130 | 1.0013
190 | 1.0028
250 | 1.0047
310 | 1.0069
370 | 1.0091
430 | 1.0112
490 | 1.0129
550 | 1.0141
610 | 1.0146
670 | 1.0143
730 | 1.0131
790 | 1.0112
850 | 1.0085
910 | 1.0053
970 | 1.0018
1030 | 0.99823
1090 | 0.99486
1150 | 0.99194
1210 | 0.9897
1270 | 0.9883
1330 | 0.98787
1390 | 0.98845
1450 | 0.99
1510 | 0.99243
1570 | 0.99558
1630 | 0.99921
1690 | 1.0031
1750 | 1.0069
1810 | 1.0104
1870 | 1.0133
1930 | 1.0154
1990 | 1.0166
2050 | 1.0167
2110 | 1.0158
2170 | 1.0138
2230 | 1.0111
2290 | 1.0078
2350 | 1.0041
2410 | 1.0003
2470 | 0.99686
2530 | 0.9939
2590 | 0.99171
2650 | 0.99048
2710 | 0.9903
2770 | 0.99122
2830 | 0.99315
2890 | 0.99597
2950 | 0.99945
3010 | 1.0033
3070 | 1.0073
3130 | 1.0111
3190 | 1.0144
3250 | 1.0169
3310 | 1.0185
3370 | 1.019
3430 | 1.0185
3490 | 1.0168
3550 | 1.0143
3610 | 1.011
3670 | 1.0074
3730 | 1.0036
3790 | 0.99995
3850 | 0.99682
3910 | 0.99444
3970 | 0.99302
4030 | 0.99269
4090 | 0.99349
4150 | 0.99537
4210 | 0.99819
4270 | 1.0017
4330 | 1.0057
4390 | 1.0098
4450 | 1.0137
4510 | 1.0171
4570 | 1.0195
4630 | 1.0209
4690 | 1.0212
4750 | 1.0201
4810 | 1.0179
4870 | 1.0147
4930 | 1.0108
4990 | 1.0066
5050 | 1.0025
5110 | 0.99885
5170 | 0.9961
5230 | 0.99456
5290 | 0.99443
5350 | 0.99578
5410 | 0.99854
5470 | 1.0025
5530 | 1.0072
5590 | 1.0121
5650 | 1.0168
5710 | 1.0202
5770 | 1.0216
5830 | 1.0198
5890 | 1.0135
5950 | 1.002
6010 | 0.98585
6070 | 0.97016
6130 | 0.97061